## Master - Study Demographics
Data Extractor | Data Validated by | Article # | Article First Author | Article Title | Publication Month/Year | DOI/PMID/Other identifier | Dominant Variant | Country | Setting/Context - Use multiple lines for multiple settings | Year/Time frame for data collection | Participant characteristics (age range, gender breakdown, rural/urban, etc) | Number of Participants | Sample Type - Use multiple lines for multiple sample types | Sample type if Other | Reference Test description - Use multiple lines for multiple Reference Tests | Reference Test if other | Reference test comments (if any) | Description of main results (include adverse events from tests) | Exclusion reasons (if any) | Notes | Need to contact Authors? Put contact info here
Amy | Sherli | 5 | Abeer Mohamed Abdelrazik | Potential Use of Antigen-Based Rapid Test for SARS-CoV-2 in Respiratory Specimens in Low-Resource Settings in Egypt for Symptomatic Patients and High-Risk Contacts | 2021-03-01 | https://doi.org/10.1093/labmed/lmaa104 | Ancestral | Egypt | College/University Campus (Medical Center/Hospital) | 2020-05-01 | 59.4% male; Median age 42 | 310 | Nasopharyngeal swabs (NP) |  | Specific Test Not Described |  |  | RAT should not be used alone in clinical settings, though there may be value in using with symptomatic pts only | N/A
Terry | Sherli | 7 | Abusrewil Z. | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-12-01 | doi: 10.1002/jmv.27186 PMID: 34241912 | Unknown - Ancestral best guess | Libya | College/University Campus (non-medical) | Oct - Dec 2020 | 56,2% male; 43.7% female; 40.8 yrs old with SD 15.23 | 231 | Nasopharyngeal swabs (NP) |  | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Terry | Sherli | 7 | Abusrewil Z. | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | doi: 10.1002/jmv.27186 PMID: 34241912 | [8 empty] | DAAN Gene RT-PCR COVID-19 (DaAnGene)
Tim | Sherli | 10 | Y. Akashi; M. Horie; Y. Takeuchi; K. Togashi; Y. Adachi; A. Ueda; S. Notake; K. Nakamura; N. Terada; Y. Kurihara; Y. Kiyasu; H. Suzuki | A prospective clinical evaluation of the diagnostic accuracy of the SARS-CoV-2 rapid antigen test using anterior nasal samples | 2022-02-01 | https://doi.org/10.1016/j.jiac.2022.02.016 |  | Japan | College/University Campus (Medical Center/Hospital) | July 7 - July 29, 2021 | suspected of contracted COVID-19 based on symptoms or close contact. IQR Age Range - 25.0-50.0. Female - 341 (42.6%). | 800 | Nasal Swabs (NS) | 2 anterior nasal swabs done. One with FLOQSwab for RT-PCR. | Japanese National Institute of Infectious Diseases (NIID) |  |  | Rapid antigen testing with anterior nasal sample had favorable diagnostic accuracy for SARS-CoV-2. |  |  | Yusaku Akashi yusaku-akashi@umin.ac.jp
Tim | Sherli | 10 | Y. Akashi; M. Horie; Y. Takeuchi; K. Togashi; Y. Adachi; A. Ueda; S. Notake; K. Nakamura; N. Terada; Y. Kurihara; Y. Kiyasu; H. Suzuki | A prospective clinical evaluation of the diagnostic accuracy of the SARS-CoV-2 rapid antigen test using anterior nasal samples | 2022-02-01 |  |  |  | Primary Care Location |  |  |  | Nasopharyngeal swabs (NP) | NP also for RT-PCR
Tim | Sherli | 10 | Y. Akashi; M. Horie; Y. Takeuchi; K. Togashi; Y. Adachi; A. Ueda; S. Notake; K. Nakamura; N. Terada; Y. Kurihara; Y. Kiyasu; H. Suzuki | A prospective clinical evaluation of the diagnostic accuracy of the SARS-CoV-2 rapid antigen test using anterior nasal samples | 2022-02-01 |  |  |  | Hospital - Inpatient
Tim | Sherli | 10 | Y. Akashi; M. Horie; Y. Takeuchi; K. Togashi; Y. Adachi; A. Ueda; S. Notake; K. Nakamura; N. Terada; Y. Kurihara; Y. Kiyasu; H. Suzuki | A prospective clinical evaluation of the diagnostic accuracy of the SARS-CoV-2 rapid antigen test using anterior nasal samples | 2022-02-01 |  |  |  | Hospital - Outpatient
Amy | Sherli | 12 | S. Al-Alawi; H. Al-Hinai; N. Al-Kindi; M. Al-Rashidi; H. Al-Kindi; I. Al-Shukri; A. Al-Rashdi; S. Jose; A. Al-Jardani | Evaluation of four rapid antigen tests for detection of SARS-CoV-2 virus | 2021-01-01 | DOI 10.5001/omj.2021.106 |  | Oman | Primary Care Location | June - July 2020 | All suspected COVID-19 cases, with symptoms or pneumonia | 306 Swabs | Nasopharyngeal swabs (NP) |  | Novel Coronavirus (2019-nCoV) Real Time Multiplex RT-PCR kit (Liferiver) |  |  | RAT worked well for high viral load (aka low CTs), but not well for low viral load (high CTs)
Amy | Sherli | 12 | S. Al-Alawi; H. Al-Hinai; N. Al-Kindi; M. Al-Rashidi; H. Al-Kindi; I. Al-Shukri; A. Al-Rashdi; S. Jose; A. Al-Jardani | Evaluation of four rapid antigen tests for detection of SARS-CoV-2 virus | 2021-01-02 | https://doi.org/10.1093/labmed/lmaa104 |  |  | Not described/unclear |  |  |  | Throat Swabs (TS) |  | Sansure Biotech COVID-19 Nucleic Acid Test kit
Amy | Sherli | 12 | S. Al-Alawi; H. Al-Hinai; N. Al-Kindi; M. Al-Rashidi; H. Al-Kindi; I. Al-Shukri; A. Al-Rashdi; S. Jose; A. Al-Jardani | Evaluation of four rapid antigen tests for detection of SARS-CoV-2 virus | 2021-01-03 | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Cobas Kits/Systems (Roche)
Amy | Sherli | 13 | S. Aleem; N. Zahoor; A. Jeelani; S. M. SalimKhan | Diagnostic Accuracy of STANDARD QCOVID-19 Antigen Detection Kit in Comparison with RT-PCR Assay using Nasopharyngeal Samples in India | 2022-01-01 | DOI: 10.7860/JCDR/2022/52286.15825 |  | India | Primary Care Location | April 1-30, 2021 | contacts of positive patients, and anyone on the facility requiring testing; 57.29% Urban; 18+; 58.56% males, mean age 38.4 | 473 | Nasopharyngeal swabs (NP) |  | Custom/In-house SARS-2 primers |  | E gene | Mean Ct for RAT+ = 19; Mean Ct for RAT- = 33; Incidence of COVID-19 16.7%
Amy | Sherli | 13 | S. Aleem; N. Zahoor; A. Jeelani; S. M. SalimKhan | Diagnostic Accuracy of STANDARD QCOVID-19 Antigen Detection Kit in Comparison with RT-PCR Assay using Nasopharyngeal Samples in India | 2022-01-01 |  |  |  | COVID-19 testing site/Screening location
Chathu | Sherli | 14 | M. Alghounaim; H. Bastaki; F. Bin Essa; H. Motlagh; S. Al-Sabah | The Performance of Two Rapid Antigen Tests During Population-Level Screening for SARS-CoV-2 Infection | 2021-12-01 | 10.3389/fmed.2021.797109 |  | Kuwait | COVID-19 testing site/Screening location | May 24 - August 12 2021 | Median age = 40.2yrs; 94% male | 991 individuals | Nasopharyngeal swabs (NP) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher)
Chathu | Sherli | 14 | M. Alghounaim; H. Bastaki; F. Bin Essa; H. Motlagh; S. Al-Sabah | The Performance of Two Rapid Antigen Tests During Population-Level Screening for SARS-CoV-2 Infection | 2021-12-01
Chathu | Sherli | 14 | M. Alghounaim; H. Bastaki; F. Bin Essa; H. Motlagh; S. Al-Sabah | The Performance of Two Rapid Antigen Tests During Population-Level Screening for SARS-CoV-2 Infection | 2021-12-01
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-01 |  |  | United States | COVID-19 testing site/Screening location | Between October and Novemebr 2020 | N/A | 18457 | Saliva (Sal) | Oral fluid | Custom/In-house SARS-2 primers |  |  | The positive percent agreement between Abbott Ag card and PCR is lower than in previous studies |  |  | Address correspondence to Lao-Tzu Allan-Blitz, lallan-blitz@partners.org.
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida |  | https://doi.org/10.1093/labmed/lmaa104 |  |  |  |  | Most likely urban-- named the counties they conducted the study in |  | Nasal Swabs (NS) | Anterior nares specimen | Custom/In-house SARS-2 primers
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida |  | https://doi.org/10.1093/labmed/lmaa104 |  |  |  |  | People presenting to test for Covid |  | Nasopharyngeal swabs (NP) |  | Custom/In-house SARS-2 primers
Ellyn | Sherli | 16 | Amer R. | Diagnostic performance of rapid antigen test for COVID-19 and the effect of viral load, sampling time, subject's clinical and laboratory parameters on test accuracy | 2021-10-01 | https://doi.org/10.1016/j.jiph.2021.06.002 | Ancestral | Egypt | College/University Campus (Medical Center/Hospital) | June 2020 - October 2020 | individuals (age 22-87) suspected of having COVID-19 infection | 83 | Nasopharyngeal swabs (NP) |  | Genesig Real-time PCR Coronavirus assay/Z-Path-COVID-19-CE (Primerdesign) |  |  | “StandardTM Q COVID-19 Ag” should not be used alone for COVID-19 diagnosis due to its low diagnostic performance relative to the RT-qPCR. RAT is best used at the early disease stage and in patients with high viral load."
Ellyn | Sherli | 16 | Amer R. | Diagnostic performance of rapid antigen test for COVID-19 and the effect of viral load, sampling time, subject's clinical and laboratory parameters on test accuracy |  | https://doi.org/10.1016/j.jiph.2021.06.002 | [6 empty] | Oropharyngeal swabs (OP)
Yati | Sherli | 17 | Anastasiou O. | Fast Detection of SARS-CoV-2 RNA Directly from Respiratory Samples Using a Loop-Mediated Isothermal Amplification (LAMP) Test | 2021-04-01 | https://doi.org/10.3390/v13050801 | Ancestral or Alpha | Germany | COVID-19 testing site/Screening location | Not given | Not given | 140 | Nasopharyngeal swabs (NP) | 54 dry NPS samples then stirred in pbs 140 NPS swabs in VTM | RealStar SARS-CoV-2 RT-PCR kit (Altona) |  |  | "These results indicate that without NA extraction, the LAMP assay cannot compete with the combination of RT-PCR and NA extraction in terms of analytical sensitivity. However, it can reliably identify individuals with a high viral load, which is relevant when aiming to reduce transmission in a community setting." |  |  | Added to the contact database - 05-16-2023
Yati | Sherli | 17 | Anastasiou O. | Fast Detection of SARS-CoV-2 RNA Directly from Respiratory Samples Using a Loop-Mediated Isothermal Amplification (LAMP) Test |  | https://doi.org/10.3390/v13050801 | [8 empty] | Custom/In-house SARS-2 primers
Karen | Sherli | 20 | R. Amer R. S. M. G. O. E.-D. N. A. A. A. A. S. W. A. A. W. M. A. | Diagnostic performance of rapid antigen test for COVID-19 and the effect of viral load, sampling time, subject's clinical and laboratory parameters on test accuracy (preprint) | 2021-03-01 |  |  | Egypt | College/University Campus (Medical Center/Hospital) | 6/2020-10/2020 | Patients in isolation units who had COVID symptoms or had been in contact with someone with COVID and had symptoms. Median age 55.5+-18.4; 59% (49) Males | 83 | Nasopharyngeal swabs (NP) |  | Genesig Real-time PCR Coronavirus assay/Z-Path-COVID-19-CE (Primerdesign) |  | For positive samples, SARS-CoV-2 RNA content was categorized according to the Ct values into high (Ct < 29), moderate (Ct = 29-36) and low (Ct = 37-39). | 69 positive, 14 negative
Karen | Sherli | 20 | R. Amer R. S. M. G. O. E.-D. N. A. A. A. A. S. W. A. A. W. M. A. | Diagnostic performance of rapid antigen test for COVID-19 and the effect of viral load, sampling time, subject's clinical and laboratory parameters on test accuracy (preprint) |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Oropharyngeal swabs (OP)
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital | 2022-04-01 | 10.3389/fpubh.2022.840984 |  | Greece | Hospital - Outpatient | December 2020 - March 2021 | Urban, 56% male, median age 50 | 263 | Oropharyngeal swabs (OP) |  | VIASURE (CerTest) |  |  | NP more reliable than OP; Variants may change these results; Low viral loads can also disrupt the results. |  | Predominant variant B.1.1.7 | Spyros Pournaras; spournaras@med.uoa.gr
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital | 2022-04-01 |  |  |  | Hospital - Inpatient |  |  |  | Nasopharyngeal swabs (NP)
Ellyn | Sherli | 23 | Babu Giridhara Mohammed | The burden of active infection and anti-SARS-CoV-2 IgG antibodies in the general population: Results from a statewide sentinel-based population survey in Karnataka, India | 2021-07-01 | https://doi.org/10.1016/j.ijid.2021.05.043 | Ancestral | India | Hospital - Outpatient | Sept 3-16, 2020 | Urban population center; healthcare workers and state agencies | 15624 | Nasopharyngeal swabs (NP) |  | Specific Test Not Described
Ellyn | Sherli | 23 | Babu Giridhara Mohammed | The burden of active infection and anti-SARS-CoV-2 IgG antibodies in the general population: Results from a statewide sentinel-based population survey in Karnataka, India |  | https://doi.org/10.1016/j.ijid.2021.05.043 |  |  | Hospital - Inpatient |  |  |  | Oropharyngeal swabs (OP)
Ellyn | Sherli | 23 | Babu Giridhara Mohammed | The burden of active infection and anti-SARS-CoV-2 IgG antibodies in the general population: Results from a statewide sentinel-based population survey in Karnataka, India |  | https://doi.org/10.1016/j.ijid.2021.05.043 |  |  | Primary Care Location
Ellyn | Sherli | 23 | Babu Giridhara Mohammed | The burden of active infection and anti-SARS-CoV-2 IgG antibodies in the general population: Results from a statewide sentinel-based population survey in Karnataka, India |  | https://doi.org/10.1016/j.ijid.2021.05.043 |  |  | Public area (not a designated screening location)
Yati | Sherli | 23 | C. M. Bachman; B. D. Grant; C. E. Anderson; L. F. Alonzo; S. Garing; S. A. Byrnes; R. Rivera; S. Burkot; A. Ball; J. W. Stafford; W. Wang; D. Banik; M. D. Keller; D. M. Cate; K. P. Nichols; B. H. Weigl; P. Dewan | Clinical validation of an open-access SARS-COV-2 antigen detection lateral flow assay, compared to commercially available assays | 2021-08-01 | https://doi.org/10.1371/journal. |  | USA | COVID-19 testing site/Screening location | December 2020 (one month) | symptomatic adult patients (age >18 years) with an onset of symp- toms within 7 days. | 170 | Nasopharyngeal swabs (NP) |  | CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel |  | The extracted RNA is amplified using a multiplexed version of the CDC 2019-nCoV Real-Time RT-qPCR targets (three targets: N1, N2, and RP) with the QuantaBio qScript XLT 1-step mix on a BioRad CFX96 Real-Time PCR Detection System | RT-qPCR. Results indicate a positive predictive agreement with NP sampling as 69% (60% -78%) OA-LFA, 74% (64% - 82%) Sofia®, and 82% (73% - 88%) BinaxNOWTM. Overall and in this patient population, amongst the rapid tests, BinaxNOWTM showed the highest sensitivity, followed by Sofia1 and OA-LFA. All three rapid tests met the WHO TPP acceptable sensitivity target of 80% relative to the AN PCR. |  | antigen levels were measured utilizing an electrochemiluminescence immunoassay on the Meso Scale Discovery (MSD) MESO Quick-Plex SQ 120 (MSD, Rockville, Maryland). An Android application (“reader app”) was developed to capture images of the completed LFA tests and to determine whether control and/or test lines could be detected in the images. | Not for quantitative data. Christine.bachman@ghlabs.org
Yati | Sherli | 23 | C. M. Bachman; B. D. Grant; C. E. Anderson; L. F. Alonzo; S. Garing; S. A. Byrnes; R. Rivera; S. Burkot; A. Ball; J. W. Stafford; W. Wang; D. Banik; M. D. Keller; D. M. Cate; K. P. Nichols; B. H. Weigl; P. Dewan | Clinical validation of an open-access SARS-COV-2 antigen detection lateral flow assay, compared to commercially available assays |  | https://doi.org/10.1371/journal. | [6 empty] | Nasal Swabs (NS)
Terry | Sherli | 28 | Basso D. | Salivary SARS-CoV-2 antigen rapid detection: A prospective cohort study | 2021-02-01 | https://doi.org/10.1016/j.cca.2021.02.014 | Ancestral | Italy | Hospital - Inpatient | 1 Aug -30 Nov, 2020 | 99 Female, 135 male; | 234 total - 138 inpatient; 96 Outpatient | Saliva (Sal) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher)
Terry | Sherli | 28 | Basso D. | Salivary SARS-CoV-2 antigen rapid detection: A prospective cohort study |  | https://doi.org/10.1016/j.cca.2021.02.014 |  |  | Hospital - Outpatient |  |  |  | Nasopharyngeal swabs (NP)
Terry-finished | Sherli | 31 | K. A. Bond; B. Smith; E. Gardiner; K. C. Liew; E. Williams; N. Walsham; M. Putland; D. A. Williamson | Utility of SARS-CoV-2 rapid antigen testing for patient triage in the emergency department: A clinical implementation study in Melbourne, Australia | 2022-05-01 |  |  | Melbourne, Australia | Hospital - Inpatient | 8/23/2021 - 12/20/2021 | Not specified. | 1762 paired samples | Other | Not specified. | Abbott RealTime SARS-CoV-2 (Abbott) |  |  | "When compared to RT-PCR, sensitivity for the RAT was 75.5% (95%CI: 69.9-80.4%), with decreasing sensitivity associated with lower assigned risk categories" | 11 paired RT-PCR and RAT were excluded from subse- quent analyses (six had invalid RAT results, and five RAT results were indeterminate) | Exact specimen type was not specified. Need to contact authors. | *Corresponding author at: Royal Melbourne Hospital at The Peter Doherty Institute for Infection and Immunity, 792 Eliza- beth Street, Melbourne, Australia, 3100. E-mail address: katherine.bond@mh.org.au (K.A. Bond).
Terry-finished | Sherli | 31 | K. A. Bond; B. Smith; E. Gardiner; K. C. Liew; E. Williams; N. Walsham; M. Putland; D. A. Williamson | Utility of SARS-CoV-2 rapid antigen testing for patient triage in the emergency department: A clinical implementation study in Melbourne, Australia |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Terry-finished | Sherli | 31 | K. A. Bond; B. Smith; E. Gardiner; K. C. Liew; E. Williams; N. Walsham; M. Putland; D. A. Williamson | Utility of SARS-CoV-2 rapid antigen testing for patient triage in the emergency department: A clinical implementation study in Melbourne, Australia |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Panther Fusion or Aptima SARS-CoV-2 assay (Hologic)
Terry-finished | Sherli | 31 | K. A. Bond; B. Smith; E. Gardiner; K. C. Liew; E. Williams; N. Walsham; M. Putland; D. A. Williamson | Utility of SARS-CoV-2 rapid antigen testing for patient triage in the emergency department: A clinical implementation study in Melbourne, Australia |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Alinity m SARS-CoV-2 AMP (Abbott)
Ellyn | Sherli | 33 | M. Borro; G. Salerno; A. Montori; A. Petrucca; P. Anibaldi; A. Marcolongo; R. Bonfini; M. Simmaco; I. Santino | SARS-CoV-2 Transmission Control Measures in the Emergency Department: The Role of Rapid Antigenic Testing in Asymptomatic Subjects | 2022-04-01 |  |  | Italy | Emergency Dept/Room | December 2021 - January 2022 | Asymptomatic subjects attending emergency departments in non-COVID-19 areas | 300 | Nasopharyngeal swabs (NP) |  | FTD SARS-CoV-2 Assay (Fast Track Diagnostics, Luxembourg) |  | viral N gene and ORF1ab region | strategic application in specific settings may limit transmission among patients
Ellyn | Sherli | 33 | M. Borro; G. Salerno; A. Montori; A. Petrucca; P. Anibaldi; A. Marcolongo; R. Bonfini; M. Simmaco; I. Santino | SARS-CoV-2 Transmission Control Measures in the Emergency Department: The Role of Rapid Antigenic Testing in Asymptomatic Subjects |  | https://doi.org/10.1093/labmed/lmaa104 |  |  |  | 23 January 2022 - 3 March 2022 | Any patient admitted to ED | 899 | Nasopharyngeal swabs (NP)
Amy | Sherli | 34 | E. Burdino; F. Cerutti; F. Panero; T. Allice; G. Gregori; M. G. Milia; G. Cavalot; A. Altavilla; F. Aprà; V. Ghisetti | SARS-CoV-2 microfluidic antigen point-of-care testing in Emergency Room patients during COVID-19 pandemic | 2021-10-01 |  |  | Italy | Emergency Dept/Room | March 1 - March 31, 2021 | Urban | 1232 | Nasopharyngeal swabs (NP) | For RT-PCR | Simplexa (DiaSorin) |  | 42.3% | New generation microfluidics based RATs are better than previous RATs and would be a good choice for screening without losing spec/senc |  | References several other papers that looked at the LumiraDx test - need to look into these to check that we found them in our screen. | E-mail address: elisa.burdino@aslcittaditorino.it (E. Burdino).
Amy | Sherli | 34 | E. Burdino; F. Cerutti; F. Panero; T. Allice; G. Gregori; M. G. Milia; G. Cavalot; A. Altavilla; F. Aprà; V. Ghisetti | SARS-CoV-2 microfluidic antigen point-of-care testing in Emergency Room patients during COVID-19 pandemic |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Nasal Swabs (NS) | For RAT | Panther Fusion or Aptima SARS-CoV-2 assay (Hologic) |  | 20.6%
Amy | Sherli | 34 | E. Burdino; F. Cerutti; F. Panero; T. Allice; G. Gregori; M. G. Milia; G. Cavalot; A. Altavilla; F. Aprà; V. Ghisetti | SARS-CoV-2 microfluidic antigen point-of-care testing in Emergency Room patients during COVID-19 pandemic |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid) |  | 15%
Amy | Sherli | 34 | E. Burdino; F. Cerutti; F. Panero; T. Allice; G. Gregori; M. G. Milia; G. Cavalot; A. Altavilla; F. Aprà; V. Ghisetti | SARS-CoV-2 microfluidic antigen point-of-care testing in Emergency Room patients during COVID-19 pandemic |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  | 13.4%
Amy | Sherli | 34 | E. Burdino; F. Cerutti; F. Panero; T. Allice; G. Gregori; M. G. Milia; G. Cavalot; A. Altavilla; F. Aprà; V. Ghisetti | SARS-CoV-2 microfluidic antigen point-of-care testing in Emergency Room patients during COVID-19 pandemic |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Cobas Kits/Systems (Roche) |  | 7%
Amy | Sherli | 34 | E. Burdino; F. Cerutti; F. Panero; T. Allice; G. Gregori; M. G. Milia; G. Cavalot; A. Altavilla; F. Aprà; V. Ghisetti | SARS-CoV-2 microfluidic antigen point-of-care testing in Emergency Room patients during COVID-19 pandemic |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Allplex assays (Seegene) |  | 1.3%
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-01 |  |  | Uganda | Hospital - Inpatient | August 2020 - March 2021 | Isolation centers (regional and national), travellers entering the country; mean age 36; 66.8% males | 1533 | Nasopharyngeal swabs (NP) |  | LightMix SarbecoV (TIB Microbiol) |  |  |  |  | TP, FP, TN, and FN came from the supplemental information on the website. | Josephine Bwogi josephinebwn12@gmail.com
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda |  | https://doi.org/10.1093/labmed/lmaa104 |  |  | Not described/unclear
Amy | Sherli | 36 | Blairon L. | Implementation of rapid SARS-CoV-2 antigenic testing in a laboratory without access to molecular methods: Experiences of a general hospital | 2020-08-01 | https://doi.org/10.1016/j.jcv.2020.104472 | Ancestral | Belgium | Hospital - Inpatient | April 5, 2020 - May 4, 2020 | Not given | 774 | Nasopharyngeal swabs (NP) |  | Custom/In-house SARS-2 primers |  | E-gene SARS-CoV-2 |  |  |  | Was in the contact database as of 12/8/2022
Chathu | Sherli | 37 | N. G. Cassuto; A. Gravier; M. Colin; A. Theillay; D. Pires-Roteira; S. Pallay; R. Serreau; L. Hocqueloux; T. Prazuck | Evaluation of a SARS-CoV-2 antigen-detecting rapid diagnostic test as a self-test: Diagnostic performance and usability | 2021-07-01 |  |  | France | College/University Campus (Medical Center/Hospital) | 2021-03-01 | adults | 239 | Nasopharyngeal swabs (NP) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher)
Ellyn Finished | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-01 |  |  | Italy | Hospital - Inpatient | 19 January - 30 May 2021 | mean age 61 +/- 18 62% male | 282 | Nasal Swabs (NS) |  | Custom/In-house SARS-2 primers |  |  | "Our study largely meets the sensitivity requirement, whereas the spevificity was slightly lower than required." |  | The FN, FP, TN, TP numbers were obtained from Figure 1 and from Table 1 | annamaria.cattelan@aopd.veneto.it
Ellyn Finished | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Nasopharyngeal swabs (NP)
Amy | Sherli | 40 | Boum Y. | Performance and operational feasibility of antigen and antibody rapid diagnostic tests for COVID-19 in symptomatic and asymptomatic patients in Cameroon: a clinical, prospective, diagnostic accuracy study | 2021-03-01 | https://doi.org/10.1016/S1473-3099(21)00132-8 | Ancestral | Cameroon | Hospital - Inpatient | June 2, 2020 - Aug 30, 3030 | 21+ years old (median age 38 years IQR 30-50); 46% female | 1195 | Nasopharyngeal swabs (NP) |  | DAAN Gene RT-PCR COVID-19 (DaAnGene)
Amy | Sherli | 40 | Boum Y. | Performance and operational feasibility of antigen and antibody rapid diagnostic tests for COVID-19 in symptomatic and asymptomatic patients in Cameroon: a clinical, prospective, diagnostic accuracy study |  | https://doi.org/10.1016/S1473-3099(21)00132-8 | [6 empty] | Blood (Bld) |  | Abbott RealTime SARS-CoV-2 (Abbott)
Ellyn finished | Sherli | 43 | C. C. Cheng; C. C. Liu; T. F. Chiu; S. Shiou-Sheng Chen | Evaluation of a Rapid Antigen Test for the Diagnosis of SARS-CoV-2 during the COVID-19 Pandemic | 2022-05-01 |  |  | Taiwan | Hospital - Inpatient | May 28 - June 6, 2021 | Anyone with covid symptoms or close contact with an affected person | 6368 |  |  | LabTurbo AIO COVID-19 RNA Testing Kit |  |  | RAD "may be feasible for differentiating individuals infected with SARS-CoV-2 from non-contagious individuals |  | Contact for Abbott Test used | Chih-Chien Cheng, Department of Education and Research, Taipei City Hospital No. 10, Sec 4, RenAi Rd, Daan Dist, Taipei City, Taipei 10619, Taiwan. DXO90@tpech.gov.tw
Sabrina | Sherli | 45 | Bulilete O. | Panbio™ rapid antigen test for SARS-CoV-2 has acceptable accuracy in symptomatic patients in primary health care | 2021-02-01 | https://doi.org/10.1016/j.jinf.2021.02.014 | Ancestral | Spain | Primary Care Location | October 2-25 2020 | 18+ (not previously diagnosed with COVID) | 1369 | Nasopharyngeal swabs (NP) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "We found that test sensitivity was higher in samples that had high viral loads (Ct < 25), even in patients who were asymptomatic at the moment of testing (86.2%, 95% CI: 68.3%, 96.1%)."
Chathu | Sherli | 46 | S. I. A. Choudhary | Validation of rapid SARS-COV-2 antigen detection test as a screening tool for detection of Covid-19 infection at district hospital in northern India | 2022-04-01 |  |  | India | College/University Campus (Medical Center/Hospital) | April 2021 to November 2021 | 452 male participants and 304 female particpants | 756 | Nasopharyngeal swabs (NP) |  | Meril COVID-19 One-Step RT-PCR Kit |  |  | The rapid antigen detection showed comparable sensitivity and specificity with RT-PCR assay. The results support that RAT is an accurate alternative to RT-PCR in areas where there is an increased testing burden etc. |  |  | Dr. Aishiya Ishrat, Lecturer, Government Medical College, Jammu Email-ishrataishiya@gmail.com Phone number-+91-7889686839
Sabrina | Sherli | 49 | Caruana Giorgia | Implementing SARS-CoV-2 Rapid Antigen Testing in the Emergency Ward of a Swiss University Hospital: The INCREASE Study | 2021-04-01 | https://doi.org/10.3390/microorganisms9040798 | Ancestral | Switzerland | College/University Campus (Medical Center/Hospital) | November 7 - December 16 2020 | Presented at ED | 572 | Nasopharyngeal swabs (NP) |  | VIASURE (CerTest) |  |  | "Given the huge gap in the sensitivity between RAT and RT-PCR(roughly corresponding to about 10,000-fold reduced analystical sensitivity), RAT could be cautiously used, especially as complementaary to PCR-based screening, in nonvulnerable outpatient subjects with one to four days of symptoms, when the viral load is likely very high."
Sabrina | Sherli | 49 | Caruana Giorgia | Implementing SARS-CoV-2 Rapid Antigen Testing in the Emergency Ward of a Swiss University Hospital: The INCREASE Study |  | https://doi.org/10.3390/microorganisms9040798 |  |  | Emergency Dept/Room | [5 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Terry | Sherli | 50 | Caruana G. | The dark side of SARS-CoV-2 rapid antigen testing: screening asymptomatic patients | 2021-05-01 | https://doi.org/10.1016/j.nmni.2021.100899 | Ancestral and Alpha | Switzerland | College/University Campus (Medical Center/Hospital) | 4/12/2020-4/1/2021 | 63 females, 53 males; median age 46.7 years | 116 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "...adopting an RAT-based screening strategy in patients without symptoms of COVID-19 might miss a significant number of asymptomatic SARS-CoV-2 carriers."
Terry | Sherli | 50 | Caruana G. | The dark side of SARS-CoV-2 rapid antigen testing: screening asymptomatic patients |  | https://doi.org/10.1016/j.nmni.2021.100899 | [8 empty] | Custom/In-house SARS-2 primers |  | E gene and RdRp gene
Brandi | Sherli | 51 | E. Cottone; F. Van Hoecke; G. A. Martens; E. De Laere; R. De Smedt; S. Vervaeke; M. Vanhee; D. De Smet | Pitfalls of SARS-CoV-2 antigen testing at emergency department | 2022-05-01 |  |  | Belgium | Emergency Dept/Room | November 1st to December 2nd 2020 | 250 males and 214 females, age range 47 to 81 years old | 464 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  |  | Main results showed inferior performance of Roche RAT for detection of SARS-CoV-2 infection with a sensitivity of 45.5%. |  |  | Eleonora Cottone, emal: eleonora.cottone@hotmail.com
Yati | Sherli | 53 | Cento V. | Frontline Screening for SARS-CoV-2 Infection at Emergency Department Admission by Third Generation Rapid Antigen Test: Can We Spare RT-qPCR? | 2021-05-01 | https://doi.org/10.3390/v13050818 | Ancestral | Italy | Emergency Dept/Room | 21 October 2020 and 9 December 2020 | urban; 57.8% male; median age 66 | 960 | Nasopharyngeal swabs (NP) |  | Simplexa (DiaSorin) |  |  | "Our study also demonstrated that the LumiraDx antigen test may fail to detect SARS-CoV-2 in a number of samples with high levels of nasopharyngeal viral-shedding and evidence of replicative intermediates."
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-01 | https://doi.org/10.1016/j.cmi.2022.02.014 |  | Spain | Primary Care Location | Feb 3 - April 14 2021 | median age 44; 63% female | 900-600 symptomatic | Nasopharyngeal swabs (NP) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "We conclude there are considerable test-to-test antigen detection variations among patients with clinical suspicion of COVID-19 and close asymptomatic contacts. Such differences should be taken into account when choosing tests for SARS-CoV-2 screening in patients attending primary care." |  |  | ** Corresponding author: Patricia Munoz, E-mail addresses: pmunoz@hggm.es (P. Munoz)
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-01 | https://doi.org/10.1016/j.cmi.2022.02.014 |  |  |  |  | median age 47; 53.7% female | 300 contacts
Chathu | Sherli | 54 | B. F. Escrivá; M. D. O. Mochón; R. M. González; C. S. García; A. T. Pla; A. S. Ricart; M. M. García; I. T. Aranda; F. G. García; C. G. Cardona | The effectiveness of rapid antigen test-based for SARS-CoV-2 detection in nursing homes in Valencia, Spain"" | 2021-08-01 |  |  | Spain | Long-term Care Facility (Nursing home, rehab centers) | November 2020 to December 2020 | Participants were either residents or caretakers | 448 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) | [5 empty] | Begona ̃ Fuster Escriva Email - begona.fuster@gmail.com
Terry | Sherli | 54 | Cerutti F. | Urgent need of rapid tests for SARS CoV-2 antigen detection: Evaluation of the SD-Biosensor antigen test for SARS-CoV-2 | Sept 2020 | https://doi.org/10.1016/j.jcv.2020.104654 | Ancestral | Italy | Emergency Dept/Room | March 3 - May 1, 2020 | Random patients in EDs | 185 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene)
Terry | Sherli | 54 | Cerutti F. | Urgent need of rapid tests for SARS CoV-2 antigen detection: Evaluation of the SD-Biosensor antigen test for SARS-CoV-2 |  | https://doi.org/10.1016/j.jcv.2020.104654 |  |  | COVID-19 testing site/Screening location | 2020-08-01 | Travelers returning from high risk countries | 145 |  |  | Simplexa (DiaSorin) |  |  | "The adoption of POCT for SARS CoV-2 testing is certainly more suitable in point of care centers for mass screening where the prevalence of COVID-19 is much lower and the pre-test probability of not having the disease is higher than that in the patients admitted to the emergency room were the pre-test probability of having COVID-19 is significantly higher and false negative results are relevant for the correct management of patients."
Terry | Sherli | 54 | Cerutti F. | Urgent need of rapid tests for SARS CoV-2 antigen detection: Evaluation of the SD-Biosensor antigen test for SARS-CoV-2 |  | https://doi.org/10.1016/j.jcv.2020.104654 | [8 empty] | Cobas Kits/Systems (Roche)
Yati | Sherli | 55 | Chaimayo C. | Rapid SARS-CoV-2 antigen detection assay in comparison with real-time RT-PCR assay for laboratory diagnosis of COVID-19 in Thailand | 2020-11-01 | https://doi.org/10.1186/s12985-020-01452-5 | Ancestral | Thailand | Hospital - Inpatient | March to May 2020 | Not given | 454 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  |  | "Even with its limitations, the rapid SARS-CoV-2 antigen test can beneft all healthcare workers in managing infected individuals in time efectively, especially in rural and outbreak areas."
Yati | Sherli | 55 | Chaimayo C. | Rapid SARS-CoV-2 antigen detection assay in comparison with real-time RT-PCR assay for laboratory diagnosis of COVID-19 in Thailand |  | https://doi.org/10.1186/s12985-020-01452-5 |  |  | Hospital - Outpatient |  |  |  | Throat Swabs (TS)
Terry - finished | Sherli | 55 | K. S. Faíco-Filho; F. E. F. Júnior; L. V. L. Moreira; P. R. G. Lins; A. F. O. Justo; N. Bellei | Evaluation of the Panbio™ COVID-19 ag rapid test at an emergency room in a hospital in São Paulo, Brazil | 2022-03-01 |  |  | Brazil | Emergency Dept/Room | Exact dates of data collection, not provided | Patients greater than or equal to. 18 years of age. Median age: 63 years. Age range 22-69. Males: 54.3%. Females: 45.7%. | 127 | Nasopharyngeal swabs (NP) |  | GeneFinder COVID-19 Plus RealAmp Kit (Osang Healthcare Co., Ltd) |  |  | "the agreement beyond chance between the PanbioTM COVID-19 Ag test and RT-PCR decreases as Ct levels increase. In particular, the agreement was 78% for samples with Ct values of ≤ 25 but decreases at higher Cts." |  | As indicated in the dates for data collection, this was not specified, specificity is also not given. Need to contact authors. | Corresponding author at: Laboratorio de Virologia, Division of -Infectious Diseases, Department of Medicine, Escola Paulista de Medicina (EPM), Rua Pedro de Toledo 781 -15 andar CEP- 04039 032, Sao Paulo, Brazil. ~ E-mail address: nbellei@uol.com.br (Nancy. Bellei).
Terry - finished | Sherli | 56 | E. Farfour; C. Amiel; E. Jolly; S. Zia-Chahabi; L. Mazaux; P. Simonin; I. Berkane; M. Cornily; T. Pascreau; V. M. Sars-Co | The Panbio COVID-19 Ag Rapid Test: Which Performances are for COVID-19 Diagnosis? | 2021-03-01 |  |  | France | Not described/unclear | Oct. 19 - Nov. 27, 2020 | Prospective Study: Age: 28 - 45.5 years old. Median age: 34.2 years. Male:female ratio 0.31. | Retrospective study: 81 participants. | Nasopharyngeal swabs (NP) |  | Alinity m SARS-CoV-2 AMP (Abbott) |  |  | "Our results show the Panbio COVID-19 Ag Rapid test device displays low performance for the detection of SARS-CoV-2 infected patients. The sensitivity of the test mainly depends on the viral load estimated by the CT value, all the samples with a CT value below 26.0 being negative with the RAD. We also show a trend in higher sensitivity when the test is performed during the 4 first days of the disease and in patients younger than 35." |  | The retrospective study involved 81 stored(frozen) specimens. No specifics given on the exact dates of collection or on the demographics of these patients. Should contact authors. | Eric Farfour, Service de Biologie Clinique Hopital Foch 40 rue Worth 92150 Suresnes France Phone: +33 6 64 98 85 61 Fax: +33 1 46 25 24 22 Email: e.farfour@hopital-foch.com ericf6598@yahoo.fr
Terry - finished | Sherli | 56 | E. Farfour; C. Amiel; E. Jolly; S. Zia-Chahabi; L. Mazaux; P. Simonin; I. Berkane; M. Cornily; T. Pascreau; V. M. Sars-Co | The Panbio COVID-19 Ag Rapid Test: Which Performances are for COVID-19 Diagnosis? |  | https://doi.org/10.1093/labmed/lmaa104 | [5 empty] | Prospective study: 330 participants. |  |  | Allplex assays (Seegene)
Terry - finished | Sherli | 58 | F. Fitoussi; S. Tonen-Wolyec; N. Awaida; R. Dupont; L. Bélec | Analytical performance of the point-of-care BIOSYNEX COVID-19 Ag BSS for the detection of SARS‐CoV‐2 nucleocapsid protein in nasopharyngeal swabs: a prospective field evaluation during the COVID-19 third wave in France | 2021-10-01 |  |  | France | COVID-19 testing site/Screening location | March-April 2021? | >/= 18 years of age. Age range: 18-95 years. Median age=34 years | 967 | Nasopharyngeal swabs (NP) |  | Novel Coronavirus (2019-nCoV) Real Time Multiplex RT-PCR kit (Liferiver) |  |  | "the sensitivity of the BIOSYNEX Ag-RDT was lower among specimens from asymptomatic persons (79.4%) than among specimens from symptomatic patients (95.0%)." |  | Need to find out exact dates of specimen collection. | Laurent Bélec laurent.belec@aphp.fr
Terry - finished | Sherli | 60 | L. Ford; M. J. Whaley; M. M. Shah; P. P. Salvatore; H. E. Segaloff; A. Delaney; D. W. Currie; L. Boyle-Estheimer; M. O'Hegarty; C. N. Morgan; J. Meece; L. Ivacic; N. J. Thornburg; A. Tamin; J. L. Harcourt; J. M. Folster; M. Medrzycki; S. Jain; P. Wong; K. Goffard; D. Gieryn; J. Kahrs; K. Langolf; T. Zochert; J. E. Tate; C. H. Hsu; H. L. Kirking | Antigen Test Performance Among Children and Adults at a SARS-CoV-2 Community Testing Site | 2021-09-01 |  |  | United States | COVID-19 testing site/Screening location | Nov. 16, 2020-Dec. 15, 2020 | Adults >/= 18 years of age. | 1885 | Nasopharyngeal swabs (NP) |  | Specific Test Not Described |  |  | "From this study and others, antigen tests had lower, although not necessarily statistically significant, sensitivity among children compared with adults" |  | This study is focused on RAT on children (under 18 years of age) and comparing those results with adults. Information presented in this spreadsheet is just data collected from the adults. | Corresponding Author: Laura Ford, PhD, COVID-19 Response Team, Centers for Disease Control and Prevention, 1600 Clifton Rd NE, Atlanta, GA 30329, USA. E-mail: qdz4@cdc.gov.
Terry - finished | Sherli | 63 | Garcia-Cardenas, Francisco; F. Alba; C. Ricardo; B. Jenny; V. Rafael; P. Fernando; F.-J. Emmanuel; C.-T. Alberto; M.-V. Alfredo; P. R.-G. Juan; H.-M. Alfredo; A. H. Luis | Analytical performances of the COVISTIX and Panbio antigen rapid tests for SARS-CoV-2 detection in an unselected population (all commers) | 2022-05-01 |  |  | Mexico | COVID-19 testing site/Screening location | 15 Dec 2021-5 Jan 2022 | Probable COVID-19 cases, symptomatic and asymptomatic. COVISTIX Assay: 391 female, 392 male. Age range (28-51), median age 40. | 783 | Nasal Swabs (NS) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "The COVISTIXTM rapid antigen test shows a high performance in all comers, thus, this test is also adequate for testing patients who have passed the peak of viral shedding or for asymptomatic patients." |  |  | Luis A. Herrera, Instituto Nacional de Medicina Genómica, Periférico Sur 4809, Arenal Tepepan, C.P. 14610, Tlalpan, CDMX, México; e-mail: lherrera@inmegen.gob.mx Alfredo Hidalgo, Instituto Nacional de Medicina Genómica, Periférico Sur 4809, Arenal Tepepan, C.P. 14610, Tlalpan, CDMX, México; e-mail: ahidalgo@inmegen.gob.mx
Terry - finished | Sherli | 63 | Garcia-Cardenas, Francisco; F. Alba; C. Ricardo; B. Jenny; V. Rafael; P. Fernando; F.-J. Emmanuel; C.-T. Alberto; M.-V. Alfredo; P. R.-G. Juan; H.-M. Alfredo; A. H. Luis | Analytical performances of the COVISTIX and Panbio antigen rapid tests for SARS-CoV-2 detection in an unselected population (all commers) |  | https://doi.org/10.1093/labmed/lmaa104 |  |  | College/University Campus (Medical Center/Hospital) | May. 1 - Aug. 16, 2021 | Panbio Assay: 979 female, 1223 male. Age range (30- 50), median age 41. | 2202 | Nasopharyngeal swabs (NP)
Terry - finished | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil | 2022-06-01 |  |  | Brazil | Hospital - Inpatient | March 22-April21, 2021 | Nonvaccinated. Median age is 56.3yrs (46-65yrs). 53.7% female. | 162 | Nasopharyngeal swabs (NP) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "Overall, slight or fair agreement was observed between Ag-RDTs and RT–PCR, except for the Ag-RDT COVID-19 (Acro Biotech), in which moderate agreement was observed." |  |  | marianalfreire@hotmail.com
Terry - finished | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Nasal Swabs (NS)
Terry - finished | Sherli | 68 | R. M. Galliez; L. Bomfim; D. Mariani; I. C. Leitão; A. C. P. Castiñeiras; C. C. A. Gonçalves; B. Ortiz da Silva; P. H. Cardoso; M. B. Arruda; P. Alvarez; R. Brindeiro; V. A. Ota; D. G. M. Rodrigues; L. J. da Costa; O. D. C. Ferreira, Jr.; T. Castiñeiras; D. S. Faffe; A. Tanuri | Evaluation of the Panbio COVID-19 Antigen Rapid Diagnostic Test in Subjects Infected with Omicron Using Different Specimens | 2022-06-01 |  |  | Brazil | College/University Campus (Medical Center/Hospital) | Jan. 17 - Feb. 7, 2022 | Age: average is 39 years (21-74). Female: 65.6%, Male: 34.4%. | 192 | Nasal Swabs (NS) |  | IDT SARS-CoV-2 (2019-nCoV) multiplex CDC qPCR probe Assay (Integrated DNA Technologies) |  | Targeting the SARS-CoV-2 N1 and N2 genes and the human RNase P (RNaseP) gene | "results from the present study demonstrated that Panbio COVID-19 Ag-RDT performance is unaffected in vaccinated individuals infected with the Omicron VOC. Furthermore, this study supported nasal specimens as the optimal specimen type for Panbio COVID-19 Ag-RDT performance compared with oral specimens." |  |  | Address correspondence to Amilcar Tanuri, atanuri@gmail.com.
Terry - finished | Sherli | 68 | R. M. Galliez; L. Bomfim; D. Mariani; I. C. Leitão; A. C. P. Castiñeiras; C. C. A. Gonçalves; B. Ortiz da Silva; P. H. Cardoso; M. B. Arruda; P. Alvarez; R. Brindeiro; V. A. Ota; D. G. M. Rodrigues; L. J. da Costa; O. D. C. Ferreira, Jr.; T. Castiñeiras; D. S. Faffe; A. Tanuri | Evaluation of the Panbio COVID-19 Antigen Rapid Diagnostic Test in Subjects Infected with Omicron Using Different Specimens |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Nasopharyngeal swabs (NP)
Terry - finished | Sherli | 68 | R. M. Galliez; L. Bomfim; D. Mariani; I. C. Leitão; A. C. P. Castiñeiras; C. C. A. Gonçalves; B. Ortiz da Silva; P. H. Cardoso; M. B. Arruda; P. Alvarez; R. Brindeiro; V. A. Ota; D. G. M. Rodrigues; L. J. da Costa; O. D. C. Ferreira, Jr.; T. Castiñeiras; D. S. Faffe; A. Tanuri | Evaluation of the Panbio COVID-19 Antigen Rapid Diagnostic Test in Subjects Infected with Omicron Using Different Specimens |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Saliva (Sal)
Terry - finished | Sherli | 68 | R. M. Galliez; L. Bomfim; D. Mariani; I. C. Leitão; A. C. P. Castiñeiras; C. C. A. Gonçalves; B. Ortiz da Silva; P. H. Cardoso; M. B. Arruda; P. Alvarez; R. Brindeiro; V. A. Ota; D. G. M. Rodrigues; L. J. da Costa; O. D. C. Ferreira, Jr.; T. Castiñeiras; D. S. Faffe; A. Tanuri | Evaluation of the Panbio COVID-19 Antigen Rapid Diagnostic Test in Subjects Infected with Omicron Using Different Specimens |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Other | Oral
Ellyn | Sherli | 69 | F. Garcia-Cardenas; F. Peñaloza; J. Bertin-Montoya; R. Valdéz-Vázquez; A. Franco; R. Cortés; E. Frias-Jimenez; A. Cedro-Tanda; A. Mendoza-Vargas; J. P. Reyes-Grajeda; A. Hidalgo-Miranda; L. A. Herrera | Analytical Performances of the COVISTIX™ Antigen Rapid Test for SARS-CoV-2 Detection in an Unselected Population (All-Comers) | 2022-05-01 |  |  | Mexico | Not described/unclear | 1st May-16 August 2021 | 391 Female, 392 Male Median age 40, IQR 28-51 | 783 | Nasopharyngeal swabs (NP) |  | Specific Test Not Described |  |  |  |  | Omicron variant found using Illumina sequencing. Reference test not specified Breakdown for Ct value TN and FP | Alfredo Hidalgo-Miranda ahidalgo@inmegen.gob.mx Luis A. Herrera lherrera@inmegen.gob.mx
Ellyn | Sherli | 69 | F. Garcia-Cardenas; F. Peñaloza; J. Bertin-Montoya; R. Valdéz-Vázquez; A. Franco; R. Cortés; E. Frias-Jimenez; A. Cedro-Tanda; A. Mendoza-Vargas; J. P. Reyes-Grajeda; A. Hidalgo-Miranda; L. A. Herrera | Analytical Performances of the COVISTIX™ Antigen Rapid Test for SARS-CoV-2 Detection in an Unselected Population (All-Comers) |  |  |  |  | Not described/unclear | 15 Dec 2021-5 Jan 2022 | No characteristics for second cohort, Omicron variant | 999 | Nasal Swabs (NS)
Terry - finished | Sherli | 71 | S. García-Fernández; D. Pablo-Marcos; S. V. de la Fuente; M. J. R. Rodríguez; M. Gozalo; J. Rodríguez-Lozano; J. M. Méndez-Legaza; J. Calvo | Evaluation of the rapid antigen detection test STANDARD F COVID-19 Ag FIA for diagnosing SARS-CoV-2: experience from an Emergency Department | 2022-03-01 |  |  | Spain | Emergency Dept/Room | Oct. 14 - Nov. 18, 2020 | Demographic data was not collected. | 663 | Nasopharyngeal swabs (NP) |  | VIASURE (CerTest) |  |  | "In this evaluation the antigen test STANDARD F COVID-19 Ag FIA presented an excellent clinical performance (84.0% sensitivity and 99.6% specificity) and an almost perfect agreement compared with the reference method, the RT-PCR, when used in patients with signs or symptoms compatible with COVID-19 that had started in the last 5 days." |  |  | * Corresponding author. Servicio de Microbiología. Hospital Universitario Marques de Valdecilla. Av. Valdecilla, 25, 39010 Santander. Spain. Tel.: +34679130556. E-mail address: segarciafe@hotmail.com (S. García-Fern andez).
Chathu finished | Sherli | 73 | S. González-Fiallo; I. Mena-Rodríguez; P. Castro-Batista; V. M. Doeste-Hernández; V. Louit-Laborit | Evaluation of SARS-CoV-2 Rapid Antigen Tests in Use on the Isle of Youth, Cuba | 2022-04-01 |  |  | Cuba | Primary Care Location | January 2021–April 2021 | Not specified | 778 | Nasopharyngeal swabs (NP) |  | RIDAGENE SARS-CoV-2 (R-Bio-pharm) |  |  | Sensitivity was high but lower than specificity
Chathu finished | Sherli | 73 | S. González-Fiallo; I. Mena-Rodríguez; P. Castro-Batista; V. M. Doeste-Hernández; V. Louit-Laborit | Evaluation of SARS-CoV-2 Rapid Antigen Tests in Use on the Isle of Youth, Cuba |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Oropharyngeal swabs (OP)
Ellyn | Sherli | 75 | Dierks S. | Diagnosing SARS-CoV-2 with Antigen Testing, Transcription-Mediated Amplification and Real-Time PCR | 2021-05-01 | https://doi.org/10.3390/jcm10112404 | Ancestral or Alpha | Germany | College/University Campus (Medical Center/Hospital) | not stated - article recieved by journal April 23, 2021 | employees at University Med Center Gottingen | 444 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "Despite the above-mentioned limitations, the presented “real world” assessment suggests superiority of the LumiraDx assay compared to the NADAL assay when applied in a low prevalence setting, based on both sensitivity and specificity considerations. Infections with low viral load, however, will be missed by both of them. In contrast, acceptable sensitivity and diagnostic accuracy similar to real-time PCR, even in the case of 1:10 pooling, could be shown for the Hologic Aptima SARS-CoV-2 TMA assay for high-throughput screenings." |  |  | Previously contacted authors
Ellyn | Sherli | 75 | Dierks S. | Diagnosing SARS-CoV-2 with Antigen Testing, Transcription-Mediated Amplification and Real-Time PCR |  | https://doi.org/10.3390/jcm10112404 | [8 empty] | Genesig Real-time PCR Coronavirus assay/Z-Path-COVID-19-CE (Primerdesign)
Ellyn | Sherli | 75 | Dierks S. | Diagnosing SARS-CoV-2 with Antigen Testing, Transcription-Mediated Amplification and Real-Time PCR |  | https://doi.org/10.3390/jcm10112404 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Chathu finished | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay | 2022-06-01 |  |  | Germany | COVID-19 testing site/Screening location | January 2022 - February 2022 |  | 447 | Nasal Swabs (NS) | First cohort | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | The within-run pre- cision analysis revealed variation of coefficients of 2.1% (252 ± 5.3 Index) and 9.7% (7.9 ± 0.8 Index). The between-run precision of the CoV2Ag was 2.3% (331 ± 7.5 Index) and 18.3% (7.3 ± 1.3 Index). | None
Chathu finished | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Chathu finished | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay |  | https://doi.org/10.1093/labmed/lmaa104 |  |  | College/University Campus (Medical Center/Hospital) | Jan 2022-March 2022 |  | 9046 | Nasal Swabs (NS) | Second cohort | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher)
Chathu finished | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Yati - done | Sherli | 86 | M. Ifko; Ž. Tkalčić Švabek; I. Friščić; M. M. Kardum Paro; I. Prkačin; L. Đerek; A. Livun; M. Skvarč | Diagnostic validation of two SARS-CoV-2 immunochromatographic tests in Slovenian and Croatian hospitals | 2021-06-01 | https://doi.org/10.3325/cmj.2021.62.513 |  | Slovenia and Croatia (only the Croatian results are entered as Slovenia was a case control study | College/University Campus (Medical Center/Hospital) | September 2020 to October 2020 | patients with at least one clinical sign of COVID-19 | 51 (Croatia only) | Nasopharyngeal swabs (NP) |  | PCR Biosystems |  | This test was used only in Croatia | "ALLTEST COVID-19 Antigen Test was validated in two Croatian hospitals. The ALLTEST COVID-19 Antigen Test had the sensitivity between 81.39% and 91.11%, with a negative predictive value between 85.45% and 98.78%." |  | This study were performed in two different countries, Slovenia (2 ER departments) and Croatia (2 hospitals). However the Slovenia study was case control whereby they used another RAT NADAL COVID-19 and tested only the ones that were positive in PCR using of Seegene Allplex 2019-nCoV test (SeeGene, Seoul, South Korea). Since this is a case control study, only data from the Croatian study were entered for analysis | Correspondence to: Martina Ifko Slovenia martina.ifko@gmail.com To ask for missing data and ask which test was used in Croatia.
Yati - done | Sherli | 86 | M. Ifko; Ž. Tkalčić Švabek; I. Friščić; M. M. Kardum Paro; I. Prkačin; L. Đerek; A. Livun; M. Skvarč | Diagnostic validation of two SARS-CoV-2 immunochromatographic tests in Slovenian and Croatian hospitals |  | https://doi.org/10.1093/labmed/lmaa104 |  |  | Emergency Dept/Room | September 2020 to October 2021 | patients with at least one clinical sign of COVID-20 | Slovenia-General Hospital Jesenice, NADAL 106; ALLTEST 90 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  | This test was used in Slovenia-General Hospital Jesenice
Yati - done | Sherli | 86 | M. Ifko; Ž. Tkalčić Švabek; I. Friščić; M. M. Kardum Paro; I. Prkačin; L. Đerek; A. Livun; M. Skvarč | Diagnostic validation of two SARS-CoV-2 immunochromatographic tests in Slovenian and Croatian hospitals |  | https://doi.org/10.1093/labmed/lmaa104 |  |  | Emergency Dept/Room | September 2020 to October 2022 | patients with at least one clinical sign of COVID-21 | Slovenia-General Hospital Slovenj Gradec Nadal 92 and ALLTEST 94 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  | This test was used in Slovenia-General Hospital Slovenj Gradec
Yati | Sherli | 87 | Fernandez-Montero A. | Validation of a rapid antigen test as a screening tool for SARS-CoV-2 infection in asymptomatic populations. Sensitivity, specificity and predictive values | 2021-06-01 | https://doi.org/10.1016/j.eclinm.2021.100954 | Ancestral | Spain | College/University Campus (non-medical) | November 2020 - January 2021 | Asymptomatic; university students and employees | 2543 | Nasopharyngeal swabs (NP) |  | BGI 2019-nCoV Real-time Fluorescent RT-PCR kit (BGI Genomics) |  |  | "the Roche SARS-CoV-2 Rapid Antigen Test showed a moderate sensitivity on asymptomatic patients, when compared with RT-PCR. Test sensitivity and specificity increased in patients with higher viral loads or higher COVID-19 risk (presence of mild symptoms or known exposure to SARS-CoV-2) but it was lower in asymptomatic low risk patients,"
Yati | Sherli | 87 | Fernandez-Montero A. | Validation of a rapid antigen test as a screening tool for SARS-CoV-2 infection in asymptomatic populations. Sensitivity, specificity and predictive values |  | https://doi.org/10.1016/j.eclinm.2021.100954 | [6 empty] | Oropharyngeal swabs (OP)
Ellyn | Sherli | 89 | Ferté T. | Accuracy of COVID-19 rapid antigenic tests compared to RT-PCR in a student population: The StudyCov study | 2021-06-01 | https://doi.org/10.1016/j.jcv.2021.104878 | Ancestral | France | COVID-19 testing site/Screening location | 28 October - 20 November, 2020 | mean age 22.95; 34% male Urban | 692 | Nasopharyngeal swabs (NP) |  | ARGENE SARS-CoV-2 R-Gene (Biomerieux) |  |  | This study highlights that AT lack sensitivity compared to RT-PCR in the context of SARS-CoV-2 diagnosis and screening. For the asymptomatic sub-group, two thirds of SARS-CoV-2 cases would be considered negative by AT even among the likely contagious ones. However, there was no false-positive in this study thus making AT reliable when positive."
Terry - finished | Sherli | 92 | K. K. Jakobsen; J. S. Jensen; T. Todsen; N. Kirkby; F. Lippert; A. M. Vangsted; M. Klokker; C. von Buchwald | Accuracy of anterior nasal swab rapid antigen tests compared with RT-PCR for massive SARS-CoV-2 screening in low prevalence population | 2021-10-01 |  |  | Denmark | COVID-19 testing site/Screening location | March 2 - 22, 2021 | >/= 18 years of age. Median age = 48 years (36-57). Female: 50.7%. | 6824 | Nasopharyngeal swabs (NP) |  | Luna Universal Probe One-Step RT-PCR for Detection of COVID-19 (SignaGen Labs) |  |  | "The study was performed in a public setting and we found a low prevalence of SARS-CoV-2 of 0.9%." |  | "Patients referred to COVID-19 testing by a doctor with symptoms of COVID-19 were tested in a separate section of the test centers and were not included in this study" | Christian von Buchwald, Department of Otorhinolaryngology, Head and Neck Surgery and Audiology, Rigshospitalet, University of Copenhagen, Inge Lehmanns Vej 8, 2100 Copenhagen, Denmark. Email: Christian.von.Buchwald@regionh.dk
Terry - finished | Sherli | 92 | K. K. Jakobsen; J. S. Jensen; T. Todsen; N. Kirkby; F. Lippert; A. M. Vangsted; M. Klokker; C. von Buchwald | Accuracy of anterior nasal swab rapid antigen tests compared with RT-PCR for massive SARS-CoV-2 screening in low prevalence population |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Nasal Swabs (NS)
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection | 2022-03-01 |  | Ancestral | Sri Lanka | Public area (not a designated screening location) | October 2020 - mid-January 2021 | People living in Colombo Municipality Council (highly dense urban) | 8470 | Blood (Bld) |  | Allplex assays (Seegene) |  |  |  |  | Calculated the combined Reference kit true positive and false negative based on the data in the tables. See back calculations at the end of the sheet. | Gathsaurie Neelika Malavige: gathsaurie.malavige@ndm.ox.ac.uk
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection |  | https://doi.org/10.1093/labmed/lmaa104 |  |  | COVID-19 testing site/Screening location |  |  |  | Nasopharyngeal swabs (NP) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher)
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Standard M nCoV Real-Time Detection Kit (SD Biosenor)
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | COVID-19 Real-time PCR kit (HBRT-COVID-19) (Chaozhou Hybribio Biochemistry Ltd., China)
Terry | Sherli | 93 | García-Fiñana MartaHughes | Performance of the Innova SARS-CoV-2 antigen rapid lateral flow test in the Liverpool asymptomatic testing pilot: population based cohort study | 2021-07-01 | doi: https://doi.org/10.1136/bmj.n1637 |  | United Kingdom | COVID-19 testing site/Screening location | Nov. 8-29, 2020 | Asymptomatic; mean age 50yrs; 54% female; 82% caucasian | 5869 | Nasal Swabs (NS) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "The Innova LFT seems, in combination with other health protection measures, to be a valuable tool in wider public health responses to covid-19 for identifying those with higher viral loads who are more likely to be infectious but do not report classic symptoms"
Terry | Sherli | 93 | García-Fiñana MartaHughes | Performance of the Innova SARS-CoV-2 antigen rapid lateral flow test in the Liverpool asymptomatic testing pilot: population based cohort study |  | doi: https://doi.org/10.1136/bmj.n1637 | [6 empty] | Throat Swabs (TS)
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-01 |  |  | Switzerland | COVID-19 testing site/Screening location | September 2021 - December 2021 | Individuals with suspected SARS-CoV-2 infection. Age >18, mean age 31.4, 221 Female | 377 | Saliva (Sal) |  | Specific Test Not Described |  | Reference test not stated |  |  | Contact author about reference test | Michael Nagler michael.nagler@insel.ch
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting |  | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Nasopharyngeal swabs (NP)
Terry - finished | Sherli | 98 | S. Jirungda; P. Khamphrom; K. Klaprajun; T. Prachakul; J. Kraisin | Clinical performance of the standard F COVID-19 AG FIA for the detection of SARS-COV-2 infection | 2022-05-01 |  |  | Thailand | Hospital - Inpatient | July 22-29, 2021 | Not provided | 423 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  |  | "The rapid antigen test for SARS-CoV-2 observed comparable sensitiv- ity and specificity with the real-time RT-PCR assay." |  |  | Jirungda, Biomolecular Laboratory, Department of Medical Technology and Clinical Pathology, Saraburi Hospital, Thailand 18000. No email address provided.
Terry - finished | Sherli | 102 | M. Kahn; L. Schuierer; C. Bartenschlager; S. Zellmer; R. Frey; M. Freitag; C. Dhillon; M. Heier; A. Ebigbo; C. Denzel; S. Temizel; H. Messmann; M. Wehler; R. Hoffmann; E. Kling; C. Römmele | Performance of antigen testing for diagnosis of COVID-19: a direct comparison of a lateral flow device to nucleic acid amplification based tests | 2021-08-01 |  |  | Germany | Emergency Dept/Room | 6/23/2020 - 11/26/2020 | Hospital patients. Demographics not included. | Not specified. Total of 3630 antigen tests. | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "This study was able to show that the real-life performance of the examined lateral-fow-based antigen test was signifcantly below the manufacturer’s specifcations. Tis indicates that the use of this test for the detection of individual patients infected with SARS-CoV2 should be discouraged." |  |  | *Correspondence: maria_kahn@gmx.de
Terry - finished | Sherli | 102 | M. Kahn; L. Schuierer; C. Bartenschlager; S. Zellmer; R. Frey; M. Freitag; C. Dhillon; M. Heier; A. Ebigbo; C. Denzel; S. Temizel; H. Messmann; M. Wehler; R. Hoffmann; E. Kling; C. Römmele | Performance of antigen testing for diagnosis of COVID-19: a direct comparison of a lateral flow device to nucleic acid amplification based tests |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Terry - finished | Sherli | 102 | M. Kahn; L. Schuierer; C. Bartenschlager; S. Zellmer; R. Frey; M. Freitag; C. Dhillon; M. Heier; A. Ebigbo; C. Denzel; S. Temizel; H. Messmann; M. Wehler; R. Hoffmann; E. Kling; C. Römmele | Performance of antigen testing for diagnosis of COVID-19: a direct comparison of a lateral flow device to nucleic acid amplification based tests |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | RealStar SARS-CoV-2 RT-PCR kit (Altona)
Terry - finished | Sherli | 102 | M. Kahn; L. Schuierer; C. Bartenschlager; S. Zellmer; R. Frey; M. Freitag; C. Dhillon; M. Heier; A. Ebigbo; C. Denzel; S. Temizel; H. Messmann; M. Wehler; R. Hoffmann; E. Kling; C. Römmele | Performance of antigen testing for diagnosis of COVID-19: a direct comparison of a lateral flow device to nucleic acid amplification based tests |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Panther Fusion or Aptima SARS-CoV-2 assay (Hologic)
Terry | Sherli | 104 | Gupta A | Rapid chromatographic immunoassay-based evaluation of COVID-19: A cross-sectional, diagnostic test accuracy study & its implications for COVID-19 management in India | 2021-01-01 | doi: 10.4103/ijmr.IJMR_3305_20 | Ancestral | India | College/University Campus (Medical Center/Hospital) | May 31, 2020 - July 24, 2020 | Symptomatic and Asymptomatic; median age 34.1 yr; male to female ratio 0.42 | 330 | Nasal Swabs (NS) |  | BGI 2019-nCoV Real-time Fluorescent RT-PCR kit (BGI Genomics) |  |  | "In conclusion, the rapid antigen test showed an excellent specificity to ‘rule-in’ COVID-19 patients within the first five days of illness and had a moderate sensitivity. Therefore, patients showing positive result need to immediately triaged and those with negative tests should be reconfirmed by an rRT-PCR."
Terry | Sherli | 104 | Gupta A | Rapid chromatographic immunoassay-based evaluation of COVID-19: A cross-sectional, diagnostic test accuracy study & its implications for COVID-19 management in India |  | doi: 10.4103/ijmr.IJMR_3305_20 |  |  | Hospital - Inpatient |  |  |  | Throat Swabs (TS)
Terry - not finished (need to contact authors for data) | Sherli: Agree with Terry. The paper is comparing on site RAT with lab RAT results | 106 | H. H. Kessler; F. Prüller; M. Hardt; E. Stelzl; E. Föderl-Höbenreich; S. Pailer; A. Lueger; P. Kreuzer; K. Zatloukal; M. Herrmann | Identification of contagious SARS-CoV-2 infected individuals by Roche's Rapid Antigen Test | 2022-03-01 |  |  | Austria | Emergency Dept/Room | Not provided | 33 females. 34 males. Mean age = 56 years. | 67 | Nasopharyngeal swabs (NP) |  | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid) |  |  | "The present results demonstrate that the SARS-CoV-2 RAT from Roche Diagnostics is a valuable tool for managing symptomatic patients in an acute care setting where rapid decision making is critical. Our results support previous studies showing that RATs are less sensitive than RT-qPCR tests, but offer good specificity." |  | Data is not clear in this paper. it appears that the authors may need to be contacted for the statistical data required. 06/14/23 AELS: From the text, can pull the numbers for the Central Lab RAT versuse RT-PCR. Not enough info from On-site RATs. | *Corresponding author: Markus Herrmann, Prof. MD, Clinical E-mail: markus.herrmann@medunigraz.at
Terry - not finished (need to contact authors for data) | Sherli: Agree with Terry. The paper is comparing on site RAT with lab RAT results | 106 | H. H. Kessler; F. Prüller; M. Hardt; E. Stelzl; E. Föderl-Höbenreich; S. Pailer; A. Lueger; P. Kreuzer; K. Zatloukal; M. Herrmann | Identification of contagious SARS-CoV-2 infected individuals by Roche's Rapid Antigen Test | 2022-03-01 | [9 empty] | Cobas Kits/Systems (Roche)
Terry/Amy | Sherli | 106 | H. H. Kessler; F. Prüller; M. Hardt; E. Stelzl; E. Föderl-Höbenreich; S. Pailer; A. Lueger; P. Kreuzer; K. Zatloukal; M. Herrmann | Identification of contagious SARS-CoV-2 infected individuals by Roche's Rapid Antigen Test | 2022-03-01 |  |  | Austria | Emergency Dept/Room | Not provided | 33 females. 34 males. Mean age = 56 years. | 67 | Nasopharyngeal swabs (NP) |  | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid) |  |  | "The present results demonstrate that the SARS-CoV-2 RAT from Roche Diagnostics is a valuable tool for managing symptomatic patients in an acute care setting where rapid decision making is critical. Our results support previous studies showing that RATs are less sensitive than RT-qPCR tests, but offer good specificity." |  | Data is not clear in this paper. it appears that the authors may need to be contacted for the statistical data required. | *Corresponding author: Markus Herrmann, Prof. MD, Clinical Institute of Medical and Chemical Laboratory Diagnostics, Medical University of Graz, Auenbruggerplatz 15, 8036 Graz, Austria, Phone: +43 316 385 13145, Fax: +43 316 385 13430, E-mail: markus.herrmann@medunigraz.at
Terry/Amy | Sherli | 106 | H. H. Kessler; F. Prüller; M. Hardt; E. Stelzl; E. Föderl-Höbenreich; S. Pailer; A. Lueger; P. Kreuzer; K. Zatloukal; M. Herrmann | Identification of contagious SARS-CoV-2 infected individuals by Roche's Rapid Antigen Test | 2022-03-01 | [9 empty] | Cobas Kits/Systems (Roche)
Terry - finished | Sherli | 109 | C. W. E. L.-d. H. M. S. R. F. C. P. R. v. K. S. King | Validation of the Panbio™ COVID-19 Antigen Rapid Test (Abbott) to screen for SARS-CoV-2 infection in Sint Maarten: a diagnostic accuracy study | Pre-print Sep 2021 |  |  | Sint Maarten, Dutch Caribbean | COVID-19 testing site/Screening location | Jan. 11 - Feb. 26, 2021 | Median age: 40.5 years. 59% female. 59% asymptomatic. 60% had no contact with COVID19+ people. | 1477 | Nasopharyngeal swabs (NP) |  | Custom/In-house SARS-2 primers |  |  | "the Abbott Panbio RAT is a valid and cheaper alternative to RT-PCR for detecting COVID-19 when used on symptomatic individuals among the general population." | "For the accuracy analysis, 66 participants were excluded because they had missing RT-PCR Ct 163 values." |  | *Corresponding author: 16 Dr. Sanne Christine van Kampen 17 Collective Prevention Services 18 Ministry of Health, Social Development and Labor 19 W.G. Buncamper Road 33, Vineyard Building 20 Philipsburg, Sint Maarten 21 Email: Sanne.vankampen@sintmaartengov.org 22 Telephone: +17215249361
Yati | Sherli | 109 | Harris David T. | SARS-CoV-2 Rapid Antigen Testing of Symptomatic and Asymptomatic Individuals on the University of Arizona Campus | 2021-05-01 | https://doi.org/10.3390/biomedicines9050539 |  | USA | College/University Campus (non-medical) | June to August 2020. | Symptomatic and Asymptomatic; Students, faculty, staff at U of A | 2436 | Nasal Swabs (NS) |  | IDT SARS-CoV-2 (2019-nCoV) multiplex CDC qPCR probe Assay (Integrated DNA Technologies) |  |  | "The rapid Ag test may be best for determining if subjects are contagious and most likely to spread the virus, while the PCR test may be a better indicator for general infection and prevalence of the virus in the general population. Overall, the rapid antigen test seems to be sufficient and more than adequate for rapid screening of both asymptomatic and symptomatic subjects."
Yati | Sherli | 109 | Harris David T. | SARS-CoV-2 Rapid Antigen Testing of Symptomatic and Asymptomatic Individuals on the University of Arizona Campus |  | https://doi.org/10.3390/biomedicines9050539 | [6 empty] | Nasopharyngeal swabs (NP) |  | Alinity m SARS-CoV-2 AMP (Abbott)
Terry - finished | Sherli | 110 | Y. Kiyasu; Y. Takeuchi; Y. Akashi; D. Kato; M. Kuwahara; S. Muramatsu; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Prospective analytical performance evaluation of the QuickNavi™-COVID19 Ag for asymptomatic individuals | 2021-07-01 |  |  | Japan | COVID-19 testing site/Screening location | October 7, 2020 - January 9, 2021 | Patients with suspected COVID infections based on symptoms. Individuals in contact with others having known COVID infections. | 1881 participants but 1934 swabs | Nasopharyngeal swabs (NP) |  | GENECUBE (Toyobo Co., Ltd.) |  |  | "The QuickNaviTM-COVID19 Ag showed lower sensitivity for the asymptomatic group than for symptomatic patients. However, its specificity was consistently high, and no false positives were found in this study."
Terry - finished | Sherli | 110 | Y. Kiyasu; Y. Takeuchi; Y. Akashi; D. Kato; M. Kuwahara; S. Muramatsu; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Prospective analytical performance evaluation of the QuickNavi™-COVID19 Ag for asymptomatic individuals |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Japanese National Institute of Infectious Diseases (NIID)
Terry - finished | Sherli | 110 | Y. Kiyasu; Y. Takeuchi; Y. Akashi; D. Kato; M. Kuwahara; S. Muramatsu; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Prospective analytical performance evaluation of the QuickNavi™-COVID19 Ag for asymptomatic individuals |  | https://doi.org/10.1093/labmed/lmaa104 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Terry - finished | Sherli | 111 | A. Klajmon; A. Olechowska-Jarząb; D. Salamon; A. Sroka-Oleksiak; M. Brzychczy-Włoch; T. Gosiewski | Comparison of antigen tests and qpcr in rapid diagnostics of infections caused by sars-cov-2 virus | 2021-12-01 |  |  | Poland | Hospital - Inpatient | November 2020 - April 2021 | Symptomatic patients. Mean age: 64.7 years. 61.1% men, 39.9% women | 192 | Nasopharyngeal swabs (NP) |  | Vitassay (Vitassay) |  |  | "study showed that effectiveness of the antigen tests in rapid laboratory diagnostics is high enough to be an alternative and support for nucleic acid amplification tests (NAAT) in the virus replication phase in the course of COVID-19." |  |  | Correspondence: tomasz.gosiewski@uj.edu.pl
Yati - done | Sherli | 113 | M. P. N. M. M. V. K. E. R. W. M. F. I. D. F. G. L. L. C. F. G. G. H. P. N. H. E. K. F. Korenkov | Assessment of SARS-CoV-2 infectivity by a Rapid Antigen Detection Test | Preprint May 2021 | doi: https://doi.org/10.1101/2021.03.30.21254624 |  | Germany | College/University Campus (Medical Center/Hospital) | Between 26th October 2020 and 08th January 2021 | Participants had a median age of 32.25 years (IQR: 26.14 - 43.12) | 2028 samples from 1849 individuals | Nasopharyngeal swabs (NP) | Combined oro- and nasopharyngeal swabs | LightMix SarbecoV (TIB Microbiol) |  |  | The Standard Q COVID-19 Ag Test yielded a positive result in 92 (4.54%) samples resulting in an overall sensitivity and specificity of 42.86% and 99.89%. While overall sensitivity was low, Standard Q COVID-19 RADT reliably detected patients with high RNA loads. |  | Stratified by adjusted Ct values the RADT had sensitivities of 100% (14/14), 98.25% (56/57), 88.64% (78/88) and 50.57% (89/176) for adjusted Ct values <20, <25, <30, <35, respectively (Table 1). Diagnostic sensitivities of 93.33%, 55.55% and 22.22% are reached for adjusted Ct values between 25-26, 27-28 and 29-30 and sensitivities of 86.36%, 29,17% and 9.68% for 6 log10, 5 log10 and 4 log10 copies/ml, respectively (Fig. 2D). We conclude that observed RADT sensitivity declines at adjusted Ct values above 27 or below 6 log10 copies/ml but reliably detects samples with higher RNA loads.29,17% and 9.68% for 6 log10, 5 log10 and 4 log10 copies/ml, respectively (Fig. 2D). We conclude that observed RADT sensitivity declines at adjusted Ct values above 27 or below 6 log10 copies/ml but reliably detects samples with higher RNA loads. | Correspondance: florian.klein@uk-koeln.de
Yati - done | Sherli | 113 | M. P. N. M. M. V. K. E. R. W. M. F. I. D. F. G. L. L. C. F. G. G. H. P. N. H. E. K. F. Korenkov | Assessment of SARS-CoV-2 infectivity by a Rapid Antigen Detection Test | Preprint May 2021 | doi: https://doi.org/10.1101/2021.03.30.21254624 |  | Germany |  |  |  |  | Oropharyngeal swabs (OP)
Yati | Sherli | 114 | Holzner C. | SARS-CoV-2 rapid antigen test: Fast-safe or dangerous? An analysis in the emergency department of an university hospital | 2021-05-01 | https://doi.org/10.1002/jmv.27033 | Ancestral | Germany | College/University Campus (Medical Center/Hospital) | October 16, 2020 and January 14, 2021. | retrospective of patients with acute complaints that presented at ED | 2375 | Nasopharyngeal swabs (NP) |  | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid) |  |  | "Due to our study results, we strongly suggest that the use of the antigen rapid tests for “rule‐out” should be critically questioned, as they suggest a supposed safety that cannot be scientifically supported. Nevertheless, this test procedure enables the rapid and easy detection of approximately two‐thirds of infectious patients in the context of large‐scale screening, even without or with only minor nonspecific symptoms, which otherwise go undetected."
Yati | Sherli | 114 | Holzner C. | SARS-CoV-2 rapid antigen test: Fast-safe or dangerous? An analysis in the emergency department of an university hospital |  | https://doi.org/10.1002/jmv.27033 |  |  | Emergency Dept/Room | [5 empty] | RealStar SARS-CoV-2 RT-PCR kit (Altona)
Yati | Sherli | 114 | Holzner C. | SARS-CoV-2 rapid antigen test: Fast-safe or dangerous? An analysis in the emergency department of an university hospital |  | https://doi.org/10.1002/jmv.27033
Yati | Sherli | 115 | Homza M. | Five Antigen Tests for SARS-CoV-2: Virus Viability Matters | 2021-04-01 | https://doi.org/10.3390/v13040684 | Ancestral or Alpha | Czech Republic | COVID-19 testing site/Screening location | Not listed - received by journal March 13, 2021 | Median age ~43 (data separated by RAT) and ~53% female | 1141 | Nasopharyngeal swabs (NP) |  | COVID-19 Multiplex RT-PCR kit (DIANA Biotech) |  |  | "This leads us to propose that (a) virus viability testing, at least of the discrepant samples, should form a part of AGT evaluation and validation process; (b) we propose that antigen testing using a well-performing AGT could represent a valid tool for infection control in high prevalence settings, and (c) that more detailed research on the patients who are “missed” by (well-performing) AGT is needed with respect to their infectiousness to determine whether or not they pose a threat from the perspective of disease spreading."
Chathu finished | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-01 |  |  | Germany | COVID-19 testing site/Screening location | November 2020 - December 2020. | All participants were asked for information on comorbidi- ties, symptom presence and duration, and severity of disease | 767 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene)
Chathu finished | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-01 |  |  | Germany |  |  |  |  | Oropharyngeal swabs (OP) |  | LightMix SarbecoV (TIB Microbiol)
Terry - finished | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-01 |  |  | Germany, Brazil | COVID-19 testing site/Screening location | Germany: April 17, 2020 - March 31, 2021, Brazil: July 27 - Sept 16, 2020 | Median age 36, with 51.1% female | 7471 total, 5742 from Germany, 1729 from Brazil | Nasopharyngeal swabs (NP) |  | Abbott RealTime SARS-CoV-2 (Abbott) |  |  | "Variability in test performance is partially explained by variable viral loads in population evaluated over the course of the pandemic. All Ag-RDTs reach high sensitivity early in the disease and in individuals with high viral loads, supporting their role in identifying transmission relevant infections." |  |  | E-mail address: claudia.denkinger@uni-heidelberg.de (C.M. Denkinger).
Terry - finished | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-01 |  |  | Germany, Brazil | Public area (not a designated screening location) |  |  |  | Oropharyngeal swabs (OP) |  | Cobas Kits/Systems (Roche)
Terry - finished | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-01 |  |  | Germany | [6 empty] | LightMix SarbecoV (TIB Microbiol)
Terry - finished | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-01 |  |  | Brazil | [6 empty] | Allplex assays (Seegene)
Yati - done | Sherli | 118 | L. J. L. A. K. Krüger; M. T. F. K. J. S. S. L. F. S. P. N. O. M. F. P. Gaeddert; J. C. V. M. Seybold; T. C. Jones; N. R. Pollock; B. W. A. W. S. S. N. S. J. S. H. P. A. T. M. P. L. Knorr; M. O. S. S. J. A. de Vos; C. D. C. Escadafal | A Multi-Center Clinical Diagnostic Accuracy Study of Surestatus - an Affordable, WHO Emergency-Use-Listed, Rapid, Point-of-Care, Antigen-Detecting Diagnostic Test for SARS-CoV-2 (preprint) | 2022-12-01 |  |  | India | Not described/unclear | Dec 2019 to Mar 2021 | Not available | 1119 | Nasopharyngeal swabs (NP) |  | Specific Test Not Described |  |  | A total of 1119 participants were included in the analysis of whom 205 (18·3%) were RT-PCR positive. SureStatus detected 169 out of 205 RT-PCR positive participants, reporting a sensitivity of 82·4% (95% CI 76·6%-87·1%) and a specificity of 98·5% (95% CI 97·4%-99·1%). |  | Article is a pre-print | Need to contact Authors but cannot find contact info
Yati - done | Sherli | 118 | L. J. L. A. K. Krüger; M. T. F. K. J. S. S. L. F. S. P. N. O. M. F. P. Gaeddert; J. C. V. M. Seybold; T. C. Jones; N. R. Pollock; B. W. A. W. S. S. N. S. J. S. H. P. A. T. M. P. L. Knorr; M. O. S. S. J. A. de Vos; C. D. C. Escadafal | A Multi-Center Clinical Diagnostic Accuracy Study of Surestatus - an Affordable, WHO Emergency-Use-Listed, Rapid, Point-of-Care, Antigen-Detecting Diagnostic Test for SARS-CoV-2 (preprint) | 2022-12-01 | https://doi.org/10.1093/labmed/lmaa104 |  | Germany |  |  |  |  | Oropharyngeal swabs (OP)
Yati - done | Sherli | 118 | L. J. L. A. K. Krüger; M. T. F. K. J. S. S. L. F. S. P. N. O. M. F. P. Gaeddert; J. C. V. M. Seybold; T. C. Jones; N. R. Pollock; B. W. A. W. S. S. N. S. J. S. H. P. A. T. M. P. L. Knorr; M. O. S. S. J. A. de Vos; C. D. C. Escadafal | A Multi-Center Clinical Diagnostic Accuracy Study of Surestatus - an Affordable, WHO Emergency-Use-Listed, Rapid, Point-of-Care, Antigen-Detecting Diagnostic Test for SARS-CoV-2 (preprint) | 2022-12-01 | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Other | Combined NP+OP
Ellyn | Sherli | 119 | Igloi Z. | Clinical Evaluation of Roche SD Biosensor Rapid Antigen Test for SARS-CoV-2 in Municipal Health Service Testing Site, the Netherlands | 2021-05-01 | doi: 10.3201/eid2705.204688 | Ancestral | Netherlands | COVID-19 testing site/Screening location | 9-15 October 2020 | Urban; average age 42; 54.7% female; 73.4% symptomatic | 970 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "We conclude that the use of Ag RDT in our drive-through test stations would provide a good method to identify most infectious patients."
Ellyn | Sherli | 119 | Igloi Z. | Clinical Evaluation of Roche SD Biosensor Rapid Antigen Test for SARS-CoV-2 in Municipal Health Service Testing Site, the Netherlands |  | doi: 10.3201/eid2705.204688 | [6 empty] | Oropharyngeal swabs (OP)
Karen - finished | Sherli | 119 | Y. Kurihara; Y. Kiyasu; Y. Akashi; Y. Takeuchi; K. Narahara; S. Mori; T. Takeshige; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | The evaluation of a novel digital immunochromatographic assay with silver amplification to detect SARS-CoV-2 | 2021-07-01 | https://doi.org/10.1016/j.jiac.2021.07.006 |  | Japan | Hospital - Inpatient | February 8 and March 24, 2021. | unknown | 1401 samples | Nasopharyngeal swabs (NP) |  | GENECUBE (Toyobo Co., Ltd.) |  | Used an in-house RT-PCR protocol in addition to reference RT-PCR | 1401 of 1416 samples were analyzed. Reference test detected SARS-CoV-2in 83 (5.9%) of the samples tested. In-house test missed one sample positive by reference test. Of the 83 samples, 36 (43.4%) were collected from symptomatic patients, and 47 (56.6%) were obtained from asymptomatic participants. |  | Of the 83 samples, 36 (43.4%) were collected from symptomatic patients, and 47 (56.6%) were obtained from asymptomatic participants. Sixty-two of the 83 samples that were positive by reference RT-PCR were also positive by the antigen test. | Yes, if we want to include data from the 4 additional Ag tests. Email address: kiyasu-tuk@umin.ac.jp (Yoshihiko Kiyasu).
Karen - finished | Sherli | 119 | Y. Kurihara; Y. Kiyasu; Y. Akashi; Y. Takeuchi; K. Narahara; S. Mori; T. Takeshige; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | The evaluation of a novel digital immunochromatographic assay with silver amplification to detect SARS-CoV-2 | 2021-07-01 | https://doi.org/10.1016/j.jiac.2021.07.006 | [8 empty] | Japanese National Institute of Infectious Diseases (NIID)
Terry - finished | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-01 |  |  | Korea | Hospital - Inpatient | Dec. 20, 2020 - March 12, 2021 | Age: avg. 68.0 (54.3-81.3), 38.2% female | 68 patients; 282 samples | Nasopharyngeal swabs (NP) |  | Real-Q 2019-nCoV Detection Kit (Biosewoom) |  |  | "In conclusion, the diagnostic sensitivity of RAT was unsatisfactory, even for specimens with high viral loads or in the initial period after symptom onset." |  | The paper only specifies sensitivity and no other info | Ji Young Park, MD, PhD Department of Pediatrics, Chung-Ang University Hospital, 102 Heukseok-ro, Dongjak- gu, Seoul 06973, Korea. Email: jypark@caumc.or.kr Seong Hwan Kim, MD, PhD Department of Orthopaedic Surgery, Chung- Ang University Hospital, 102 Heukseok-ro, Dongjak-gu, Seoul 06973, Korea. Email: ksh170177@caumc.or.kr
Ellyn | Sherli | 123 | Jakobsen K. | Accuracy and cost description of rapid antigen test compared with reverse transcriptase-polymerase chain reaction for SARS-CoV-2 detection | 2021-06-01 | https://www.researchgate.net/publication/352374494 | Ancestral | Denmark | COVID-19 testing site/Screening location | 26-31 December 2020 | People who booked an appointment at public testing centers | 4697 | Oropharyngeal swabs (OP) |  | Luna Universal Probe One-Step RT-PCR for Detection of COVID-19 (SignaGen Labs) |  |  | "The AG test is a good and relevant supplement to RT-PCR testing in public SARS-CoV-2 screenings"
Ellyn | Sherli | 123 | Jakobsen K. | Accuracy and cost description of rapid antigen test compared with reverse transcriptase-polymerase chain reaction for SARS-CoV-2 detection |  | https://www.researchgate.net/publication/352374494 | [6 empty] | Nasopharyngeal swabs (NP)
Terry - finished | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-01 |  |  | Greece | Emergency Dept/Room | March 19, 2021 - April 24, 2021 | Not provided | 624 | Nasal Swabs (NS) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "Rapid Test Ag 2019-nCoV (ProGnosis Biotech SA) showed an overall sensitivity of 85.5% and an inter-rater agreement (weighted Kappa) when compared with RT-PCR of 0.89, indicating that it could accurately identify SARS-CoV-2 antigens in people with suspected COVID-19." |  |  | Correspondence: xhatzi@uth.gr
Terry - finished | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-01 | https://doi.org/10.1093/labmed/lmaa104 |  |  | COVID-19 testing site/Screening location |  |  |  | Nasopharyngeal swabs (NP)
Terry - finished | Sherli | 125 | L. Landaverde; J. Turcinovic; L. Doucette-Stamm; K. Gonzales; J. Platt; J. H. Connor; C. Klapperich | Comparison of BinaxNOW TM and SARS-CoV-2 qRT-PCR detection of the Omicron Variant from Matched Anterior Nares Swabs (preprint) | 2022-03-01 |  |  | United States, Boston, Massachusetts | College/University Campus (Medical Center/Hospital) | January - February 2022 | Boston University Campus Community | 300 participants; 319 samples | Nasal Swabs (NS) |  | Specific Test Not Described |  | Test kit not specified | "Our results suggest that RDT tests could be used in our setting to confirm SARS-CoV-2 infection in individuals with substantial viral load, but that a significant fraction of infected individuals would be missed if we used RDT tests exclusively to rule out infection." |  |  | Not specified
Terry - finished | Sherli | 126 | E. Layer; S. Hoehl; M. Widera; D. Bojkova; T. Westphal; R. Gottschalk; B. Boeddinghaus; J. Schork; S. Ciesek; U. Goetsch | SARS-CoV-2 screening strategies for returning international travellers: Evaluation of a rapid antigen test approach | 2022-02-01 |  |  | Germany | COVID-19 testing site/Screening location | February - May 2021 | "Aged 18 and 85 (mean age 39.2) years. Approximately 53.3% (n = 793) were male, 46.0% (n = 685) were female, and 0.7% (n = 10) were nonbinary" | 1488 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "Overall, we conclude that a RAT no earlier than day 5, when performed under study conditions, is an appropriate method for early termination of mandatory quarantine after travel. The fact that none of the negative participants tested positive during thefollowing 14 days indicates that day 5 presents a suitable time to test (Kiang et al., 2021)." |  | Need to contact authors for additional data. | Corresponding author: Dr. Udo Goetsch, Address: Breite Gasse 28, 60313 Frank- furt am Main, Germany, Tel: +49 69 212 36981. E-mail address: udo.goetsch@stadt-frankfurt.de (U. Goetsch).
Terry - finished | Sherli | 126 | E. Layer; S. Hoehl; M. Widera; D. Bojkova; T. Westphal; R. Gottschalk; B. Boeddinghaus; J. Schork; S. Ciesek; U. Goetsch | SARS-CoV-2 screening strategies for returning international travellers: Evaluation of a rapid antigen test approach | 2022-02-01 | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Oropharyngeal swabs (OP) |  | Alinity m SARS-CoV-2 AMP (Abbott)
Ellyn | Sherli | 126 | Jegerlehner Sabrina | Diagnostic accuracy of a SARS-CoV-2 rapid antigen test in real-life clinical settings | 2021-08-01 | https://doi.org/10.1016/j.ijid.2021.07.010 | Alpha | Switzerland | COVID-19 testing site/Screening location | January - March 2021 | Suspected SARS-2 infection | 1465 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "The accuracy of the SARS-CoV-2 Roche/SD Biosensor rapid antigen test in diagnosing SARS- CoV-2 infections in a primary/secondary care testing facility was considerably lower compared with the manufacturer’s data. Widespread application in such a setting might lead to a considerable number of individuals falsely classified as SARS-CoV-2 negative."
Ellyn | Sherli | 126 | Jegerlehner Sabrina | Diagnostic accuracy of a SARS-CoV-2 rapid antigen test in real-life clinical settings |  | https://doi.org/10.1016/j.ijid.2021.07.010 |  |  | College/University Campus (Medical Center/Hospital) | [5 empty] | Allplex assays (Seegene)
Karen/Amy | Sherli | 128 | J. LeGoff; S. Kernéis; C. Elie; S. Mercier-Delarue; N. Gastli; L. Choupeaux; J. Fourgeaud; M. L. Alby; P. Quentin; J. Pavie; P. Brazille; M. L. Néré; M. Minier; A. Gabassi; C. Leroy; B. Parfait; J. M. Tréluyer; C. Delaugerre | Evaluation of a saliva molecular point of care for the detection of SARS-CoV-2 in ambulatory care | 2021-10-01 | https://doi.org/10.1038/s41598-021-00560-8 |  | France | COVID-19 testing site/Screening location | 11/4/2020-2/15/2021 | 26-52 years (37 median); 55%female - One to three symptoms were observed in 530/1712 (31%) participants. | 1718 | Nasopharyngeal swabs (NP) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "When compared to a composite reference standard (including the results of nasopharyngeal RT-PCR and diferent RT-PCR tests on saliva), the sensitivity of saliva RT-PCR was higher than nasopharyngeal RT-PCR." |  | Authors were actually testing a POC RT-LAMP assay, but also provided AG test sensitivity and specificity data for the nasopharyngeal samples | Jérôme LeGoff email: jerome.le-goff@aphp.fr
Karen/Amy | Sherli | 128 | J. LeGoff; S. Kernéis; C. Elie; S. Mercier-Delarue; N. Gastli; L. Choupeaux; J. Fourgeaud; M. L. Alby; P. Quentin; J. Pavie; P. Brazille; M. L. Néré; M. Minier; A. Gabassi; C. Leroy; B. Parfait; J. M. Tréluyer; C. Delaugerre | Evaluation of a saliva molecular point of care for the detection of SARS-CoV-2 in ambulatory care | 2021-10-01 | https://doi.org/10.1038/s41598-021-00560-8 | [6 empty] | Saliva (Sal)
Ellyn | Sherli | 136 | Kim D.Lee | Development and Clinical Evaluation of an Immunochromatography-Based Rapid Antigen Test (GenBody™ COVAG025) for COVID-19 Diagnosis | 2021-04-01 | https://doi.org/10.3390/v13050796 | Ancestral | Prospective: India | COVID-19 testing site/Screening location | 25 January 2021-3 February 2021 | Unknown - not listed | 200 | Nasopharyngeal swabs (NP) |  | EURORealTime SARS-CoV-2 (Euroimmun) |  |  | "Given the speed, low complexity and accuracy of the GenBodyTM COVID-19 Ag test COVAG025, it is predicted to be suitable for the rapid identification of positive patients. Thus, we believe the GenBodyTM COVID-19 Ag test has potential for use as a simple and rapid SARS-CoV-2 antigen detection test, especially in high prevalence areas." | Retrospective study done as case control. Use only prospective study data.
Second Pass: AELS | Sherli | 141 | D. K. Mandal; B. R. Bhattarai; S. Pokhrel; M. Chhusyabaga; P. Bhandari; M. P. Bhatt; S. B. Marhattha | Diagnostic Performance of SARS-CoV-2 Rapid Antigen Test in relation to RT-PCR CqValue | 2022-05-01 | https://doi.org/10.1155/2022/9245248 |  | Nepal | Hospital - Inpatient | Nov 2020 - Jan 2021 | median age was 35 years (IQR 27–46.5). Male 67.1% (n = 143/213) | 213 | Nasopharyngeal swabs (NP) |  | Specific Test Not Described |  | Authors did not specify a commercial RT-PCR kit but provided info about the mastermix used | The promising performance of Ag-RDT renders it useful as screening tool alongside RT-PCR to reduce transmission via improving contact tracing, implementation of local mitigation strategies, and rening existing testing protocol for diagnosis of COVID-19. |  | A bit concerned that this might be a case-controlled study - based on the note as in the sections below: Methods section: Each individual specimen was initially screened for SARS-CoV-2 N-Ag using a rapid antigen test (Espline®,Japan) based on the principle of lateral flow assay. Results section: Among samples tested for possible SARS-CoV-2 infection by rapid antigen test method, 36.2% (n = 77/213) indicated positive results and 63.8% (n = 136/213) showed negative test results. All the test results were then confirmed by real time RT-PCR assay, where 55.4% (n = 118/213) tested positive, while 44.6% (n = 95/213) tested negative for SARS-CoV-2 viral RNA genome as shown in Table 1.
Second Pass: AELS | Sherli | 141 | D. K. Mandal; B. R. Bhattarai; S. Pokhrel; M. Chhusyabaga; P. Bhandari; M. P. Bhatt; S. B. Marhattha | Diagnostic Performance of SARS-CoV-2 Rapid Antigen Test in relation to RT-PCR CqValue | 2022-05-01 | https://doi.org/10.1155/2022/9245248 |  |  | Hospital - Outpatient |  |  |  | Throat Swabs (TS) | [6 empty] | Not case-control -AELS 06/2/23
Ellyn | Sherli | 144 | Klein J. | Head-to-head performance comparison of self-collected nasal versus professional-collected nasopharyngeal swab for a WHO-listed SARS-CoV-2 antigen-detecting rapid diagnostic test | 2021-05-01 | https://doi.org/10.1007/s00430-021-00710-9 |  | Germany | COVID-19 testing site/Screening location | Dec 15, 2020-Jan 19, 2021 | Symptomatic or close contacts; median age 42.7 yrs; 52.4% female | 292 | Nasopharyngeal swabs (NP) |  | LightMix SarbecoV (TIB Microbiol) |  |  | "Our study suggests that supervised NMT self-sampling leads to results comparable to NP sampling for the Panbio Ag-RDT."
Terry | Sherli | 148 | Kohmer N. | Article the comparative clinical performance of four SARS-CoV-2 rapid antigen tests and their correlation to infectivity in vitro | 2021-01-01 | https://doi.org/10.3390/jcm10020328 | Ancestral | Germany | Long-term Care Facility (Nursing home, rehab centers) | 2020-11-01 | Not described beyond in sharing living facility | 100 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "The clinical sensitivity within the potential infectious range and moderate agreement with cell culture, including the above-mentioned limitation, allow large-scale SARS-CoV-2 Ag-RDTs-based testing to be considered as a surrogate marker for identifying potentially infective individuals in a population, reducing the spread to others."
Terry | Sherli | 152 | Korenkov M. | Evaluation of a rapid antigen test to detect SARS-CoV-2 infection and identify potentially infectious individuals | 2021-08-01 | DOI: 10.1128/JCM.00896-21 | Ancestral | Germany | College/University Campus (Medical Center/Hospital) | Oct. 26, 2020 - Jan. 11, 2021 | Not given | 2028 | Oropharyngeal swabs (OP) |  | Custom/In-house SARS-2 primers |  | N-gene | "Furthermore, one might suspect that RT-qPCR positive, but RADT-negative individuals do not pose a high risk of transmissions, since all samples remained culture negative in our experimental setup. However, individual results must be interpreted with caution, as SARS-CoV-2 infection could remain undetected in early stages. In summary, our results suggest that SARS-CoV-2 transmission could be reduced by systematic RADT, despite the fact that some infected individuals will not be detected by the test"
Terry | Sherli | 152 | Korenkov M. | Evaluation of a rapid antigen test to detect SARS-CoV-2 infection and identify potentially infectious individuals |  | DOI: 10.1128/JCM.00896-21 | [6 empty] | Nasopharyngeal swabs (NP) |  | LightMix SarbecoV (TIB Microbiol)
Terry | Sherli | 152 | Korenkov M. | Evaluation of a rapid antigen test to detect SARS-CoV-2 infection and identify potentially infectious individuals |  | DOI: 10.1128/JCM.00896-21 | [8 empty] | Cobas Kits/Systems (Roche)
Terry | Sherli | 152 | Korenkov M. | Evaluation of a rapid antigen test to detect SARS-CoV-2 infection and identify potentially infectious individuals |  | DOI: 10.1128/JCM.00896-21 | [8 empty] | Alinity m SARS-CoV-2 AMP (Abbott)
Terry | Sherli | 152 | Korenkov M. | Evaluation of a rapid antigen test to detect SARS-CoV-2 infection and identify potentially infectious individuals |  | DOI: 10.1128/JCM.00896-21 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Terry - finished | Sherli | 155 | A. Mizrahi; J. C. Nguyen Van; N. El Helali; J. Lourtet-Hascoet; I. Jabnoune; M. L. Pacreau; Y. Talb; J. Fourgeaud; M. Leruez-Ville; B. Pilmis; V. Avettand-Fenoel; A. Le Monnier | The Coris BioConcept COVID 19 Ag Respi-Strip, a field experience feedback | 2021-11-01 |  |  | France | Hospital - Inpatient | April 6 - 10, 2020; April 12-16, 2020 | not specified | 294 | Nasopharyngeal swabs (NP) |  | RealStar SARS-CoV-2 RT-PCR kit (Altona) |  |  | "the Coris-Ag system could find a use in laboratories not equipped with RT-PCR or as a point-of-care doctor’s test on the condition that negative results be confirmed by RT-PCR." |  |  | Corresponding author at: Service de Microbiologie clinique et dosages des anti-infectieux, Groupe Hospitalier Paris Saint-Joseph, 185 rue Raymond Losserand, 75014, Paris, France. E-mail addresses: amizrahi@ghpsj.fr, amizrahi@hpsj.fr (A. Mizrahi).
Terry - finished | Sherli | 156 | I. J. B. Møller; A. R. Utke; U. K. Rysgaard; L. J. Østergaard; S. Jespersen | Diagnostic performance, user acceptability, and safety of unsupervised SARS-CoV-2 rapid antigen-detecting tests performed at home | 2022-01-01 |  |  | Denmark | COVID-19 testing site/Screening location | March 25, 2021, to April 4, 2021, for the Hangzhou RADT onsite | mean age 42 years; 50.5% were female | 827 | Oropharyngeal swabs (OP) |  | Specific Test Not Described |  |  | "this study showed that self-RADTs for analyzing nasal swab samples was a reliable, user-acceptable, safe complementary test to PCR analyses of professionally collected oropharyngeal swab samples." |  |  | ∗ Corresponding author: Ida Johanne Bocher Møller, E-mail addresses: idajbm@gmail.com (I.J.B. Møller)
Terry - finished | Sherli | 156 | I. J. B. Møller; A. R. Utke; U. K. Rysgaard; L. J. Østergaard; S. Jespersen | Diagnostic performance, user acceptability, and safety of unsupervised SARS-CoV-2 rapid antigen-detecting tests performed at home | 2022-01-01 |  |  |  |  | March 10, 2021, to March 31, 2021, for the Hangzhou RADT Offsite |  |  | Nasopharyngeal swabs (NP) | NP were for the Abbott test
Terry - finished | Sherli | 156 | I. J. B. Møller; A. R. Utke; U. K. Rysgaard; L. J. Østergaard; S. Jespersen | Diagnostic performance, user acceptability, and safety of unsupervised SARS-CoV-2 rapid antigen-detecting tests performed at home | 2022-01-01 |  |  |  |  | February 16, 2021, to March 10, 2021, for the DNA Diagnostic RADT Offsite
Terry - finished | Sherli | 156 | I. J. B. Møller; A. R. Utke; U. K. Rysgaard; L. J. Østergaard; S. Jespersen | Diagnostic performance, user acceptability, and safety of unsupervised SARS-CoV-2 rapid antigen-detecting tests performed at home | 2022-01-01 |  |  |  |  | January 21, 2021, to January 25, 2021, for both the DNA Diagnostic RADT and the Abbott RADT onsite
Terry |  | 159 | Krüger L. | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | https://doi.org/10.1371/journal.pone.0247918 | Ancestral | Germany | COVID-19 testing site/Screening location | Sept 28, 2020 - Oct 30, 2020 | Urban; Symptomatic or close contact; mean age 39.4yrs; 50.7% female | 1108 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  |  | "In summary, the favourable ease-of-use results and the limited infrastructure required for the Ag-RDT testing procedure, its high specificity in addition to the high sensitivity of the test in persons with high viral load, can empower control of population transmission if implemented in well-designed testing programs"
Terry |  | 159 | Krüger L. | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use |  | https://doi.org/10.1371/journal.pone.0247918 |  |  | Urgent Care location |  |  |  | Oropharyngeal swabs (OP) |  | Cobas Kits/Systems (Roche)
Terry |  | 159 | Krüger L. | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use |  | https://doi.org/10.1371/journal.pone.0247918 | [8 empty] | LightMix SarbecoV (TIB Microbiol)
Terry - finished | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-01 |  |  | Germany | College/University Campus (Medical Center/Hospital) | November 30, 2020 - January 18, 2021 | Average age: 34.6 years with 46.7% females | 228 | Nasal Swabs (NS) |  | Specific Test Not Described |  |  | "Professional anterior nasal and nasal mid-turbinate sampling are of equivalent accuracy for an antigen-detecting rapid diagnostic test in ambulatory symptomatic adults." |  |  | Andreas K. Lindner andreas.lindner@charite.de
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-01 | https://doi.org/10.1371/journal.pone.0262399 |  | Hungary | Hospital - Inpatient | 10/2/2020-2/28/2021 | Can calcluate from supplemental data but I didn't. The numbers in the paper include children. | 1572 | Nasopharyngeal swabs (NP) |  | PerkinElmer SARS-CoV-2 Real-time RT-PCR Assay |  | The outcome of the Ag-RTD was unclear in the case of 4 cases, and RT-qPCR gave inconclusive result with 13 samples | "it is easy and convenient to use antigen-based RTDs as a tool for point-of-care testing among symptomatic patients; however, on-field applicability for screening of asymptomatic individuals in schools, offices or factories remains concern." |  |  | palyi.bernadett@nnk.gov.hu
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-01 | https://doi.org/10.1371/journal.pone.0262399 |  |  | COVID-19 testing site/Screening location
Terry - finished | Sherli | 167 | G. A. Okoye; H. I. Kamara; M. Strobeck; T. A. Mellman; J. Kwagyan; A. Sullivan; A. S. Byrd; B. Shokrani; H. E. Mighty | Diagnostic accuracy of a rapid diagnostic test for the early detection of COVID-19 | 2021-10-01 |  |  | United States | College/University Campus (non-medical) | September to December 2020 | not specified | 3810 nasal swab samples (#participants not specified) | Nasal Swabs (NS) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "The results of this study of the BinaxNOW rapid antigen test in symptomatic and asymptomatic individuals demonstrate a lower sensitivity (91.84%) than that reported by the manufacturer (97.1%). However, it is important to note that the sensitivity of this test is higher in symptomatic individuals"
Terry - finished | Sherli | 168 | S. N. Onsongo; K. Otieno; S. van Duijn; E. Adams; M. Omollo; I. A. Odero; A. K'Oloo; N. Houben; E. Milimo; R. Aroka; H. C. Barsosio; F. Oluoch; A. Odhiambo; S. Kariuki; T. F. R. de Wit | Performance of a rapid antigen test for SARS-CoV-2 in Kenya | 2021-11-01 |  |  | Kenya | Primary Care Location | December 28, 2020 - March 31, 2021 | Median age: 37 years (28-49 years). 53.8% male. 46.2% female. | 997 | Nasopharyngeal swabs (NP) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "The performance of the kit almost met the WHO cut-off specifications in detecting COVID-19 when compared to the current gold standard method. We note that The BioNote NowCheck rapid antigen test is a realistic testing option for usage in Africa" |  |  | Corresponding author. Tel.: +254 720 345 110. E-mail address: simonya@live.com (S.N. Onsongo).
Terry - finished | Sherli | 168 | S. N. Onsongo; K. Otieno; S. van Duijn; E. Adams; M. Omollo; I. A. Odero; A. K'Oloo; N. Houben; E. Milimo; R. Aroka; H. C. Barsosio; F. Oluoch; A. Odhiambo; S. Kariuki; T. F. R. de Wit | Performance of a rapid antigen test for SARS-CoV-2 in Kenya | 2021-11-01 | [7 empty] | Oropharyngeal swabs (OP)
Terry - finished | Sherli | 171 | K. C. Paap; A. M. van Loon; F. M. Koene; L. W. van Buul; S. Jurriaans; M. Smalbrugge; M. D. de Jong; C. Hertogh | Clinical evaluation of single-swab sampling for rapid COVID-19 detection in outbreak settings in Dutch nursing homes | 2021-11-01 |  |  | Netherlands | Long-term Care Facility (Nursing home, rehab centers) | December 1, 2020 - March 31, 2021 | not specified. Only identified as nursing home residents. | 461 | Other | Samples were only referred to as collected by the single swab method. They were not described any other way. | Panther Fusion or Aptima SARS-CoV-2 assay (Hologic) |  |  | "Given the low sensitivity and specifcity of the single-swab method, the reliability and advantages of this method in the NH setting became questionable. Yet, sensitivity increased when considering the lower CT values, similar to a laboratory setting, indicating a higher viral load and thereby identifying the most infectious persons. Specifcity was much higher when the single-swab method was performed in a laboratory setting." |  |  | Anouk M. van Loon a.vanloon@amsterdamumc.nl
Terry - finished | Sherli | 172 | A. K. Pandey; A. Mohanty; V. Hada; R. S. Rath; S. Kumar; S. Kishore; R. Kant | Comparison of the Rapid Antigen Testing Method With RT-qPCR for the Diagnosis of COVID-19 | 2021-08-01 |  |  | India | Hospital - Inpatient | not specified | Mean age was 29.4 +-10.11, 81% male | 318 | Nasopharyngeal swabs (NP) |  | Specific Test Not Described |  |  | "It was seen in this study that the rapid antigen test performed well in the patients with a low CT value, i.e., a high viral load. Thus, all such cases if diagnosed early can help in breaking the chain of transmission." |  | Should contact authors for some of the information which I designated as not specified. | Aroop Mohanty, aroopmohanty7785@yahoo.com
Amy | Sherli | 174 | J. Y. Park; J. H. Lee; B. K. Cha; B. S. Kim; H. J. Lee; G. H. Kim; K. T. Kang; Y. S. Lee; S. K. Ahn; S. H. Kim | Analysis of the Efficacy of Universal Screening of Coronavirus Disease with Antigen-Detecting Rapid Diagnostic Tests at Point-or-Care Settings and Sharing the Experience of Admission Protocol—A Pilot Study | 2022-02-01 | https://doi.org/10.3390/jpm12020319 |  | Korea | Hospital - Inpatient | 1 December 2020, and 28 February 2021, | 47.5% female; mean age 70.1 ± 13.1 (PCR+) 63.4 ± 19.9 (PCR-) | 646 | Other | Not specified | Allplex assays (Seegene) |  |  | RT-PCR can be slow and the usage of RAT for initial screening is useful for triage. |  |  | Seong Hwan Kim: ksh170177@nate.com; Tel.: +82-2-6299-3143
Ellyn | Sherli | 175 | Leixner G. | Evaluation of the AMP SARS-CoV-2 rapid antigen test in a hospital setting | 2021-07-01 | https://doi.org/10.1016/j.ijid.2021.05.063 | Ancestral | Austria | Hospital - Inpatient | 2 Nov - 29 Dec 2020 | Symptomatic; median age 70; 49% female | 392 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "In conclusion, the AMP RAT showed good test performance in patients with low Ct values and might be a valuable tool for frontline testing whenever quick rRT-PCR testing is not feasible."
Ellyn | Sherli | 175 | Leixner G. | Evaluation of the AMP SARS-CoV-2 rapid antigen test in a hospital setting |  | https://doi.org/10.1016/j.ijid.2021.05.063 |  |  | Hospital - Outpatient |  |  |  | Oropharyngeal swabs (OP) |  | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Ellyn | Sherli | 175 | Leixner G. | Evaluation of the AMP SARS-CoV-2 rapid antigen test in a hospital setting |  | https://doi.org/10.1016/j.ijid.2021.05.063 | [8 empty] | BD Max (Becton-Dickinson)
Ellyn | Sherli | 175 | Leixner G. | Evaluation of the AMP SARS-CoV-2 rapid antigen test in a hospital setting |  | https://doi.org/10.1016/j.ijid.2021.05.063 | [8 empty] | Simplexa (DiaSorin)
Amy | Sherli | 177 | W. F. Peacock; K. M. Soto-Ruiz; S. L. House; C. M. Cannon; G. Headden; B. Tiffany; S. Motov; K. Merchant-Borna; A. M. Chang; C. Pearson; B. W. Patterson; A. E. Jones; J. Miller; J. Varon; A. Bastani; C. Clark; Z. Rafique; B. Kea; J. Eppensteiner; J. M. Williams; S. A. Mahler; B. E. Driver; P. Hendry; E. Quackenbush; D. Robinson; J. W. Schrock; J. P. D'Etienne; C. J. Hogan; A. Osborne; R. Riviello; S. Young | Utility of COVID-19 antigen testing in the emergency department | 2022-01-01 | https://doi.org/10.1002/emp2.12605 |  | United States | Emergency Dept/Room | July 20 to October 28, 2020, | mean age 47; 57.4% female | 735 | Nasopharyngeal swabs (NP) |  | Abbott RealTime SARS-CoV-2 (Abbott) |  |  | "The clinical utility of the BinaxNOW is that it can be performed and resulted in 15 minutes by nonlaboratorians without significant equipment." |  |  | W. Frank Peacock,MD, 800Wharf Street, SW, Washington DC, 20021. Email: Frankpeacock@gmail.com
Amy | Sherli | 177 | W. F. Peacock; K. M. Soto-Ruiz; S. L. House; C. M. Cannon; G. Headden; B. Tiffany; S. Motov; K. Merchant-Borna; A. M. Chang; C. Pearson; B. W. Patterson; A. E. Jones; J. Miller; J. Varon; A. Bastani; C. Clark; Z. Rafique; B. Kea; J. Eppensteiner; J. M. Williams; S. A. Mahler; B. E. Driver; P. Hendry; E. Quackenbush; D. Robinson; J. W. Schrock; J. P. D'Etienne; C. J. Hogan; A. Osborne; R. Riviello; S. Young | Utility of COVID-19 antigen testing in the emergency department | 2022-01-01 | https://doi.org/10.1002/emp2.12605 |  |  | Hospital - Inpatient |  |  |  | Nasal Swabs (NS)
Amy | Sherli | 180 | C. Peronace; R. Tallerico; M. Colosimo; V. Sacco; R. Talarico; M. De Fazio; F. Pasceri; I. Talotta; G. Panduri; J.-H. Kim; E. Cione; P. Minchella | Validation of GeneFinder COVID-19 Ag Plus Rapid Test and Its Potential Utility to Slowing Infection Waves: A Single-Center Laboratory Evaluation Study | 2022-05-01 | https://doi.org/10.3390/diagnostics12051126 |  | Italy | Emergency Dept/Room | 2021-10-01 | Not given | 603 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  |  | "Therefore, based on our findings, we suggested that the GeneFinder COVID-19 Ag Plus Rapid Test can be used as a screening test for early SARS-CoV-2 diagnosis to slow down virus spread." |  |  | Correspondence: Cinzia Peronace cinziaperonace@hotmail.it (C.P.); Erika Cione erika.cione@unical.it (E.C.)
Ellyn | Sherli | 181 | Linares M. | Panbio antigen rapid test is reliable to diagnose SARS-CoV-2 infection in the first 7 days after the onset of symptoms | 2020-10-01 | https://doi.org/10.1016/j.jcv.2020.104659 | Ancestral | Spain | Emergency Dept/Room | Sept 10 - 15, 2020 | Urban, 50% female ED; 53.3% female PC; mean age 51.5 ED; mean age 39.0 PC | 255 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  |  | "The RDT evaluated in this study showed a high sensitivity and specificity in samples mainly obtained during the first week of symptoms and with high viral loads (Ct<25)."
Ellyn | Sherli | 181 | Linares M. | Panbio antigen rapid test is reliable to diagnose SARS-CoV-2 infection in the first 7 days after the onset of symptoms |  | https://doi.org/10.1016/j.jcv.2020.104659 |  |  | Primary Care Location
Ellyn | Sherli | 182 | Lindner A. | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with self-collected nasal swab versus professional-collected nasopharyngeal swab | 2021-04-01 | DOI: 10.1183/13993003.03961-2020 | Ancestral | Germany | COVID-19 testing site/Screening location | Sept 23 - Oct 14, 2020 | 34.7 yrs median age; 42.9% female; Urban | 289 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "In conclusion, this study demonstrates that supervised self-sampling from the nose is a reliable alternative to professional nasopharyngeal sampling with STANDARD Q."
Ellyn | Sherli | 182 | Lindner A. | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with self-collected nasal swab versus professional-collected nasopharyngeal swab |  | DOI: 10.1183/13993003.03961-2020 | [6 empty] | Nasal Swabs (NS) |  | LightMix SarbecoV (TIB Microbiol)
Ellyn | Sherli | 182 | Lindner A. | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with self-collected nasal swab versus professional-collected nasopharyngeal swab |  | DOI: 10.1183/13993003.03961-2020 | [6 empty] | Oropharyngeal swabs (OP)
Ellyn | Sherli | 183 | Lindner A. | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with professional-collected nasal versus nasopharyngeal swab | 2021-05-01 | DOI: 10.1183/13993003.03961-2020 | Ancestral | Germany | COVID-19 testing site/Screening location | Nov 11-18, 2020 | 36.2 yrs median age; 48.0% female; Urban | 179 | Nasal Swabs (NS) |  | Cobas Kits/Systems (Roche) |  |  | "In conclusion, this study demonstrates that sensitivity of a WHO-listed SARS-CoV-2 Ag-RDT using professional nasal sampling kit is at least equal to that of NP sampling kit, although confidence intervals overlap."
Ellyn | Sherli | 183 | Lindner A. | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with professional-collected nasal versus nasopharyngeal swab |  | DOI: 10.1183/13993003.03961-2020 | [6 empty] | Nasopharyngeal swabs (NP) |  | LightMix SarbecoV (TIB Microbiol)
Ellyn | Sherli | 183 | Lindner A. | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with professional-collected nasal versus nasopharyngeal swab |  | DOI: 10.1183/13993003.03961-2020 | [6 empty] | Oropharyngeal swabs (OP)
Ellyn | Sherli | 184 | Lindner A. | Diagnostic accuracy and feasibility of patient self-testing with a SARS-CoV-2 antigen-detecting rapid test | 2021-05-01 | https://doi.org/10.1016/j.jcv.2021.104874 | Ancestral | Germany | COVID-19 testing site/Screening location | Nov 30-Dec 11, 2020 | 35.0 mean age; 51.4% female; Urban; Symptomatic | 146 | Nasal Swabs (NS) |  | Cobas Kits/Systems (Roche) |  |  | "In conclusion, we demonstrate that symptomatic laypersons can reliably perform a SARS-CoV-2 lateral-flow Ag-RDT and test themselves."
Ellyn | Sherli | 184 | Lindner A. | Diagnostic accuracy and feasibility of patient self-testing with a SARS-CoV-2 antigen-detecting rapid test |  | https://doi.org/10.1016/j.jcv.2021.104874 | [6 empty] | Nasopharyngeal swabs (NP) |  | LightMix SarbecoV (TIB Microbiol)
Ellyn | Sherli | 184 | Lindner A. | Diagnostic accuracy and feasibility of patient self-testing with a SARS-CoV-2 antigen-detecting rapid test |  | https://doi.org/10.1016/j.jcv.2021.104874 | [6 empty] | Oropharyngeal swabs (OP)
Chathu - completed | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-01 |  |  | Germany | College/University Campus (Medical Center/Hospital) | Dec 6-13, 2021 | 420 (60.3%) female; participants with an overall median age of 31 years (IQR: 26 - 41; | 696 samples | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche)
Chathu - completed | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-01 | [7 empty] | Oropharyngeal swabs (OP) |  | Alinity m SARS-CoV-2 AMP (Abbott)
Chathu - completed | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-01 | [9 empty] | LightMix SarbecoV (TIB Microbiol)
Chathu - finished | Sherli | 193 | M. Prost | Evaluation of a Rapid in Vitro Diagnostic Test Device for Detection of SARS-CoV-2 Antigen in Nasal Swabs | 2021-12-01 |  |  | France | College/University Campus (Medical Center/Hospital) | September to October 2021 | Participant information is not given | 420 NS samples | Nasal Swabs (NS) |  | Vitassay (Vitassay) |  |  | The specificity of Ag tests was generally high and it showed that false positive test results were unlikely.
Chathu - finished | Sherli | 193 | M. Prost | Evaluation of a Rapid in Vitro Diagnostic Test Device for Detection of SARS-CoV-2 Antigen in Nasal Swabs | 2021-12-01 | [7 empty] | Nasopharyngeal swabs (NP)
Ellyn | Sherli | 194 | Masiá M. | Nasopharyngeal Panbio COVID-19 Antigen Performed at Point-of-Care Has a High Sensitivity in Symptomatic and Asymptomatic Patients with Higher Risk for Transmission and Older Age | 2021-02-01 | https://doi.org/10.1093/ofid/ofab059 | Ancestral | Spain | Primary Care Location | September 15-October 29, 2020 | 46.3% Male; 40.6 years median age; Symptomatic and Asymptomatic | 931 | Nasopharyngeal swabs (NP) |  | LightMix SarbecoV (TIB Microbiol) |  |  | "In conclusion, the sensitivity of the NP Panbio COVID-19 antigen RDT is closely related to the Ct values, age, and the presence and duration of symptoms."
Ellyn | Sherli | 194 | Masiá M. | Nasopharyngeal Panbio COVID-19 Antigen Performed at Point-of-Care Has a High Sensitivity in Symptomatic and Asymptomatic Patients with Higher Risk for Transmission and Older Age |  | https://doi.org/10.1093/ofid/ofab059 |  |  | Emergency Dept/Room |  |  |  | Nasal Swabs (NS)
Ellyn | Sherli | 194 | Masiá M. | Nasopharyngeal Panbio COVID-19 Antigen Performed at Point-of-Care Has a High Sensitivity in Symptomatic and Asymptomatic Patients with Higher Risk for Transmission and Older Age |  | https://doi.org/10.1093/ofid/ofab059 | [6 empty] | Saliva (Sal)
Amy | Sherli | 195 | M. M. Rahman; A. F. Hoque; Y. Karim; Z. Kawser; A. B. Siddik; M. K. Sumiya; A. Siddika; M. S. Alam; A. N. Alam; M. Alam; M. E. Hossain; S. Banu; F. Qadri; T. Shirin; M. Rahman; M. Z. Rahman | Clinical evaluation of SARS-CoV-2 antigen-based rapid diagnostic test kit for detection of COVID-19 cases in Bangladesh | 2021-11-01 | https://doi.org/10.1016/j.heliyon.2021.e08455 |  | Bangladesh | Hospital - Inpatient | Dec 2020 - April 2021 | Characteristics are combined for adults and children | 861 adults | Nasopharyngeal swabs (NP) |  | Custom/In-house SARS-2 primers |  | Chinese CDC recommend of RdRp and N genes | The RDT showed strong clinical performance in detecting SARS-CoV-2 infections among the COVID-19 suspected cases in this study. |  | Out of the 205 subjected to sequencing, only 179 sequences were retrievable of which 12 (6.7%) were Alpha (B.1.1.7), 109 (60.89%) were Beta (B.1.351) and 58 (32.40%) were Wuhan-Hu-1 like. Note: only the sensitivity was given for adults. All other data was combined with kids
Amy | Sherli | 195 | M. M. Rahman; A. F. Hoque; Y. Karim; Z. Kawser; A. B. Siddik; M. K. Sumiya; A. Siddika; M. S. Alam; A. N. Alam; M. Alam; M. E. Hossain; S. Banu; F. Qadri; T. Shirin; M. Rahman; M. Z. Rahman | Clinical evaluation of SARS-CoV-2 antigen-based rapid diagnostic test kit for detection of COVID-19 cases in Bangladesh | 2021-11-01 | https://doi.org/10.1016/j.heliyon.2021.e08455 |  |  | Hospital - Outpatient
Lupe | Sherli | 196 | A. F. A. S. N. Y. M. T. A. Raja Kamran | DIAGNOSTIC ACCURACY OF PANBIO COVID-19 RAPID ANTIGEN METHOD FOR SCREENING IN EMERGENCY CASES | 2021-09-01 |  |  | Pakistan | College/University Campus (Medical Center/Hospital) | Jan. to March 2021 | inclusion criteria for this study was patients that present as an emergency medical or surgical case | 1539 | Nasopharyngeal swabs (NP) |  | Sansure Biotech COVID-19 Nucleic Acid Test kit |  |  | PANBIO RAT had an "excellent" diagnostic accuracy of 99.8% and the authors recommend its use for screening |  |  | Dr. Afzal Kamran Raja (could not find email)
Lupe | Sherli | 196 | A. F. A. S. N. Y. M. T. A. Raja Kamran | DIAGNOSTIC ACCURACY OF PANBIO COVID-19 RAPID ANTIGEN METHOD FOR SCREENING IN EMERGENCY CASES | 2021-09-01 | https://doi.org/10.1093/labmed/lmaa104 | [6 empty] | Oropharyngeal swabs (OP)
Lupe | Sherli | 197 | N. Rana; P. Thakur; S. Saxena; V. Manchanda; R. Arora | Evaluation of the currently used antigen-based Rapid Diagnostic Test for the detection of SARS CoV-2 virus in respiratory specimens | Sept 2021 |  |  | India | College/University Campus (Medical Center/Hospital) | not specified | patients visiting the opthalmology department | 564 | Nasopharyngeal swabs (NP) |  | Specific Test Not Described |  | reference test was not specified | RATs had a negative correlation with Ct values (the values were not specified) |  |  | could not find author information (and I even tried the medical school website)
Lupe | Sherli | 203 | J. Rohde; W. Himmel; C. Hofinger; T. T. Lâm; H. Schrader; J. Wallstabe; O. Kurzai; I. Gágyor | Diagnostic accuracy and feasibility of a rapid SARS-CoV-2 antigen test in general practice - a prospective multicenter validation and implementation study | 2022-02-01 |  |  | Germany | Primary Care Location | Dec 2020 - June 2021 | Suspected cases of COVID-19, 53% female, median age of 40 | 1450 | Nasopharyngeal swabs (NP) |  | Specific Test Not Described |  | Reference test was not specified, NP swabs for RT-PCR was sent to 6 different labs | Sensitivity of the POC test was lower than reported by the manufacturer. Decreased viral load lead to a decrease in sensitivity. |  |  | Jorn Rohde: rohde_j@ukw.de
Terry | Sherli | 211 | Nagura-Ikeda M. | Clinical Evaluation of Self-Collected Saliva by Quantitative Reverse Transcription-PCR (RT-qPCR), Direct RT-qPCR, Reverse Transcription-Loop-Mediated Isothermal Amplification, and a Rapid Antigen Test To Diagnose COVID-19 | 2020-08-01 | https://doi.org/10.1128/JCM.01438-20 | Ancestral | Japan | Hospital - Inpatient | Feb. 11 - May 13, 2020 | 46 years median age; 64.1% male; symptomatic and asymptomatic | 103 | Nasopharyngeal swabs (NP) |  | Custom/In-house SARS-2 primers |  | 2 targets on the N gene | "Self-collected saliva in the early phase of symptom onset is an alternative specimen option for diagnosing COVID-19. The RT-qPCR LDT, the cobas SARS-CoV-2 high-throughput system, direct RT-qPCR kits (except for one commercial kit), and RT-LAMP showed different sensitivities for detecting viral RNA in saliva specimens, but each can be selectively used according to the clinical setting and facilities if close attention is paid to any false-negative results. The rapid SARS-CoV-2 antigen test alone is not recommended for use at this time due to its low sensitivity."
Terry | Sherli | 211 | Nagura-Ikeda M. | Clinical Evaluation of Self-Collected Saliva by Quantitative Reverse Transcription-PCR (RT-qPCR), Direct RT-qPCR, Reverse Transcription-Loop-Mediated Isothermal Amplification, and a Rapid Antigen Test To Diagnose COVID-19 |  | https://doi.org/10.1128/JCM.01438-20 | [6 empty] | Saliva (Sal) |  | Takara Bio SARS-CoV-2 direct detection RT-qPCR kit
Terry | Sherli | 211 | Nagura-Ikeda M. | Clinical Evaluation of Self-Collected Saliva by Quantitative Reverse Transcription-PCR (RT-qPCR), Direct RT-qPCR, Reverse Transcription-Loop-Mediated Isothermal Amplification, and a Rapid Antigen Test To Diagnose COVID-19 |  | https://doi.org/10.1128/JCM.01438-20 | [6 empty] | Throat Swabs (TS) |  | Shimadzu Ampdirect 2019 novel coronavirus detection kit
Terry | Sherli | 211 | Nagura-Ikeda M. | Clinical Evaluation of Self-Collected Saliva by Quantitative Reverse Transcription-PCR (RT-qPCR), Direct RT-qPCR, Reverse Transcription-Loop-Mediated Isothermal Amplification, and a Rapid Antigen Test To Diagnose COVID-19 |  | https://doi.org/10.1128/JCM.01438-20 | [8 empty] | GENECUBE (Toyobo Co., Ltd.)
Terry | Sherli | 211 | Nagura-Ikeda M. | Clinical Evaluation of Self-Collected Saliva by Quantitative Reverse Transcription-PCR (RT-qPCR), Direct RT-qPCR, Reverse Transcription-Loop-Mediated Isothermal Amplification, and a Rapid Antigen Test To Diagnose COVID-19 |  | https://doi.org/10.1128/JCM.01438-20 | [8 empty] | Cobas Kits/Systems (Roche)
Lupe | Sherli | 213 | G. L. Salvagno; G. Gianfilippi; L. Pighi; S. De Nitto; B. M. Henry; G. Lippi | Real-world assessment of Fluorecare SARS-CoV-2 Spike Protein Test Kit | 2021-05-01 |  |  | Italy | Hospital - Outpatient | April 2-19, 2021 | Suspected cases of symptomatic COVID-19 or in close contact with COVID-19 cases. Mean age 47+-20 years, 55.1% women | 354 | Nasopharyngeal swabs (NP) |  | RealStar SARS-CoV-2 RT-PCR kit (Altona) |  |  | RAT worked well as a screening test at a higher viral load |  |  | Prof. Guiseppe Lippi giuseppe.lippi@univr.it
Lupe | Sherli | 213 | G. L. Salvagno; G. Gianfilippi; L. Pighi; S. De Nitto; B. M. Henry; G. Lippi | Real-world assessment of Fluorecare SARS-CoV-2 Spike Protein Test Kit | 2021-05-01 | https://doi.org/10.1093/labmed/lmaa104 |  |  | Hospital - Inpatient
Yati | Sherli | 225 | Osmanodja B. | Accuracy of a novel sars-cov-2 antigen-detecting rapid diagnostic test from standardized self-collected anterior nasal swabs | 2021-05-01 | https://doi.org/10.3390/jcm10102099 | Alpha | Germany | COVID-19 testing site/Screening location | 12 February and 22 March 2021 | 34.0 aver. age; 53.3% female; urban | 379 | Nasal Swabs (NS) |  | Cobas Kits/Systems (Roche) |  |  | "The standardized sampling and test procedure of the Ag-RDT investigated in this study may allow for more reliable self-testing, which can increase testing frequency and can have significant impact on the pandemic."
Yati | Sherli | 225 | Osmanodja B. | Accuracy of a novel sars-cov-2 antigen-detecting rapid diagnostic test from standardized self-collected anterior nasal swabs |  | https://doi.org/10.3390/jcm10102099 | [6 empty] | Nasopharyngeal swabs (NP)
Yati | Sherli | 225 | Osmanodja B. | Accuracy of a novel sars-cov-2 antigen-detecting rapid diagnostic test from standardized self-collected anterior nasal swabs |  | https://doi.org/10.3390/jcm10102099 | [6 empty] | Oropharyngeal swabs (OP)
Ellyn | Sherli | 237 | Peña M. | Performance of SARS-CoV-2 rapid antigen test compared with real-time RT-PCR in asymptomatic individuals | 2021-06-01 | https://doi.org/10.1016/j.ijid.2021.04.087 | Ancestral | Chile | COVID-19 testing site/Screening location | 14-17 January 2021 | 36.67 mean age; 51% male | 842 | Nasopharyngeal swabs (NP) |  | GenomeCoV19 Detection kit (ABM) |  |  | "Given the high predictive values in asymptomatic individuals and the fast test result implying faster tracing of infected individuals, these results support and provide policymakers with evidence that RAT might have a significant role in COVID-19 screening, testing and contact tracing strategies to control the COVID-19 pandemic in i) areas that lack suitable laboratories to perform SARS-CoV-2 real-time RT-PCR diagnostics, ii) areas where results take more than 24–48 h and iii) areas with high traffic of individuals such as border/customs, airports, interregional bus and train stations or in any mass testing campaign requiring rapid results."
Ellyn | Sherli | 237 | Peña M. | Performance of SARS-CoV-2 rapid antigen test compared with real-time RT-PCR in asymptomatic individuals |  | https://doi.org/10.1016/j.ijid.2021.04.087 |  |  | Public area (not a designated screening location)
Ellyn | Sherli | 237 | Peña M. | Performance of SARS-CoV-2 rapid antigen test compared with real-time RT-PCR in asymptomatic individuals |  | https://doi.org/10.1016/j.ijid.2021.04.087 |  |  | Long-term Care Facility (Nursing home, rehab centers)
Ellyn | Sherli | 238 | Peña-Rodríguez M. | Performance evaluation of a lateral flow assay for nasopharyngeal antigen detection for SARS-CoV-2 diagnosis | 2021-03-01 | https://doi.org/10.1002/jcla.23745 | Ancestral | Mexico |  | October - November 2020 | 36.6 mean age; 41.7% Male; urban | 369 | Nasopharyngeal swabs (NP) |  | DeCov19 kit Triplex (Genes2Life) |  |  | "In conclusion, there is an urgent need for rapid diagnosis so that the transmission burden is dampened. Several challenges must be overcome in the timely detection of SARS-CoV-2. The SD BIOSENSOR is a useful assay for this; nonetheless, important caveats must be considered before the community implementation of this diagnostic test."
Ellyn | Sherli | 238 | Peña-Rodríguez M. | Performance evaluation of a lateral flow assay for nasopharyngeal antigen detection for SARS-CoV-2 diagnosis |  | https://doi.org/10.1002/jcla.23745 | [6 empty] | Oropharyngeal swabs (OP)
Ellyn | Sherli | 241 | Pilarowski GenayLebel | Performance characteristics of a rapid SARS-CoV-2 antigen detection assay at a public plaza testing site in San Francisco | 2021-01-01 | doi: 10.1101/2020.11.02.20223891 | Ancestral | United States | COVID-19 testing site/Screening location | 3 days in September 2020 | 54% male; 77% 18-50 years | 878 | Nasal Swabs (NS) |  | Custom/In-house SARS-2 primers |  | Chan Zuckerberg Biohub, and UCSF | "In summary, under field conditions with supplementary technician training, the Binax-CoV2 assay accurately detected SARS-CoV-2 infection with high viral loads in both asymptomatic and symptomatic individuals."
Amy | Sherli | 244 | Pollock N. | Performance and Implementation Evaluation of the Abbott BinaxNOW Rapid Antigen Test in a High-Throughput Drive-Through Community Testing Site in Massachusetts | 2021-04-01 | DOI:https://doi.org/10.1128/JCM.00083-21 | Ancestral | USA | Hospital - Inpatient | Oct 26 2020 - Nov 12 2020 | 58.9% female; urban | 1380 | Nasal Swabs (NS) |  | CRSP SARS-CoV-2 (Clinical Research Sequencing Platform, Harvard/MIT) |  |  | "Thus, the test does not appear to be optimal for ruling out SARS-CoV-2 infection in asymptomatic adults or children; use in serial testing programs and for testing of contacts of known cases deserves independent study."
Amy | Sherli | 244 | Pollock N. | Performance and Implementation Evaluation of the Abbott BinaxNOW Rapid Antigen Test in a High-Throughput Drive-Through Community Testing Site in Massachusetts |  | DOI:https://doi.org/10.1128/JCM.00083-21 |  |  | COVID-19 testing site/Screening location
Amy | Sherli | 262 | Rastawicki W. | Evaluation of PCL rapid point of care antigen test for detection of SARS-CoV-2 in nasopharyngeal swabs | 2021-01-01 | https://doi.org/10.1002/jmv.26765 | Ancestral | Poland | Hospital - Outpatient | May - Sept 2020 | 57.7 yrs median age; 50% female | 42 | Nasopharyngeal swabs (NP) |  | Novel Coronavirus (2019-nCoV) Real Time Multiplex RT-PCR kit (Liferiver) |  |  | "In conclusion, we observed low sensitivity for the PCL COVID‐19 Ag rapid FIA in detecting active SARS‐CoV‐2 infection in hospitalized patients."
Amy | Sherli | 262 | Rastawicki W. | Evaluation of PCL rapid point of care antigen test for detection of SARS-CoV-2 in nasopharyngeal swabs |  | https://doi.org/10.1002/jmv.26765 | [8 empty] | DiaPlexQ (Solgent)
Amy | Sherli | 262 | Rastawicki W. | Evaluation of PCL rapid point of care antigen test for detection of SARS-CoV-2 in nasopharyngeal swabs |  | https://doi.org/10.1002/jmv.26765 | [8 empty] | Vitassay (Vitassay)
Amy | Sherli | 266 | Reza SoleimaniCorentin | Rapid COVID-19 antigenic tests: Usefulness of a modified method for diagnosis | 2021-05-01 | https://doi.org/10.1002/jmv.27094 | Ancestral | Belgium, Numar |  | November 13 - December 9, 2020 | Not given | 401 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  |  | "Our modified method of using these tests (using UTM tubes) could be easily integrated into the laboratory workflow without impacting the preanalytical process."
Amy | Sherli | 274 | Salvagno G. | Clinical assessment of the Roche SARS-CoV-2 rapid antigen test | 2021-01-01 | https://doi.org/10.1515/dx-2020-0154 | Ancestral | Italy | Hospital - Inpatient | Nov 16 - 30 2020 | Mean age 46 yrs; 56.4% female | 321 | Nasal Swabs (NS) |  | Allplex assays (Seegene) |  |  | "In conclusion, the results of our clinical evaluation of Roche SARS-CoV-2 Rapid Antigen Test demonstrates that this method seems a suitable approach for rapid nasopharyngeal swab screening of patients with high SARS-CoV-2 viral load, especially those with symptoms or in the pre-symptomatic phase."
Amy | Sherli | 274 | Salvagno G. | Clinical assessment of the Roche SARS-CoV-2 rapid antigen test |  | https://doi.org/10.1515/dx-2020-0154 |  | Italy |  |  |  |  | Throat Swabs (TS)
Amy | Sherli | 278 | Schildgen V. | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | https://doi.org/10.3390/pathogens10010038 | Ancestral | Germany | Hospital - Inpatient | Not given | Not given | 73 | Bronchoalveolar Lavage (BAL)/Bronchial sample |  | RealStar SARS-CoV-2 RT-PCR kit (Altona) |  |  | "Consequently, although it actually seems that exclusive antigen tests have research character rather than being true in vitro diagnostics (IVDs) and unfortunately cannot replace PCR assays, they should additionally be used to gain deeper insights into infectivity and the course of infection to develop more advanced testing strategies."
Amy | Sherli | 288 | Smith R. | Clinical Evaluation of Sofia Rapid Antigen Assay for Detection of Severe Acute Respiratory Syndrome Coronavirus 2 among Emergency Department to Hospital Admissions | 2021-06-01 | DOI: https://doi.org/10.1017/ice.2021.281 | Ancestral | USA | Hospital - Inpatient | Sept 21 - Dec 3, 2020 | Not given | 2887 | Nasopharyngeal swabs (NP) |  | Abbott RealTime SARS-CoV-2 (Abbott) |  |  | "Despite not meeting the requirements to replace the RT-PCR assay for detection of SARS-CoV-2 infection, our findings suggest the usefulness of Sofia SARS rapid antigen FIA to guide initial patient placement in a burdened, limited rapid PCR-capacity hospital setting for both symptomatic and asymptomatic patients."
Amy | Sherli | 288 | Smith R. | Clinical Evaluation of Sofia Rapid Antigen Assay for Detection of Severe Acute Respiratory Syndrome Coronavirus 2 among Emergency Department to Hospital Admissions |  | DOI: https://doi.org/10.1017/ice.2021.281 |  |  | Hospital - Outpatient | [5 empty] | Cobas Kits/Systems (Roche)
Amy | Sherli | 288 | Smith R. | Clinical Evaluation of Sofia Rapid Antigen Assay for Detection of Severe Acute Respiratory Syndrome Coronavirus 2 among Emergency Department to Hospital Admissions |  | DOI: https://doi.org/10.1017/ice.2021.281 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Amy | Sherli | 296 | Terpos E. | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 | 2021-05-01 | https://doi.org/10.3390/diagnostics11060995 | Ancestral or Alpha | Slovenia | COVID-19 testing site/Screening location | Not given | All symptomatic with less than 7 days of symptoms | 358 | Nasal Swabs (NS) |  | Allplex assays (Seegene) |  |  | "Nasopharyngeal and nasal swabs have comparable high sensitivity, specificity, and accuracy in symptomatic patients with COVID-19 symptoms experienced for less than 7 days."
Amy | Sherli | 296 | Terpos E. | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 |  | https://doi.org/10.3390/diagnostics11060995 |  |  | Primary Care Location |  |  |  | Nasopharyngeal swabs (NP)
Amy | Sherli | 296 | Terpos E. | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 |  | https://doi.org/10.3390/diagnostics11060995 |  |  | Hospital - Outpatient
Amy | Sherli | 296 | Terpos E. | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 |  | https://doi.org/10.3390/diagnostics11060995 |  |  | Emergency Dept/Room
Amy | Sherli | 297 | Thakur P. | Utility of Antigen-Based Rapid Diagnostic Test for Detection of SARS-CoV-2 Virus in Routine Hospital Settings | 2021-04-01 | https://doi.org/10.1093/labmed/lmab033 | Ancestral | India | Hospital - Inpatient | Oct 27 - Nov 23 2020 | Mean age 44.78 yrs; 54.9% male | 677 | Nasopharyngeal swabs (NP) |  | Standard M nCoV Real-Time Detection Kit (SD Biosenor) |  |  | "The high specificity of RDT warrants it as a diagnostic tool in patients with a positive result for SARS-CoV-2 infection. However, its low sensitivity does not qualify its use as a standalone test, especially in asymptomatic or presymptomatic patients, for whom RT-qPCR continues to be the gold-standard test. The utility of RDT is in diagnosing symptomatic patients and may not be particularly suited as a screening tool for patients with a low viral load."
Amy | Sherli | 297 | Thakur P. | Utility of Antigen-Based Rapid Diagnostic Test for Detection of SARS-CoV-2 Virus in Routine Hospital Settings |  | https://doi.org/10.1093/labmed/lmab033 | [6 empty] | Oropharyngeal swabs (OP)
Amy | Sherli | 299 | Toptan T. | Evaluation of a SARS-CoV-2 rapid antigen test: Potential to help reduce community spread? | 2020-12-01 | https://doi.org/10.1016/j.jcv.2020.104713 | Ancestral | Germany | College/University Campus (Medical Center/Hospital) | Not given | excess diagnostic samples - deidentified | 137 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "Our results suggest that the rapid antigen test can detect SARS-CoV-2 infected individuals with high viral loads and has potential in determining highly contagious individuals."
Amy | Sherli | 299 | Toptan T. | Evaluation of a SARS-CoV-2 rapid antigen test: Potential to help reduce community spread? |  | https://doi.org/10.1016/j.jcv.2020.104713 | [6 empty] | Oropharyngeal swabs (OP) |  | Luna Universal Probe One-Step RT-PCR for Detection of COVID-19 (SignaGen Labs)
Amy | Sherli | 305 | Trobajo-Sanmartín | Evaluation of the rapid antigen test CerTest SARS-CoV-2 as an alternative COVID-19 diagnosis technique | 2021-03-01 | https://doi.org/10.1080/23744235.2021.1902563 | Delta | Spain |  | Nov 17 2021 - Dec 10 2021 | Aver age 43.56; 62.14% female | 240 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "In conclusion, CerTest SARS-CoV-2 can be used as the first diagnostic method with the aim to reduce the number of RT-qPCR. Nevertheless those patients diagnosed as negatives by the rapid test must be confirmed by RT-qPCR, especially if a suspicion of COVID-19 exists."
Amy | Sherli | 305 | Trobajo-Sanmartín | Evaluation of the rapid antigen test CerTest SARS-CoV-2 as an alternative COVID-19 diagnosis technique |  | https://doi.org/10.1080/23744235.2021.1902563 | [8 empty] | Alinity m SARS-CoV-2 AMP (Abbott)
Amy | Sherli | 305 | Trobajo-Sanmartín | Evaluation of the rapid antigen test CerTest SARS-CoV-2 as an alternative COVID-19 diagnosis technique |  | https://doi.org/10.1080/23744235.2021.1902563 | [8 empty] | Allplex assays (Seegene)
Amy | Sherli | 306 | Turcato Gianni | Clinical application of a rapid antigen test for the detection of SARS-CoV-2 infection in symptomatic and asymptomatic patients evaluated in the emergency department: A preliminary report | 2021-03-01 | https://doi.org/10.1016/j.jinf.2020.12.012 | Ancestral | Italy | Emergency Dept/Room | July 1 2020 through Nov 10 2020 | Not given | 3410 |  | Unclear - mentions Swabs but not what type | Specific Test Not Described |  |  | "In conclusion, the use of a RAT implemented for infectious risk assessment directly in triage should be considered in EDs and could be an additional tool to address the challenge of containing the SARS-CoV-2 pandemic."
Ellyn | Sherli | 310 | Van der Moeren N. | Evaluation of the test accuracy of a SARS-CoV-2 rapid antigen test in symptomatic community dwelling individuals in the Netherlands | 2021-05-01 | https://doi.org/10.1371/journal.pone.0250886 | Ancestral | Netherlands | COVID-19 testing site/Screening location | Part 1 - 9/28/2020-9/30/2020 | adults, male and female - no specifics | 354 | Nasopharyngeal swabs (NP) |  | Cobas Kits/Systems (Roche) |  |  | "In conclusion, the VRD is a promising diagnostic test for testing of symptomatic community-dwelling individuals within seven days after symptom onset for the purpose of disease control."
Ellyn | Sherli | 310 | Van der Moeren N. | Evaluation of the test accuracy of a SARS-CoV-2 rapid antigen test in symptomatic community dwelling individuals in the Netherlands |  | https://doi.org/10.1371/journal.pone.0250886 |  |  |  | Part 2 - 9/26/2020-10/6/2020 | Part 2 - case control |  | Oropharyngeal swabs (OP) |  | Abbott RealTime SARS-CoV-2 (Abbott)
Amy | Sherli | 72 | A. Goga; E. S. Mayne; K. Woeber; S. Takuva; D. Nsibande; M. Lekalakala; S. Jaumdally; P. Mutevedzi; H. Vreede; B. Daniels; C. Kufe; K. Dheda; K. Chetty; G. E. Gray | UTILITY of COVID-19 POINT-OF-CARE ANTIGEN TESTS in LOW-MIDDLE INCOME SETTINGS | 2022-03-01 | N/A |  | South Africa | Primary Care Location | October 2020 - October 2021 | Not listed | 1816 | Nasal Swabs (NS) |  | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid) |  |  | "In a real world setting, during beta and delta waves, compared with PCR the sensitivity of rapid antigen tests ranged from 35-68%. This may reflect low viral loads at diagnosis." |  | Conference abstract | None listed
Amy | Sherli | 72 | A. Goga; E. S. Mayne; K. Woeber; S. Takuva; D. Nsibande; M. Lekalakala; S. Jaumdally; P. Mutevedzi; H. Vreede; B. Daniels; C. Kufe; K. Dheda; K. Chetty; G. E. Gray | UTILITY of COVID-19 POINT-OF-CARE ANTIGEN TESTS in LOW-MIDDLE INCOME SETTINGS | 2022-03-01 | [7 empty] | Nasopharyngeal swabs (NP) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher)
Amy | Sherli | 72 | A. Goga; E. S. Mayne; K. Woeber; S. Takuva; D. Nsibande; M. Lekalakala; S. Jaumdally; P. Mutevedzi; H. Vreede; B. Daniels; C. Kufe; K. Dheda; K. Chetty; G. E. Gray | UTILITY of COVID-19 POINT-OF-CARE ANTIGEN TESTS in LOW-MIDDLE INCOME SETTINGS | 2022-03-01 | [9 empty] | Allplex assays (Seegene)
Amy | Sherli | 212 | N. Salcedo; B. F. Sena; X. Qu; B. B. Herrera | Comparative Evaluation of Rapid Isothermal Amplification and Antigen Assays for Screening Testing of SARS-CoV-2 | 2022-02-01 | https://doi.org/10.3390/v14030468 |  | United States | Primary Care Location | Not Listed | Not described | 173 | Nasal Swabs (NS) |  | IDT SARS-CoV-2 (2019-nCoV) multiplex CDC qPCR probe Assay (Integrated DNA Technologies) |  |  | "Antigen tests and upcoming isothermal amplification assays fit these needs and could be scaled to millions of tests per day." |  |  | Correspondence: Bobby Brooke Herrera bbherrera@e25bio.com
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-01 | https://doi.org/10.1371/journal.pone.0270715 |  | United Kingdom | COVID-19 testing site/Screening location | March 31, 2021 - May 21, 2021 | mean age 40 years; 41.7% Male - for both recruitment periods | 100 | Throat Swabs (TS) | Both sample types in both recruitment sets | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "In conclusion, self-taken swabs for COVID-19 testing offer an accurate alternative to healthcare worker taken swabs for use with RDTs." |  | Also showed cumulative sensitivity for each CT value (e.g., CT<24.9, CT<29.9); No overall because the Self-taken and HCW taken were from the same subject and used the same PCR result. | Emily R. Adams Emily.adams@lstmed.ac.uk
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-01 | https://doi.org/10.1371/journal.pone.0270715 |  |  |  | July 2021 - August 9, 2021 |  | 150 | Nasal Swabs (NS)
Amy | Sherli | 227 | R. A. P. Selvabai; L. V. Koshy; P. Shanmugam | Diagnostic Efficacy of COVID-19 Rapid Antigen Detection Card in Diagnosis of SARS-CoV-2 | 2022-04-01 | https://doi.org/10.1055/s-0042-1744237 |  | India | Hospital - Inpatient | Not listed | not listed | 599 | Nasal Swabs (NS) |  | Standard M nCoV Real-Time Detection Kit (SD Biosenor) |  | Targets RdRP and E genes | "Athenese-Dx COVID-19 Ag Rapid test met the ICMR acceptance criteria, and it could be a potential tool for point-of-care testing purpose." |  |  | Priyadarshini Shanmugam, MD, PhD (e-mail: priyadarshini0018@gmail.com).
Amy | Sherli | 227 | R. A. P. Selvabai; L. V. Koshy; P. Shanmugam | Diagnostic Efficacy of COVID-19 Rapid Antigen Detection Card in Diagnosis of SARS-CoV-2 | 2022-04-01 | https://doi.org/10.1055/s-0042-1744237 | [6 empty] | Throat Swabs (TS)
Amy | Sherli | 227 | R. A. P. Selvabai; L. V. Koshy; P. Shanmugam | Diagnostic Efficacy of COVID-19 Rapid Antigen Detection Card in Diagnosis of SARS-CoV-2 | 2022-04-01 | https://doi.org/10.1055/s-0042-1744237 | [6 empty] | Nasopharyngeal swabs (NP)
Amy | Sherli | 230 | J. L. V. Shaw; V. Deslandes; J. Smith; M. Desjardins | Evaluation of the Abbott Panbio(TM) COVID-19 Ag rapid antigen test for the detection of SARS-CoV-2 in asymptomatic Canadians | 2021-07-01 | https://doi.org/10.1016/j.diagmicrobio.2021.115514 |  | Canada | COVID-19 testing site/Screening location | not listed | asymptomatic only | 3007 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  |  | "The Abbott PanbioTM COVID-19 Ag rapid antigen detection lateral flow assay showed reduced positive concordance with the reference, laboratory-based molecular assays for the diagnosis of SARS-CoV-2 infections in the asymptomatic population compared to the data available in the literature for symptomatic patients." |  |  | E-mail address: julshaw@eorla.ca (J.L.V. Shaw).
Amy | Sherli | 230 | J. L. V. Shaw; V. Deslandes; J. Smith; M. Desjardins | Evaluation of the Abbott Panbio(TM) COVID-19 Ag rapid antigen test for the detection of SARS-CoV-2 in asymptomatic Canadians | 2021-07-02 | https://doi.org/10.1016/j.diagmicrobio.2021.115515 |  |  | Public area (not a designated screening location) | [5 empty] | Cobas Kits/Systems (Roche)
Amy | Sherli | 230 | J. L. V. Shaw; V. Deslandes; J. Smith; M. Desjardins | Evaluation of the Abbott Panbio(TM) COVID-19 Ag rapid antigen test for the detection of SARS-CoV-2 in asymptomatic Canadians | 2021-07-03 | https://doi.org/10.1016/j.diagmicrobio.2021.115516 | [8 empty] | Panther Fusion or Aptima SARS-CoV-2 assay (Hologic) |  | Orf1a and Orf1b gene regions
Amy | Sherli | 230 | J. L. V. Shaw; V. Deslandes; J. Smith; M. Desjardins | Evaluation of the Abbott Panbio(TM) COVID-19 Ag rapid antigen test for the detection of SARS-CoV-2 in asymptomatic Canadians | 2021-07-05 | https://doi.org/10.1016/j.diagmicrobio.2021.115518 | [8 empty] | Custom/In-house SARS-2 primers |  | Public Health Ontario Laboratory; E gene
Amy | Sherli | 233 | Z. K. Siddiqui; M. Chaudhary; M. L. Robinson; A. B. McCall; R. Peralta; R. Esteve; C. W. Callahan; Y. C. Manabe; J. D. Campbell; J. K. Johnson; M. Elhabashy; M. Kantsiper; J. R. Ficke | Implementation and Accuracy of BinaxNOW Rapid Antigen COVID-19 Test in Asymptomatic and Symptomatic Populations in a High-Volume Self-Referred Testing Site | 2021-12-01 | https://doi.org/10.1128/Spectrum.01008-21 |  | United States | COVID-19 testing site/Screening location | Dec 23, 2020 - Jan 11, 2021 | female 53%; median age 31 years | 6061 | Nasal Swabs (NS) |  | CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel |  | Targets N Gene | "Rapid antigen tests could be considered for screening asymptomatic people when a 24-h turnaround RT-PCR is unavailable, especially for repeat testing in congregate settings." |  |  | Address correspondence to Zishan K. Siddiqui, zsiddiq1@jhu.edu.
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-01 | https://doi.org/10.3390/diagnostics12020475 |  | Mozambique | Hospital - Inpatient | Jan to March 2021 | median age 38 years; 54.3% female | 1277 | Nasopharyngeal swabs (NP) |  | Abbott RealTime SARS-CoV-2 (Abbott) |  |  | "Laboratory-based molecular diagnosis with good sensitivity, as RT-PCR, will still be required for SARS-CoV-2 triage among contacts of positive cases, asymptomatic cases, and symptomatic patients with negative results in Ag-RDTs." |  |  | Correspondence: Nádia Sitoe nadia.sitoe@ins.gov.mz
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-01 | https://doi.org/10.3390/diagnostics12020475 |  |  | COVID-19 testing site/Screening location | [5 empty] | Cobas Kits/Systems (Roche)
Amy | Sherli | 236 | M. Skvarč | Clinical validation of two immunochromatographic SARS-CoV-2 antigen tests in near hospital facilities | 2021-04-01 | doi:10.3855/jidc.14953 |  | Slovenia | COVID-19 testing site/Screening location | 2020-12-01 | Not given | 1127 | Nasopharyngeal swabs (NP) |  | Allplex assays (Seegene) |  |  | "We had some false-negative cases that were explicable with the low viral burden on mucosal barriers and mild disease, or with more than 5 days since the start of first symptoms when SARS-CoV-2 viral proteins already disappear from the mucosal barrier of the upper respiratory tract due to efficient immune response." |  |  | Miha Skvarč, MD, PhD mihaskvarc@hotmail.com
Amy | Sherli | 236 | M. Skvarč | Clinical validation of two immunochromatographic SARS-CoV-2 antigen tests in near hospital facilities | 2021-04-01 | doi:10.3855/jidc.14953 |  |  | Hospital - Outpatient | [5 empty] | Cobas Kits/Systems (Roche)
Amy | Sherli | 236 | M. Skvarč | Clinical validation of two immunochromatographic SARS-CoV-2 antigen tests in near hospital facilities | 2021-04-01 | doi:10.3855/jidc.14953 |  |  | Hospital - Inpatient | [5 empty] | Biofire
Amy | Sherli | 241 | J. Stohr; V. F. Zwart; G. Goderski; A. Meijer; C. R. S. Nagel-Imming; M. F. Q. Kluytmans-van den Bergh; S. D. Pas; F. van den Oetelaar; M. Hellwich; K. H. Gan; A. Rietveld; J. J. Verweij; J. L. Murk; W. van den Bijllaardt; J. Kluytmans | Self-testing for the detection of SARS-CoV-2 infection with rapid antigen tests for people with suspected COVID-19 in the community | 2021-08-01 | https://doi.org/10.1016/j.cmi.2021.07.039 |  | Netherlands | COVID-19 testing site/Screening location | Dec 23, 2020 - Jan 17, 2021 | Median age 41 years; 43.9% male | 3201 | Oropharyngeal swabs (OP) |  | Alinity m SARS-CoV-2 AMP (Abbott) |  |  | "In conclusion, we showed that self-testing using currently available RDT has a high specificity and relatively high sensitivity (75% - 80%) to identifyindividuals with a high probability of contagiousness. This application has the potential for frequent and extensive testing, which may be an aid to lift current restrictions to society." |  |  | Corresponding author. Joep J.J.M. Stohr; joep.stohr@gmail.com (J.J.J.M. Stohr).
Amy | Sherli | 241 | J. Stohr; V. F. Zwart; G. Goderski; A. Meijer; C. R. S. Nagel-Imming; M. F. Q. Kluytmans-van den Bergh; S. D. Pas; F. van den Oetelaar; M. Hellwich; K. H. Gan; A. Rietveld; J. J. Verweij; J. L. Murk; W. van den Bijllaardt; J. Kluytmans | Self-testing for the detection of SARS-CoV-2 infection with rapid antigen tests for people with suspected COVID-19 in the community | 2021-08-01 | https://doi.org/10.1016/j.cmi.2021.07.039 | [6 empty] | Nasopharyngeal swabs (NP) |  | Custom/In-house SARS-2 primers
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-01 | https://doi.org/10.1016/j.jiac.2021.10.024 |  | Japan | COVID-19 testing site/Screening location | April 20, 2021 - May 31, 2021 |  | 1127 | Nasopharyngeal swabs (NP) |  | Japanese National Institute of Infectious Diseases (NIID) |  | targets N and N2 genes | "In conclusion, the sensitivity of DIA tests was superior to that of visually analyzed antigen tests, but the rate of false-positive results increased with the introduction of a DIA device. When positive DIA results are recorded without a visually recognizable red line at the positive line location on the test cassette, additional RT-PCR evaluation should be performed due to possible false-positives." |  |  | Hiromichi Suzuki hsuzuki@md.tsukuba.ac.jp, hiromichisuzuki@hotmail.com (H. Suzuki),
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-01 | https://doi.org/10.12688/wellcomeopenres.16842.1 |  | United Kingdom | COVID-19 testing site/Screening location | Nov 26, 2020 - March 11, 2021 | 69.7% Female (82.5% Northumberland; 37.7% Yorkshire). Average age 38.0 yrs Northumberland and 68.7 years Yorkshire | 258 | Throat Swabs (TS) |  | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid) |  |  | "The Mologic Ag-RDT has replicated diagnostic accuracy in a variety of settings important for COVID-19 control (hospitals, ambulatory services, and drive-through centres) and on self-collected swab specimens." |  | Cites STARD guidelines; Northumberland and Yorkshire only | Emily R. Adams (emily.adams@lstmed.ac.uk)
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-01 | https://doi.org/10.12688/wellcomeopenres.16842.1 |  |  | Emergency Dept/Room |  |  |  | Nasal Swabs (NS) |  | Panther Fusion or Aptima SARS-CoV-2 assay (Hologic) |  | targets ORF1ab
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-01 | https://doi.org/10.12688/wellcomeopenres.16842.1 |  |  | Hospital - Outpatient |  |  |  | Nasopharyngeal swabs (NP)
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-01 | https://doi.org/10.1038/s41598-022-11284-8 |  | India | COVID-19 testing site/Screening location | June 19 - July 1, 2021 | mean age 30.6 phase I, 29.5 Phase II; 21.3% female Phase I, 17.9% in Phase II | 529 | Other | Oral swab | Covidsure Multiplex RT-PCR kit (Trivitron Healthcare Labsystems Diagnostics) |  |  | "To conclude, our study demonstrated moderate sensitivity of oral swab specimen for SARS-CoV-2 detection antigen detection in symptomatic individuals. Oral swab, was the preferred sampling by almost all participants in terms of convenience and comfort as compared to NP swab." |  |  | Arati Mane email: amane@nariindia.org
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-01 | https://doi.org/10.1038/s41598-022-11284-8 | [6 empty] | Nasopharyngeal swabs (NP)
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-01 | https://doi.org/10.1038/s41598-022-11284-8 | [6 empty] | Oropharyngeal swabs (OP)
Amy | Sherli | 301 | M. Maniscalco; P. Ambrosino; A. Ciullo; S. Fuschillo; V. Valente; C. Gaudiosi; D. Paris; R. Cobuccio; F. Stefanelli; A. Motta | A Rapid Antigen Detection Test to Diagnose SARS-CoV-2 Infection Using Exhaled Breath Condensate by A Modified Inflammacheck(®) Device | 2021-08-01 | https://doi.org/10.3390/s21175710 |  | Italy | Hospital - Outpatient | Not listed | Mean age 58.4 years; 39.0% female | 105 | Other | Breath | NeuMoDx SARS-CoV-2 Assay (Qiagen) |  | targets Nsp2 and N genes | "Our preliminary results demonstrate that Inflammacheck® enables the qualitative detection of SARS-CoV-2 in EBC samples in a reliable manner, with excellent overall, positive and negative percentage agreements and with a high κ value." |  |  | Mauro Maniscalco Correspondence: mauro.maniscalco@icsmaugeri.it
Amy | Sherli | 301 | M. Maniscalco; P. Ambrosino; A. Ciullo; S. Fuschillo; V. Valente; C. Gaudiosi; D. Paris; R. Cobuccio; F. Stefanelli; A. Motta | A Rapid Antigen Detection Test to Diagnose SARS-CoV-2 Infection Using Exhaled Breath Condensate by A Modified Inflammacheck(®) Device | 2021-08-01 | https://doi.org/10.3390/s21175710 |  |  | Hospital - Inpatient
Amy | Sherli | 243 | K. Surasi; K. J. Cummings; C. Hanson; M. K. Morris; M. Salas; D. Seftel; L. Ortiz; R. Thilakaratne; C. Stainken; D. A. Wadford | Effectiveness of Abbott BinaxNOW Rapid Antigen Test for Detection of SARS-CoV-2 Infections in Outbreak among Horse Racetrack Workers, California, USA | 2021-11-01 | https://doi.org/10.3201/eid2711.211449 |  | United States | Not described/unclear | Nov 25 - Dec 22 2020 | median age 52 years; 25.7% female | 769 | Nasal Swabs (NS) |  | TaqPath COVID-19 Combo kit (Applied Biosystems/ThermoFisher) |  |  | "Our results indicate that BinaxNOW performs better at identifying rRT-PCR–positive specimens with lower Ct (suggestive of higher viral loads) and positive viral cultures, although these factors are not precise proxies for infectiousness." |  | Setting is Horse Track workers | Krishna Surasi, email: okt3@cdc.gov
Amy | Sherli | 244 | H. Suzuki; Y. Akashi; D. Kato; Y. Takeuchi; Y. Kiyasu; N. Terada; Y. Kurihara; M. Kuwahara; S. Muramatsu; A. Ueda; S. Notake; K. Nakamura | Analytical performance of rapid antigen tests for the detection of SARS-CoV-2 during widespread circulation of the Omicron variant | 2022-05-01 | https://doi.org/10.1101/2022.05.17.22275034 |  | Japan | COVID-19 testing site/Screening location | Dec 28, 20221 -Feb 16, 2022 |  | 1073 | Nasopharyngeal swabs (NP) |  | Japanese National Institute of Infectious Diseases (NIID) |  | targets N and N2 genes | "In conclusion, the current study showed that the QuickNavi-COVID19 Ag test had a high diagnostic performance for the detection of SARS-CoV-2 Omicron sublineages BA.1 and BA.2 in nasopharyngeal samples." |  |  | Hiromichi Suzuki, MD, PhD, E-mail: hsuzuki@md.tsukuba.ac.jp
Amy | Sherli | 244 | H. Suzuki; Y. Akashi; D. Kato; Y. Takeuchi; Y. Kiyasu; N. Terada; Y. Kurihara; M. Kuwahara; S. Muramatsu; A. Ueda; S. Notake; K. Nakamura | Analytical performance of rapid antigen tests for the detection of SARS-CoV-2 during widespread circulation of the Omicron variant | 2022-05-01 | https://doi.org/10.1101/2022.05.17.22275034 | [8 empty] | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid)
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-01 | https://doi.org/10.1371/journal.pone.0259527 |  | Austria | Emergency Dept/Room | Oct 30, 2020 - Dec 13, 2020 | ED median age 58; PHC median age 37; ED 49.0% female; PHC 64.1% female | 541 | Other | Swab - type not specified | Specific Test Not Described |  |  | "This prospective study demonstrated a performance of the SARS-CoV-2 Rapid Antigen Test (Roche Diagnostics) with an overall sensitivity of 80.3% compared to rt-PCR, which, in case of a negative result, needs to be interpreted together with the duration of the disease at the time of testing, the viral load, and likely the diligence of the generation of the specimen." |  |  | * Rainer ThellI; rainer.thell@meduniwien.ac.at (RT); Marton Szell marton.szell@gesundheitsverbund.at (MS)
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-01 | https://doi.org/10.1371/journal.pone.0259527 |  |  | Primary Care Location
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-01 | https://doi.org/10.1016/j.ijid.2021.10.027 |  | Mexico | Hospital - Inpatient | Not listed | Mean age 46.6 years; 56.3% women | 1060 | Nasopharyngeal swabs (NP) |  | Custom/In-house SARS-2 primers |  | Targest N, E, ORF and RdRp genes | "As expected, sensitivity was highly associated with time from the onset of symptoms and with viral load. The low sensitivity and high specificity observed in the Panbio test make it necessary to confirm all negative results, especially in the presence of symptoms compatible with COVID-19 and in settings with a high community infection rate." |  | Some data came from the supplemental file which is now uploaded into the re-run full text folder | Correspondence: Dr. Rogelio Pérez-Padilla, E-mail address: perezpad@gmail.com (Dr.J.R. Pérez-Padilla).
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-01 | https://doi.org/10.1016/j.ijid.2021.10.027 |  |  | Emergency Dept/Room
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-01 | https://doi.org/10.1016/j.ijid.2021.10.027 |  |  | COVID-19 testing site/Screening location
Amy | Sherli | 251 | V. Tonelotto; A. Davini; L. Cardarelli; M. Calderone; P. Marin | Efficacy of Fluorecare SARS-CoV-2 Spike Protein Test Kit for SARS-CoV-2 detection in nasopharyngeal samples of 121 individuals working in a manufacturing company | 2022-01-01 | https://doi.org/10.1371/journal.pone.0262174 |  | Italy | Not described/unclear | Mar 16 - Apr 2 2021 |  | 253 | Nasopharyngeal swabs (NP) |  | PerkinElmer SARS-CoV-2 Real-time RT-PCR Assay |  | targets N and ORF1ab | "This real-life scenario gave us the possibility not only to evaluate the clinical performance of Fluorecare SARS-CoV-2 Spike Protein Test Kit, a new antigen assay, but also to experience the success of a prompt and clever strategy on preserving public health and work activities." |  | Samples were from leftover diagnostic samples at a lab and worker screening at a local manufacturer | Valentina Tonelotto - valentina.tonelotto@gmail.com
Amy | Sherli | 251 | V. Tonelotto; A. Davini; L. Cardarelli; M. Calderone; P. Marin | Efficacy of Fluorecare SARS-CoV-2 Spike Protein Test Kit for SARS-CoV-2 detection in nasopharyngeal samples of 121 individuals working in a manufacturing company | 2022-01-01 | https://doi.org/10.1371/journal.pone.0262174 | [8 empty] | REALQUALITY RQ-SARS-CoV-2 kit (AB Analitica) |  | targets RdRp and S genes
Amy | Sherli | 251 | V. Tonelotto; A. Davini; L. Cardarelli; M. Calderone; P. Marin | Efficacy of Fluorecare SARS-CoV-2 Spike Protein Test Kit for SARS-CoV-2 detection in nasopharyngeal samples of 121 individuals working in a manufacturing company | 2022-01-01 | https://doi.org/10.1371/journal.pone.0262174 | [8 empty] | MutaPLEX Coronavirus Real-time-RT-PCR kit (Immundiagnostik AG) |  | targets RdRp, S, E and N genes
Amy | Sherli | 252 | G. Turcato; A. Zaboli; N. Pfeifer; S. Sibilio; G. Tezza; A. Bonora; L. Ciccariello; D. Ausserhofer | Rapid antigen test to identify COVID-19 infected patients with and without symptoms admitted to the Emergency Department | 2022-01-01 | https://doi.org/10.1016/j.ajem.2021.10.022 |  | Italy | Emergency Dept/Room | July 1, 2020 - Dec 10, 2020 | 48.9% female; median age 69 years | 3899 | Nasopharyngeal swabs (NP) |  | Xpert Xpress SARS-CoV-2/GeneXpert (Cepheid) |  |  | "A negative antigen test in these symptomatic patients should however be followed by RT-PCR testing to confirm the absence of disease." |  |  | Gianni Turcato, E-mail address: gianni.turcato@yahoo.it (G. Turcato).
Amy | Sherli | 253 | E. Van Honacker; K. Van Vaerenbergh; A. Boel; H. De Beenhouwer; I. Leroux-Roels; L. Cattoir | COMPARISON OF FIVE SARS-COV-2 RAPID ANTIGEN TESTS IN A HOSPITAL SETTING AND PERFORMANCE OF ONE ANTIGEN ASSAY IN ROUTINE PRACTICE. A USEFUL TOOL TO GUIDE ISOLATION PRECAUTIONS? | 2021-03-01 | https://doi.org/10.1016/j.jhin.2021.03.021 |  | Belgium | Emergency Dept/Room | Nov 5, 2020 - Mar 14, 2021 | 54.4% female; 68 years median age | 4195 | Nasopharyngeal swabs (NP) |  | Custom/In-house SARS-2 primers |  |  | "Exclusive use of a RAD test cannot replace rRT-PCR in a hospital setting because a negative RAD test result cannot rule out a diagnosis of COVID-19." |  | Implementation study only. | E. Van Honacker E-mail address: eveline.vanhonacker@ugent.be (E. Van Honacker).
Amy | Sherli | 254 | T. von Ahnen; M. von Ahnen; U. Wirth; H. M. Schardey; S. Herdtle | Evaluation of a rapid-antigen test for COVID-19 in an asymptomatic collective : A prospective study | 2021-09-01 | https://doi.org/10.1007/s10354-021-00883-1 |  | Germany | Hospital - Inpatient | 2021-01-01 | Mean age 41.5 years; 76% female | 919 | Nasopharyngeal swabs (NP) |  | Eurofins ViroBOAR kit (Eurofins Genomic GmbH) |  |  | "We therefore recommend repeating the tests regularly over several days in order to identify as early as possible all asymptomatic persons in a sensitive facility, if an RT-PCR test is not available." |  |  | Dr. T. von Ahnen; vonahnen2002@yahoo.de
Amy | Sherli | 258 | C. Wertenauer; G. Brenner Michael; A. Dressel; C. Pfeifer; U. Hauser; E. Wieland; C. Mayer; C. Mutschmann; M. Roskos; H. J. Wertenauer; A. P. Moissl; S. Lorkowski; W. März | Diagnostic Performance of Rapid Antigen Testing for SARS-CoV-2: The COVid-19 AntiGen (COVAG) study | 2022-03-01 | 10.3389/fmed.2022.774550 |  | Germany | COVID-19 testing site/Screening location | Feb 1, 2021 - Mar 31, 2021 | Mean age 39.9 Years; 54.7% female | 2215 | Nasopharyngeal swabs (NP) |  | RIDAGENE SARS-CoV-2 (R-Bio-pharm) |  | targets E gene | "Of note, the significant negative association between positivity of RDT results and presence of the alpha variant remains statistically significant when we considered age, sex, Ct-value, reason for testing, presence of comorbidities and presence of clinical symptoms in the same logistic regression model. This implies that the RDTs are less sensitive toward the alpha variant even adjusting for other conditions affecting sensitivity." |  |  | Christoph Wertenauer christoph.wertenauer@uni-heidelberg.de
Amy | Sherli | 258 | C. Wertenauer; G. Brenner Michael; A. Dressel; C. Pfeifer; U. Hauser; E. Wieland; C. Mayer; C. Mutschmann; M. Roskos; H. J. Wertenauer; A. P. Moissl; S. Lorkowski; W. März | Diagnostic Performance of Rapid Antigen Testing for SARS-CoV-2: The COVid-19 AntiGen (COVAG) study | 2022-03-01 | 10.3389/fmed.2022.774550 | [6 empty] | Oropharyngeal swabs (OP)

## Master - Study Data
Data Extractor | Data Validated by | Article # | Article First Author | Identifier | Article Title | Month Year | Index test description - Use one line per test and subgroup. For example, If Standard Q and BinaxNOW were both used in asymptomatic and symptomatic subjects, there would be 4 lines: (1) Standard Q Asymptomatic; (2) Standard Q Symptomatic; (3) BinaxNOW asymptomatic; and (4) BinaxNOW symptomatic | Index test if Other | Index test comments (if any) | Subgroups (if any; include Overall) | True Positive (TP) | False Positive (FP) | True Negative (TN) | False Negative (FN) | Sensitivity (TP/(TP+FN)) | Sensitivity 95% CI (low,high) | Specificity (TN/(TN+FP)) | Specificity 95% CI (low,high) | Positive Predictive Value PPV (TP/(TP+FP)) | Negative Predictive Values NPV (TN/(FN+TN) |  | Calculated Sensitivity Based on entered #'s | Calculated Specificity Based on entered #'s | Calculated PPV Based on entered #'s | Calculated NPV Based on entered #'s
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | PanBio (Abbott) |  |  | Overall | 10 | 0 | 10 | 3 | 76.92 | 46.19,94.96 | 100 |  | 100 | 92.1 |  | 76.9231 | 100 | 100 | 76.9231
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | PanBio (Abbott) |  |  | Ct ≤25 |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | PanBio (Abbott) |  |  | Symptom onset ≤6 days |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Flowflex COVID-19 Antigen test (ACON Labs) |  |  | Overall | 15 | 0 | 10 | 0 | 100 | 78.20,100 | 100 |  | 100 | 100 |  | 100 | 100 | 100 | 100
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Flowflex COVID-19 Antigen test (ACON Labs) |  |  | Ct ≤25 |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Flowflex COVID-19 Antigen test (ACON Labs) |  |  | Symptom onset ≤6 days |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | AMP Rapid Test SARS-CoV-2 Ag (AMP Diagnostics) |  |  | Overall | 6 | 0 | 8 | 1 | 85.71 | 42.13,99.64 | 100 |  | 100 | 94.98 |  | 85.7143 | 100 | 100 | 88.8889
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | AMP Rapid Test SARS-CoV-2 Ag (AMP Diagnostics) |  |  | Ct ≤25 |  |  |  |  | 85.7 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | AMP Rapid Test SARS-CoV-2 Ag (AMP Diagnostics) |  |  | Symptom onset ≤6 days |  |  |  |  | 75 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | COVID-19 Antigen Rapid Test (Assut Europe) |  |  | Overall | 5 | 0 | 24 | 2 | 71.43 | 29.04,96.33 | 100 |  | 100 | 90.44 |  | 71.4286 | 100 | 100 | 92.3077
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | COVID-19 Antigen Rapid Test (Assut Europe) |  |  | Ct ≤25 |  |  |  |  | 83.3 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | COVID-19 Antigen Rapid Test (Assut Europe) |  |  | Symptom onset ≤6 days |  |  |  |  | 80 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Novel Corona Virus (SARS-CoV-2) Ag Rapid Test kit (Bioperfectus) |  |  | Overall | 8 | 0 | 11 | 2 | 80 | 44.39,94.78 | 100 |  | 100 | 93.11 |  | 80 | 100 | 100 | 84.6154
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Novel Corona Virus (SARS-CoV-2) Ag Rapid Test kit (Bioperfectus) |  |  | Ct ≤25 |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Novel Corona Virus (SARS-CoV-2) Ag Rapid Test kit (Bioperfectus) |  |  | Symptom onset ≤6 days |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | CerTest SARS-CoV-2 (Certest Biotech) |  |  | Overall | 5 | 0 | 9 | 3 | 62.5 | 24.49,91.48 | 100 |  | 100 | 87.82 |  | 62.5 | 100 | 100 | 75
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | CerTest SARS-CoV-2 (Certest Biotech) |  |  | Ct ≤25 |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | CerTest SARS-CoV-2 (Certest Biotech) |  |  | Symptom onset ≤6 days |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Espline SARS-CoV-2 (Fujirebio) |  |  | Overall | 8 | 0 | 7 | 2 | 80 | 44.39,97.48 | 100 |  | 100 | 93.11 |  | 80 | 100 | 100 | 77.7778
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Espline SARS-CoV-2 (Fujirebio) |  |  | Ct ≤25 |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Espline SARS-CoV-2 (Fujirebio) |  |  | Symptom onset ≤6 days |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Fluorecare (Colloidal Gold/Fluorescent) SARS-CoV-2 Spike Protein Test kit (Shenzen Microprofit) |  |  | Overall | 11 | 0 | 9 | 1 | 91.67 | 61.52,99.79 | 100 |  | 100 | 97.01 |  | 91.6667 | 100 | 100 | 90
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Fluorecare (Colloidal Gold/Fluorescent) SARS-CoV-2 Spike Protein Test kit (Shenzen Microprofit) |  |  | Ct ≤25 |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Fluorecare (Colloidal Gold/Fluorescent) SARS-CoV-2 Spike Protein Test kit (Shenzen Microprofit) |  |  | Symptom onset ≤6 days |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Other | Mixed (Fluorecare, Espline, Rapigen, Assut, Oreint Gene, AMP, ACON, Abbott, Certest Biotec, Bioperfects) |  | Overall | 83 | 0 | 123 | 25 | 76.85 | 67.75,84.43 | 100 |  | 100 | 92.1 |  | 76.8519 | 100 | 100 | 83.1081
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Other | Mixed (Fluorecare, Espline, Rapigen, Assut, Oreint Gene, AMP, ACON, Abbott, Certest Biotec, Bioperfects) |  | Ct ≤25 |  |  |  |  | 96.4 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Other | Mixed (Fluorecare, Espline, Rapigen, Assut, Oreint Gene, AMP, ACON, Abbott, Certest Biotec, Bioperfects) |  | Symptom onset ≤6 days |  |  |  |  | 96.43 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Orient Gene Rapid Covid-19 (Antigen) Self-Test (Orient Gene) |  |  | Overall | 5 | 0 | 12 | 5 | 50 | 18.71,81.29 | 100 |  | 100 | 84.39 |  | 50 | 100 | 100 | 70.5882
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Orient Gene Rapid Covid-19 (Antigen) Self-Test (Orient Gene) |  |  | Ct ≤25 |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | Orient Gene Rapid Covid-19 (Antigen) Self-Test (Orient Gene) |  |  | Symptom onset ≤6 days |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | RapiGen (BioCredit) |  |  | Overall | 10 | 0 | 23 | 6 | 62.5 | 35.4,84.80 | 100 |  | 100 | 87.82 |  | 62.5 | 100 | 100 | 79.3103
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | RapiGen (BioCredit) |  |  | Ct ≤25 |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 7 | Abusrewil Z. | Abusrewil, et. al. Dec 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays | 2021-05-01 | RapiGen (BioCredit) |  |  | Symptom onset ≤6 days |  |  |  |  | 100 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Tim | Sherli | 10 | Y. Akashi; M. Horie; Y. Takeuchi; K. Togashi; Y. Adachi; A. Ueda; S. Notake; K. Nakamura; N. Terada; Y. Kurihara; Y. Kiyasu; H. Suzuki | Akashi, et. al. Feb 2022 | A prospective clinical evaluation of the diagnostic accuracy of the SARS-CoV-2 rapid antigen test using anterior nasal samples | 2022-02-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall | 80 | 0 | 690 | 30 | 72.7 | 63.4,80.8 | 100 | 99.5,100 | 100 | 95.8 | PPV 95%CI 95.5,100 NPV 95%CI 94.1,97.1
Tim | Sherli | 10 | Y. Akashi; M. Horie; Y. Takeuchi; K. Togashi; Y. Adachi; A. Ueda; S. Notake; K. Nakamura; N. Terada; Y. Kurihara; Y. Kiyasu; H. Suzuki | Akashi, et. al. Feb 2022 | A prospective clinical evaluation of the diagnostic accuracy of the SARS-CoV-2 rapid antigen test using anterior nasal samples | 2022-02-02 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptomatic | 50 | 0 | 274 | 9 | 84.7 | 73.0,92.8 | 100 | 98.7,100 | 100 | 96.8 | PPV 95%CI 92.9,100 NPV 95%CI 94,98.5
Tim | Sherli | 10 | Y. Akashi; M. Horie; Y. Takeuchi; K. Togashi; Y. Adachi; A. Ueda; S. Notake; K. Nakamura; N. Terada; Y. Kurihara; Y. Kiyasu; H. Suzuki | Akashi, et. al. Feb 2022 | A prospective clinical evaluation of the diagnostic accuracy of the SARS-CoV-2 rapid antigen test using anterior nasal samples | 2022-02-03 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Asymptomatic | 30 | 0 | 416 | 21 | 58.8 | 44.2,72.4 | 100 | 91.1,100 | 100 | 95.2 | PPV 95%CI 88.4,100 NPV 95%CI 92.7,97
Amy | Sherli | 12 | S. Al-Alawi; H. Al-Hinai; N. Al-Kindi; M. Al-Rashidi; H. Al-Kindi; I. Al-Shukri; A. Al-Rashdi; S. Jose; A. Al-Jardani | Al-Alawi, et. al. Jan 2021 | Evaluation of four rapid antigen tests for detection of SARS-CoV-2 virus |  | Standard Q COVID-19 Ag (SD Biosensor) |  | 66 samples | Overall |  |  |  |  | 65.8 | 48.65,80.37 | 100 | 87.66,100 | 100 | 96.7 | NPV CI95 = 94.89-97.82 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  | 87 samples | Overall |  |  |  |  | 69.8 | 55.66,81.66 | 94.1 | 80.32,99.28 | 54.6 | 96.9 | PPV CI95 = 26.35-82.35 NPV CI95 = 95.30-97.91 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | RapiGen (BioCredit) |  | 75 samples | Overall |  |  |  |  | 64 | 49.19,77.08 | 100 | 86.28,100 | 100 | 96.5 | NPV CI95 = 94.99-97.54 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  | 78 samples | Overall |  |  |  |  | 64.3 | 50.36,76.64 | 100 | 84.56,100 | 100 | 96.5 | NPV CI95 = 95.11-97.52 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT <= 35 |  |  |  |  | 71 | 51.96,58.78 |  |  | 100 | 97.1 | NPV CI95 = 95.15-98.33 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT <= 30 |  |  |  |  | 85 | 62.11,96.79 |  |  | 100 | 98.5 | NPV CI95 = 95.86-99.47 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT <= 25 |  |  |  |  | 92.3 | 63.97,99.81 |  |  | 100 | 99.2 | NPV CI95 = 95.13-99.88 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT <= 35 |  |  |  |  | 75.5 | 61.13,99.28 |  |  | 56.5 | 97.4 | PPV CI95 = 25.14-83.44 NPV CI95 = 95.84-98.42 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT <= 30 |  |  |  |  | 91.9 | 78.09,98.30 |  |  | 61.3 | 99.1 | PPV CI95 = 29.14-85.90 NPV CI95 = 97.47-99.71 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT <= 25 |  |  |  |  | 100 | 85.75,100 |  |  | 63.3 | 100 | PPV CI95 = 30.99-86.86 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | RapiGen (BioCredit) |  |  | CT <= 35 |  |  |  |  | 71.1 | 55.69,83.63 |  |  | 100 | 97.2 | NPV CI95 = 95.58-98.18 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | RapiGen (BioCredit) |  |  | CT <= 30 |  |  |  |  | 85.7 | 69.74,95.19 |  |  | 100 | 98.6 | NPV CI95 = 96.84-99.36 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | RapiGen (BioCredit) |  |  | CT <= 25 |  |  |  |  | 92.3 | 74.87,99.05 |  |  | 100 | 99.2 | NPV CI95 = 98.13-99.79 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | CT <= 35 |  |  |  |  | 66 | 51.73,78.48 |  |  | 100 | 96.7 | NPV CI95 = 95.23-97.69 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 7 | Abusrewil Z. | Al-Alawi, et. al. Jan 2021 | Time scale performance of rapid antigen testing for SARS-COV-2: evaluation of ten rapid antigen assays |  | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | CT <= 20 |  |  |  |  | 100 | 89.11,100 |  |  | 100 | 100 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 13 | S. Aleem; N. Zahoor; A. Jeelani; S. M. SalimKhan | Aleem, et. al. Jan 2022 | Diagnostic Accuracy of STANDARD QCOVID-19 Antigen Detection Kit in Comparison with RT-PCR Assay using Nasopharyngeal Samples in India |  | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 43 | 3 | 391 | 36 | 54.43 | 42.83,65.69 | 99.24 | 97.79,99.84 | 93.48 | 91.57 | PPV 95% [82.02,97.83] NPV 95% [89.51,93.25 | 54.4304 | 99.2386 | 93.4783 | 91.5691
Chathu | Sherli | 14 | M. Alghounaim; H. Bastaki; F. Bin Essa; H. Motlagh; S. Al-Sabah | Alghounaim, et. al. Dec 2021 | The Performance of Two Rapid Antigen Tests During Population-Level Screening for SARS-CoV-2 Infection |  | Liaison |  |  | Overall | 26 | 1 | 836 | 34 | 43.3 | 30.6–56.8 | 99.9 | 99.3–100 | 96.3 | 96.1 | PPV 95%CI 81,99.9 NPV 95% CI 94.6,97.3 | 43.3333 | 99.8805 | 96.2963 | 96.092
Chathu | Sherli | 14 | M. Alghounaim; H. Bastaki; F. Bin Essa; H. Motlagh; S. Al-Sabah | Alghounaim, et. al. Dec 2021 | The Performance of Two Rapid Antigen Tests During Population-Level Screening for SARS-CoV-2 Infection |  | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 19 | 11 | 898 | 43 | 30.6 | 19.6–43.7 | 98.8 | 97.8–99.4 | 63.3 | 95.4 | PPV 95%CI 43.9,80.1 NPV 95% CI 93.9,96.7 | 30.6452 | 98.7899 | 63.3333 | 95.4304
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-01 | BinaxNOW (Abbott) |  |  | Asymptomatic--Oral Fluid PCR | 374 | 81 | 6776 | 480 | 43.8 | 40.4,47.2 | 98.8 | 98.5,99.1
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-02 | BinaxNOW (Abbott) |  |  | Symptomatic--Oral Fluid PCR | 1024 | 101 | 6556 | 948 | 51.9 | 49.7,54.2 | 98.5 | 98.2,98.8
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-03 | BinaxNOW (Abbott) |  |  | Unknown symptoms--Oral Fluid PCR | 26 | 4 | 307 | 32 | 44.8 | 31.7,58.5 | 98.7 | 96.7,99.7
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-04 | BinaxNOW (Abbott) |  |  | Overall--Oral Fluid PCR | 1424 | 186 | 13639 | 1460 | 49.37 | 47.5,51.2 | 98.7 | 98.5,98.8
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-05 | BinaxNOW (Abbott) |  |  | Asymptomatic--Nasopharyngeal PCR | 21 | 1 | 449 | 37 | 36.2 | 24,49.9 | 99.8 | 98.8,100
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-06 | BinaxNOW (Abbott) |  |  | Symptomatic--Nasopharyngeal PCR | 38 | 1 | 191 | 32 | 54.3 | 41.9,66.3 | 99.5 | 97.1,100
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-07 | BinaxNOW (Abbott) |  |  | Unknown symptoms--Nasopharyngeal PCR | 0 | 0 | 3 | 0 | 0 | -- | 100 | 29.2,100
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-08 | BinaxNOW (Abbott) |  |  | Overall--Nasopharyngeal Fluid PCR | 59 | 2 | 643 | 69 | 46.1 | 37.3,55.1 | 99.7 | 98.9,100
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-09 | BinaxNOW (Abbott) |  |  | Asymptomatic--Anterior Nares PCR | 25 | 1 | 421 | 21 | 54.4 | 39,69.1 | 99.8 | 98.7,100
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-10 | BinaxNOW (Abbott) |  |  | Symptomatic--Anterior Nares PCR | 31 | 0 | 276 | 34 | 47.7 | 35.2,60.5 | 100 | 98.7,100
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-11 | BinaxNOW (Abbott) |  |  | Uknown symptoms--Anterior Nares PCR | 11 | 0 | 136 | 19 | 36.7 | 19.9,56.1 | 100 | 97.3,100
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-12 | BinaxNOW (Abbott) |  |  | Overall--Anterior Nares PCR | 67 | 1 | 833 | 74 | 47.5 | 39.1,56.1 | 100 | 99.3,100
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-13 | BinaxNOW (Abbott) |  |  | Asymptomatic--all sample types PCR | 420 | 83 | 7646 | 538 | 43.8 | 40.7,47.1 | 98.9 | 98.7,99.1
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-14 | BinaxNOW (Abbott) |  |  | Symptomatic--all sample types PCR | 1093 | 102 | 7023 | 1014 | 51.9 | 49.7,54 | 98.6 | 98.3,98.8
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-15 | BinaxNOW (Abbott) |  |  | Unknown symptoms --all sample types PCR | 37 | 4 | 446 | 51 | 42.1 | 31.6,53.1 | 99.1 | 97.7,99.8
Sarina | Sherli | 15 | L. T. Allan-Blitz; J. D. Klausner | Allan-Blitz, et. al. Sep 2021 | A Real-World Comparison of SARS-CoV-2 Rapid Antigen Testing versus PCR Testing in Florida | 2021-09-16 | BinaxNOW (Abbott) |  |  | Overall--all sample types PCR | 1550 | 189 | 15115 | 1603 | 49.2 | 47.4,50.9 | 98.8 | 98.6,98.9
Ellyn | Sherli | 16 | Amer R. | Amer, et. al. Oct 2021 | Diagnostic performance of rapid antigen test for COVID-19 and the effect of viral load, sampling time, subject's clinical and laboratory parameters on test accuracy | 2021-06-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 54 | 5 | 9 | 15 | 78.2 | 67,86 | 64.2 | 83,99 | 91.5 | 37.5 |  | 78.2609 | 64.2857 | 91.5254 | 37.5
Karen | Sherli | 20 | R. Amer R. S. M. G. O. E.-D. N. A. A. A. A. S. W. A. A. W. M. A. | Amer, et. al.Mar 2022 | Diagnostic performance of rapid antigen test for COVID-19 and the effect of viral load, sampling time, subject's clinical and laboratory parameters on test accuracy (preprint) | 2022-03-01 | Standard Q COVID-19 Ag (SD Biosensor) |  | data in table 2 | Overall | 54 | 5 | 9 | 15 | 78.2 | 67, 86 | 64.2 | 38,83 | 91.5 | 37.5
Karen | Sherli | 20 | R. Amer R. S. M. G. O. E.-D. N. A. A. A. A. S. W. A. A. W. M. A. | Amer, et. al.Mar 2022 | Diagnostic performance of rapid antigen test for COVID-19 and the effect of viral load, sampling time, subject's clinical and laboratory parameters on test accuracy (preprint) | 2022-03-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct < 29 | 28 | 28 | 22 | 1 | 96.5 | 83,99 | 44 | 31,57 | 50 | 95.7
Karen | Sherli | 20 | R. Amer R. S. M. G. O. E.-D. N. A. A. A. A. S. W. A. A. W. M. A. | Amer, et. al.Mar 2022 | Diagnostic performance of rapid antigen test for COVID-19 and the effect of viral load, sampling time, subject's clinical and laboratory parameters on test accuracy (preprint) | 2022-03-03 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | 29< Ct < 36 | 21 | 38 | 16 | 8 | 72.4 | 54,85 | 29.6 | 19,42 | 35.6 | 66.7
Karen | Sherli | 20 | R. Amer R. S. M. G. O. E.-D. N. A. A. A. A. S. W. A. A. W. M. A. | Amer, et. al.Mar 2022 | Diagnostic performance of rapid antigen test for COVID-19 and the effect of viral load, sampling time, subject's clinical and laboratory parameters on test accuracy (preprint) | 2022-03-04 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | 37 < Ct < 39 | 5 | 54 | 17 | 6 | 45.4 | 21,71 | 23.9 | 15,35 | 8.5 | 73.9
Yati | Sherli | 17 | Anastasiou O. | Anastasiou, et. al. Apr 2021 | Fast Detection of SARS-CoV-2 RNA Directly from Respiratory Samples Using a Loop-Mediated Isothermal Amplification (LAMP) Test | 2021-04-01 | Custom/Novel/In-house |  | New England BioLabs | Overall | 55 | 0 | 47 | 25 | 68.8 | 57.3,78.4 | 100 | 90.6,100 | 100 | 65.3 |  | 68.75 | 100 | 100 | 65.2778
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Avgoulea, et. al. Apr 2022 | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital |  | Custom/Novel/In-house |  | NG-Test from NG-Biotech Laboratories, France | NP sample | 108 |  |  | 26 | 81 | 73,87 | 99 | 95100 | 99 | 83 | PPV 95%CI 94-100 NPV 95% CI 76-88 | 80.597 | #DIV/0! | 100 | 0
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Avgoulea, et. al. Apr 2022 | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital |  | Custom/Novel/In-house |  | NG-Test from NG-Biotech Laboratories, France | OP Sample | 68 |  |  | 66 | 51 | 42,59 | 100 | 96-100 | 100 | 66 | PPV 95% CI 93-100 NPV 95% CI 59-73 | 50.7463 | #DIV/0! | 100 | 0
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Avgoulea, et. al. Apr 2022 | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital |  | Custom/Novel/In-house |  | NG-Test from NG-Biotech Laboratories, France | NP sample - High viral Load (CT=< 25) |  |  |  |  | 99 | 92100 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Avgoulea, et. al. Apr 2022 | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital |  | Custom/Novel/In-house |  | NG-Test from NG-Biotech Laboratories, France | NP sample - intermediate viral load (25<CT=<30) |  |  |  |  | 72 | 55,85 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Avgoulea, et. al. Apr 2022 | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital |  | Custom/Novel/In-house |  | NG-Test from NG-Biotech Laboratories, France | NP sample - low viral load (30<CT=<35) |  |  |  |  | 29 | 12,52 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Avgoulea, et. al. Apr 2022 | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital |  | Custom/Novel/In-house |  | NG-Test from NG-Biotech Laboratories, France | OP sample - High viral Load (CT=< 25) |  |  |  |  | 71 | 60,81 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Avgoulea, et. al. Apr 2022 | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital |  | Custom/Novel/In-house |  | NG-Test from NG-Biotech Laboratories, France | OP sample - intermediate viral load (25<CT=<30) |  |  |  |  | 31 | 17,48 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 22 | K. Avgoulea; M. I. Beredaki; S. Vourli; M. Siopi; N. Siafakas; S. Pournaras | Avgoulea, et. al. Apr 2022 | Field Evaluation of the New Rapid NG-Test(®) SARS-CoV-2 Ag for Diagnosis of COVID-19 in the Emergency Department of an Academic Referral Hospital |  | Custom/Novel/In-house |  | NG-Test from NG-Biotech Laboratories, France | OP sample - low viral load (30<CT=<35) |  |  |  |  | 10 | 2,32 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Yati | Sherli | 23 | C. M. Bachman; B. D. Grant; C. E. Anderson; L. F. Alonzo; S. Garing; S. A. Byrnes; R. Rivera; S. Burkot; A. Ball; J. W. Stafford; W. Wang; D. Banik; M. D. Keller; D. M. Cate; K. P. Nichols; B. H. Weigl; P. Dewan | Bachman, et. al. Aug 2021 | Clinical validation of an open-access SARS-COV-2 antigen detection lateral flow assay, compared to commercially available assays | 2021-08-01 | Custom/Novel/In-house |  | open-access lateral flow assay (OA-LFA) | PCR collected by NP | 75 | 2 | 57 | 33 | 69 | 60,78 | 97 | 88100
Yati | Sherli | 23 | C. M. Bachman; B. D. Grant; C. E. Anderson; L. F. Alonzo; S. Garing; S. A. Byrnes; R. Rivera; S. Burkot; A. Ball; J. W. Stafford; W. Wang; D. Banik; M. D. Keller; D. M. Cate; K. P. Nichols; B. H. Weigl; P. Dewan | Bachman, et. al. Aug 2021 | Clinical validation of an open-access SARS-COV-2 antigen detection lateral flow assay, compared to commercially available assays | 2021-08-02 | Custom/Novel/In-house |  | open-access lateral flow assay (OA-LFA) | PCR collected by NS | 75 | 2 | 75 | 15 | 83 | 74,90 | 97 | 91100
Yati | Sherli | 23 | C. M. Bachman; B. D. Grant; C. E. Anderson; L. F. Alonzo; S. Garing; S. A. Byrnes; R. Rivera; S. Burkot; A. Ball; J. W. Stafford; W. Wang; D. Banik; M. D. Keller; D. M. Cate; K. P. Nichols; B. H. Weigl; P. Dewan | Bachman, et. al. Aug 2021 | Clinical validation of an open-access SARS-COV-2 antigen detection lateral flow assay, compared to commercially available assays | 2021-08-03 | BinaxNOW (Abbott) |  |  | PCR collected by NP | 89 | 0 | 60 | 20 | 82 | 73,88 | 100 | 94100
Yati | Sherli | 23 | C. M. Bachman; B. D. Grant; C. E. Anderson; L. F. Alonzo; S. Garing; S. A. Byrnes; R. Rivera; S. Burkot; A. Ball; J. W. Stafford; W. Wang; D. Banik; M. D. Keller; D. M. Cate; K. P. Nichols; B. H. Weigl; P. Dewan | Bachman, et. al. Aug 2021 | Clinical validation of an open-access SARS-COV-2 antigen detection lateral flow assay, compared to commercially available assays | 2021-08-04 | BinaxNOW (Abbott) |  |  | PCR collected by NS | 84 | 5 | 72 | 8 | 91 | 84,96 | 94 | 85,98
Yati | Sherli | 23 | C. M. Bachman; B. D. Grant; C. E. Anderson; L. F. Alonzo; S. Garing; S. A. Byrnes; R. Rivera; S. Burkot; A. Ball; J. W. Stafford; W. Wang; D. Banik; M. D. Keller; D. M. Cate; K. P. Nichols; B. H. Weigl; P. Dewan | Bachman, et. al. Aug 2021 | Clinical validation of an open-access SARS-COV-2 antigen detection lateral flow assay, compared to commercially available assays | 2021-08-05 | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | PCR collected by NP | 81 | 1 | 59 | 29 | 74 | 64,82 | 98 | 91100
Yati | Sherli | 23 | C. M. Bachman; B. D. Grant; C. E. Anderson; L. F. Alonzo; S. Garing; S. A. Byrnes; R. Rivera; S. Burkot; A. Ball; J. W. Stafford; W. Wang; D. Banik; M. D. Keller; D. M. Cate; K. P. Nichols; B. H. Weigl; P. Dewan | Bachman, et. al. Aug 2021 | Clinical validation of an open-access SARS-COV-2 antigen detection lateral flow assay, compared to commercially available assays | 2021-08-06 | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | PCR collected by NS | 79 | 3 | 75 | 13 | 86 | 77,92 | 96 | 89,99
Terry | Sherli | 28 | Basso D. | Basso, et. al. Feb 2021 | Salivary SARS-CoV-2 antigen rapid detection: A prospective cohort study | 2021-02-01 | Espline SARS-CoV-2 (Fujirebio) |  |  | Overall |  |  |  |  | 48 |  | 100 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 28 | Basso D. | Basso, et. al. Feb 2021 | Salivary SARS-CoV-2 antigen rapid detection: A prospective cohort study | 2021-02-01 | Other | Mixed (All tests) |  | Overall |  |  |  |  | 81.6 | 71.0,89.5 | 93.8 | 86.2,98.0 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry-finished | Sherli | 31 | K. A. Bond; B. Smith; E. Gardiner; K. C. Liew; E. Williams; N. Walsham; M. Putland; D. A. Williamson | Bond, et. al. May 2022 | Utility of SARS-CoV-2 rapid antigen testing for patient triage in the emergency department: A clinical implementation study in Melbourne, Australia | 2022 | PanBio (Abbott) |  |  | Overall | 206 | 0 | 1489 | 67 | 75.5 | 69.9,80.4 | 100 | 99.8,100 | 100 | 95.7 | 95%CI PPV 98.6,100; 95%CI NPV 94.6,96.7
Terry-finished | Sherli | 31 | K. A. Bond; B. Smith; E. Gardiner; K. C. Liew; E. Williams; N. Walsham; M. Putland; D. A. Williamson | Bond, et. al. May 2022 | Utility of SARS-CoV-2 rapid antigen testing for patient triage in the emergency department: A clinical implementation study in Melbourne, Australia | 2022 | PanBio (Abbott) |  |  | High Risk (Symptoms/epidemiology) | 3 | 0 | 17 | 0 | 100 | 29.2,100 | 100 | 83.8,100 | 100 | 100 | 95%CI PPV 98.6,100; 95%CI NPV 94.6,96.7
Terry-finished | Sherli | 31 | K. A. Bond; B. Smith; E. Gardiner; K. C. Liew; E. Williams; N. Walsham; M. Putland; D. A. Williamson | Bond, et. al. May 2022 | Utility of SARS-CoV-2 rapid antigen testing for patient triage in the emergency department: A clinical implementation study in Melbourne, Australia | 2022 | PanBio (Abbott) |  |  | Low Risk (symptoms/no epidemiology) | 89 | 0 | 937 | 29 | 75.4 | 66.7,82.9 | 100 | 99.7,100 | 100 | 97 | 95%CI PPV 96.7,100; 95%CI NPV 95.7,98
Terry-finished | Sherli | 31 | K. A. Bond; B. Smith; E. Gardiner; K. C. Liew; E. Williams; N. Walsham; M. Putland; D. A. Williamson | Bond, et. al. May 2022 | Utility of SARS-CoV-2 rapid antigen testing for patient triage in the emergency department: A clinical implementation study in Melbourne, Australia | 2022 | PanBio (Abbott) |  |  | At Risk (no symptoms/epidemiology) | 58 | 0 | 102 | 13 | 81.7 | 70.7,89.9 | 100 | 97.1,100 | 100 | 88.7 | 95%CI PPV 95,100; 95%CI NPV 81.5,93.8
Terry-finished | Sherli | 31 | K. A. Bond; B. Smith; E. Gardiner; K. C. Liew; E. Williams; N. Walsham; M. Putland; D. A. Williamson | Bond, et. al. May 2022 | Utility of SARS-CoV-2 rapid antigen testing for patient triage in the emergency department: A clinical implementation study in Melbourne, Australia | 2022 | PanBio (Abbott) |  |  | No Risk (no symptoms/no epidemiology) | 21 | 0 | 420 | 8 | 72.4 | 52.8,87.3 | 100 | 99.3,100 | 100 | 98.1 | 95%CI PPV 86.7,100; 95%CI NPV 96.4,99.2
Ellyn | Sherli | 33 | M. Borro; G. Salerno; A. Montori; A. Petrucca; P. Anibaldi; A. Marcolongo; R. Bonfini; M. Simmaco; I. Santino | Borro, et. al. Apr 2022 | SARS-CoV-2 Transmission Control Measures in the Emergency Department: The Role of Rapid Antigenic Testing in Asymptomatic Subjects | 2022-04-01 | Green Spring "SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold)" |  | Standard protocol = manufacturers instructions UTM-modified protocol for rRT-PCR = 3 drops of NP material collected in UTM | Standard protocol - overall | 79 | 0 | 201 | 20 | 79.8 |  | 100 |  | 100 | 90.9
Ellyn | Sherli | 33 | M. Borro; G. Salerno; A. Montori; A. Petrucca; P. Anibaldi; A. Marcolongo; R. Bonfini; M. Simmaco; I. Santino | Borro, et. al. Apr 2022 | SARS-CoV-2 Transmission Control Measures in the Emergency Department: The Role of Rapid Antigenic Testing in Asymptomatic Subjects | 2022-04-02 | Green Spring "SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold)" |  | Standard protocol = manufacturers instructions UTM-modified protocol for rRT-PCR = 3 drops of NP material collected in UTM | UTM-modified protocol - Overall | 70 | 0 | 201 | 29 | 70.7 |  | 100 |  | 100 | 87.4
Ellyn | Sherli | 33 | M. Borro; G. Salerno; A. Montori; A. Petrucca; P. Anibaldi; A. Marcolongo; R. Bonfini; M. Simmaco; I. Santino | Borro, et. al. Apr 2022 | SARS-CoV-2 Transmission Control Measures in the Emergency Department: The Role of Rapid Antigenic Testing in Asymptomatic Subjects | 2022-04-03 | Green Spring "SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold)" |  | Standard protocol = manufacturers instructions UTM-modified protocol for rRT-PCR = 3 drops of NP material collected in UTM | Standard protocol Ct≤30 (N = 278) | 74 | 0 | 201 | 3 | 96.1 |  | 100 |  | 100 | 98.5
Ellyn | Sherli | 33 | M. Borro; G. Salerno; A. Montori; A. Petrucca; P. Anibaldi; A. Marcolongo; R. Bonfini; M. Simmaco; I. Santino | Borro, et. al. Apr 2022 | SARS-CoV-2 Transmission Control Measures in the Emergency Department: The Role of Rapid Antigenic Testing in Asymptomatic Subjects | 2022-04-04 | Green Spring "SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold)" |  | Standard protocol = manufacturers instructions UTM-modified protocol for rRT-PCR = 3 drops of NP material collected in UTM | UTM-modified protocol Ct≤30 (N = 278) | 66 | 0 | 21 | 11 | 85.7 |  | 100 |  | 100 | 94.8
Ellyn | Sherli | 33 | M. Borro; G. Salerno; A. Montori; A. Petrucca; P. Anibaldi; A. Marcolongo; R. Bonfini; M. Simmaco; I. Santino | Borro, et. al. Apr 2022 | SARS-CoV-2 Transmission Control Measures in the Emergency Department: The Role of Rapid Antigenic Testing in Asymptomatic Subjects | 2022-04-05 | Green Spring "SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold)" |  | Doubled Testing Strategy | UTM-modified protocol overall | 43 | 0 | 801 | 55 | 43.9 |  | 100 |  | 100 | 93.6
Ellyn | Sherli | 33 | M. Borro; G. Salerno; A. Montori; A. Petrucca; P. Anibaldi; A. Marcolongo; R. Bonfini; M. Simmaco; I. Santino | Borro, et. al. Apr 2022 | SARS-CoV-2 Transmission Control Measures in the Emergency Department: The Role of Rapid Antigenic Testing in Asymptomatic Subjects | 2022-04-06 | Green Spring "SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold)" |  | Doubled Testing Strategy | UTM-modified protocol Ct≤30 (N = 852) | 41 | 0 | 801 | 10 | 80.4 |  | 100 |  | 100 | 98.8
Ellyn | Sherli | 33 | M. Borro; G. Salerno; A. Montori; A. Petrucca; P. Anibaldi; A. Marcolongo; R. Bonfini; M. Simmaco; I. Santino | Borro, et. al. Apr 2022 | SARS-CoV-2 Transmission Control Measures in the Emergency Department: The Role of Rapid Antigenic Testing in Asymptomatic Subjects | 2022-04-07 | Green Spring "SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold)" |  | Doubled Testing Strategy | UTM-modified protocol Ct≤25 (N = 839) | 36 | 0 | 801 | 2 | 94.7 |  | 100 |  | 100 | 99.7
Amy | Sherli | 40 | Boum Y. | Boum, et. al. May 2021 | Performance and operational feasibility of antigen and antibody rapid diagnostic tests for COVID-19 in symptomatic and asymptomatic patients in Cameroon: a clinical, prospective, diagnostic accuracy study | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 170 | 54 | 745 | 121 | 58.4 | 53.0,64.8 | 93.2 | 88.0,97.0 | 84 | 76 |  | 58.4192 | 93.2416 | 75.8929 | 86.0277
Amy | Sherli | 40 | Boum Y. | Boum, et. al. May 2021 | Performance and operational feasibility of antigen and antibody rapid diagnostic tests for COVID-19 in symptomatic and asymptomatic patients in Cameroon: a clinical, prospective, diagnostic accuracy study | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | >14 days |  |  |  |  | 19 | 6.0,68.0 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 40 | Boum Y. | Boum, et. al. May 2021 | Performance and operational feasibility of antigen and antibody rapid diagnostic tests for COVID-19 in symptomatic and asymptomatic patients in Cameroon: a clinical, prospective, diagnostic accuracy study | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | 0-7 days |  |  |  |  | 80 | 71.0,88.0 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 40 | Boum Y. | Boum, et. al. May 2021 | Performance and operational feasibility of antigen and antibody rapid diagnostic tests for COVID-19 in symptomatic and asymptomatic patients in Cameroon: a clinical, prospective, diagnostic accuracy study | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | 8-14 days |  |  |  |  | 76 | 59.0,88.0 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 40 | Boum Y. | Boum, et. al. May 2021 | Performance and operational feasibility of antigen and antibody rapid diagnostic tests for COVID-19 in symptomatic and asymptomatic patients in Cameroon: a clinical, prospective, diagnostic accuracy study | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic |  |  |  |  | 37 | 27.0,48.0 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 40 | Boum Y. | Boum, et. al. May 2021 | Performance and operational feasibility of antigen and antibody rapid diagnostic tests for COVID-19 in symptomatic and asymptomatic patients in Cameroon: a clinical, prospective, diagnostic accuracy study | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptomatic |  |  |  |  | 69 | 62.0,75.0 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Sabrina | Sherli | 45 | Bulilete O. | Bulilete, et. al. Feb 2021 | Panbio™ rapid antigen test for SARS-CoV-2 has acceptable accuracy in symptomatic patients in primary health care | 2021-02-01 | PanBio (Abbott) |  |  | Overall | 98 | 2 | 1222 | 40 | 71.4 | 63.1,78.7 | 99.8 | 99.4,99.9 | 98 | 96.8 |  | 71.0145 | 99.8366 | 98 | 96.8304
Sabrina | Sherli | 45 | Bulilete O. | Bulilete, et. al. Feb 2021 | Panbio™ rapid antigen test for SARS-CoV-2 has acceptable accuracy in symptomatic patients in primary health care | 2021-02-01 | PanBio (Abbott) |  |  | Asymptomatic |  |  |  |  | 56.6 | 42.3,70.1 | 100 | 99.4,100.0 | 100 | 96.4 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Sabrina | Sherli | 45 | Bulilete O. | Bulilete, et. al. Feb 2021 | Panbio™ rapid antigen test for SARS-CoV-2 has acceptable accuracy in symptomatic patients in primary health care | 2021-02-01 | PanBio (Abbott) |  |  | Symptomatic |  |  |  |  | 80.4 | 70.5,88.1 | 99.6 | 98.7,99.9 | 97.2 | 97.1 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 34 | E. Burdino; F. Cerutti; F. Panero; T. Allice; G. Gregori; M. G. Milia; G. Cavalot; A. Altavilla; F. Aprà; V. Ghisetti | Burdino, et. al. Oct 2021 | SARS-CoV-2 microfluidic antigen point-of-care testing in Emergency Room patients during COVID-19 pandemic | 2021-10-01 | SARS-CoV-2 Ag (LumiraDx) |  |  | Overall |  |  |  |  | 90.1 | 86.2,93.1 | 99.4 | 98.6,99.8 | 98.2 | 96.5 | PPV 95CI: 95.7-99.3 NPV 95CI: 95.0-97.5
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-01 | Immupass VivaDiag (VivaChek Biotech) |  | 263 participants | overall | 13 | 13 | 207 | 30 | 30.2 | 18.0,46.1 | 94.1 | 90.1,96.6 |  |  | False positive rate: 50.0% (30.5-69.5) False negative rate: 12.7% (9.0-17.6)
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-02 | Immupass VivaDiag (VivaChek Biotech) |  |  | CT=<33 |  |  |  |  | 31.6 | 18.4,48.6
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-03 | Immupass VivaDiag (VivaChek Biotech) |  |  | CT=<29 |  |  |  |  | 80 | 37.8,96.3
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-04 | MEDsan SARS-CoV-2 Antigne Rapid Test |  | 243 participants | overall | 16 | 0 | 120 | 107 | 13 | 8.1,20.3 | 100 | 96.9,100 |  |  | False positive rate: 0.0% (0.0-19.4) False negative rate: 87.0% (79.7-91.9)
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-05 | MEDsan SARS-CoV-2 Antigne Rapid Test |  |  | CT=<33 |  |  |  |  | 23.7 | 44.7,74.3
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-06 | MEDsan SARS-CoV-2 Antigne Rapid Test |  |  | CT=<29 |  |  |  |  | 43.5 | 24.0,65.3
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-07 | Novegent COVID-19 Antigen Rapid Test Kit (Colloidal gold) |  | 229 participants | overall | 46 | 13 | 116 | 54 | 46 | 36.3,56.0 | 89.9 | 83.3,94.1 |  |  | False positive rate: 22.0% (13.0-34.8) False Negative rate: 31.8% (25.1-39.2)
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-08 | Novegent COVID-19 Antigen Rapid Test Kit (Colloidal gold) |  |  | CT=<33 |  |  |  |  | 58.2 | 44.5,70.8
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-09 | Novegent COVID-19 Antigen Rapid Test Kit (Colloidal gold) |  |  | CT=<29 |  |  |  |  | 66.7 | 46.0,82.5
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-10 | PanBio (Abbott) |  | 185 participants | overall | 42 | 0 | 100 | 43 | 49.4 | 38.7,60.1 | 100 | 96.4,100 |  |  | False positive rate: 0% (0-8.4) False Negative rate: 50.6% (39.9-61.2)
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-11 | PanBio (Abbott) |  |  | CT=<33 |  |  |  |  | 72 | 57.6,83.0
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-12 | PanBio (Abbott) |  |  | CT=<29 |  |  |  |  | 85.7 | 66.0,94.9
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-13 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  | 172 participants | overall | 35 | 8 | 71 | 58 | 37.6 | 28.2,48.1 | 89.9 | 80.8,94.9 |  |  | False positive rate: 10.1% (5.1-19.2) False negative rate: 62.4% (51.9-71.8)
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-14 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT=<33 |  |  |  |  | 71 | 51.8,84.8
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-15 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT=<29 |  |  |  |  | 77.8 | 57.0,90.2
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-16 | RapiGen (BioCredit) |  | 247 participants | overall | 37 | 2 | 98 | 110 | 27.4 | 20.5,35.6 | 98.2 | 93.1,99.6 |  |  | False Positive rate: 5.1% (1.2-19.3) False Negative Rate: 72.6% (64.4-79.5)
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-17 | RapiGen (BioCredit) |  |  | CT=<33 |  |  |  |  | 44.7 | 33.8,56.3
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-18 | RapiGen (BioCredit) |  |  | CT=<29 |  |  |  |  | 60 | 45.5,72.9
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-19 | RapiGen (BioCredit) |  |  | Symptom onset 0-7 days |  |  |  |  | 39.3 | 27.7,52.4
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-20 | RapiGen (BioCredit) |  |  | Symptom onset >7 days |  |  |  |  | 11.1 | 3.3,31.1
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-21 | RapiGen (BioCredit) |  |  | Asymptomatic |  |  |  |  | 21.3 | 11.6-35.8
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-22 | Respi-Strip (Coris BioConcept) |  | 194 participants | overall | 13 | 1 | 140 | 53 | 19.4 | 11.5,30.9 | 99.2 | 94.5,99.9 |  |  | False positive rate: 0.8% (0.1-5.5) False negative rate: 80.6% (69.1-88.5)
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-23 | Respi-Strip (Coris BioConcept) |  |  | CT=<33 |  |  |  |  | 35.1 | 21.1,52.4
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-24 | Respi-Strip (Coris BioConcept) |  |  | CT=<29 |  |  |  |  | 46.4 | 28.2,65.7
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-25 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | Symptom onset 0-7 days |  |  |  |  | 26.3 | 10.4,52.4
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-26 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | Symptom onset >7 days |  |  |  |  | 11.1 | 0.9,62.6
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-27 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | Asymptomatic |  |  |  |  | 46 | 33.9,58.7
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-28 | MEDsan SARS-CoV-2 Antigne Rapid Test |  |  | Symptom onset 0-7 days |  |  |  |  | 12 | 3.96,33.3
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-29 | MEDsan SARS-CoV-2 Antigne Rapid Test |  |  | Symptom onset >7 days |  |  |  |  | 21.7 | 6.7,44.8
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-30 | PanBio (Abbott) |  |  | Symptom onset 0-7 days |  |  |  |  | 51.1 | 36.3,65.8
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-31 | PanBio (Abbott) |  |  | Symptom onset >7 days |  |  |  |  | 44.4 | 13.4,80.5
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-01 | PanBio (Abbott) |  |  | Asymptomatic |  |  |  |  | 48.4 | 30.8,66.4
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-01 | Novegent COVID-19 Antigen Rapid Test Kit (Colloidal gold) |  |  | Symptom onset 0-7 days |  |  |  |  | 65 | 38.6,71.4
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-01 | Novegent COVID-19 Antigen Rapid Test Kit (Colloidal gold) |  |  | Symptom onset >7 days |  |  |  |  | 42.9 | 28.4,58.7
Amy | Sherli | 35 | J. Bwogi; T. Lutalo; P. Tushabe; H. Bukenya; J. P. Eliku; I. Ssewanyana; S. Nabadda; C. Nsereko; M. Cotten; R. Downing; J. Lutwama; P. Kaleebu | Bwogi, et. al. May 2022 | Field evaluation of the performance of seven Antigen Rapid diagnostic tests for the diagnosis of SARs-CoV-2 virus infection in Uganda | 2022-05-01 | Immupass VivaDiag (VivaChek Biotech) |  |  | Symptom onset 0-7 days |  |  |  |  | 26.3 | 14.4,43.2
Sabrina | Sherli | 49 | Caruana Giorgia | Caruana, et. al. Apr 2021 | Implementing SARS-CoV-2 Rapid Antigen Testing in the Emergency Ward of a Swiss University Hospital: The INCREASE Study | 2021-04-01 | PanBio (Abbott) |  |  | Overall |  |  |  |  | 41.2 |  | 99.5 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Sabrina | Sherli | 49 | Caruana Giorgia | Caruana, et. al. Apr 2021 | Implementing SARS-CoV-2 Rapid Antigen Testing in the Emergency Ward of a Swiss University Hospital: The INCREASE Study | 2021-04-01 | BD Veritor COVID-19 Rapid Antigen Test (Becton-Dickinson) |  |  | Overall |  |  |  |  | 41.2 |  | 99.7 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Sabrina | Sherli | 49 | Caruana Giorgia | Caruana, et. al. Apr 2021 | Implementing SARS-CoV-2 Rapid Antigen Testing in the Emergency Ward of a Swiss University Hospital: The INCREASE Study | 2021-04-01 | Exdia (Precision Biosensor) |  |  | Overall |  |  |  |  | 48.3 |  | 99.5 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Sabrina | Sherli | 49 | Caruana Giorgia | Caruana, et. al. Apr 2021 | Implementing SARS-CoV-2 Rapid Antigen Testing in the Emergency Ward of a Swiss University Hospital: The INCREASE Study | 2021-04-01 | Other | Mixed (Standard Q, Panbio, Exdia COVID 19 Ag, BD Veritor) |  | Overall | 47 | 1 | 417 | 67 | 41.2 |  | 99.8 |  | 97.9 | 86.2 |  | 41.2281 | 99.7608 | 97.9167 | 86.157
Sabrina | Sherli | 49 | Caruana Giorgia | Caruana, et. al. Apr 2021 | Implementing SARS-CoV-2 Rapid Antigen Testing in the Emergency Ward of a Swiss University Hospital: The INCREASE Study | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall |  |  |  |  | 41.2 |  | 99.7 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 50 | Caruana G. | Caruana, et. al. May 2021 | The dark side of SARS-CoV-2 rapid antigen testing: screening asymptomatic patients | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall |  |  |  |  | 28.6 |  | 98.2 |  | 50 |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Chathu | Sherli | 37 | N. G. Cassuto; A. Gravier; M. Colin; A. Theillay; D. Pires-Roteira; S. Pallay; R. Serreau; L. Hocqueloux; T. Prazuck | Cassuto, et. al. Jul 2021 | Evaluation of a SARS-CoV-2 antigen-detecting rapid diagnostic test as a self-test: Diagnostic performance and usability | 2021-07-01 | COVID‐VIRO nasal swab test |  |  | Overall | 31 | 0 | 202 | 1 | 96.88 | 83.78,99.92 | 100 | 98.19,100.00 | 100 | 99.5 | NPV 95% CI 96.7,99.93
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-01 | SARS-CoV-2 Ag (LumiraDx) |  |  | Overall | 174 | 3 | 51 | 54 | 76.3 | 70.8,81.8 | 94.4 | 88.3,100 | 98.3 | 48.6 | PPV 95% CI (96.4,100) NPV 95%CI (39,58.2)
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-02 | SARS-CoV-2 Ag (LumiraDx) |  |  | Asymptomatic | 3 | 1 | 25 | 3 | 50 | 10,90 | 96.2 | 88.0,100 | 75 | 89.3 | PPV 95% CI (32.6,100) NPV 95%CI (77.8,100)
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-03 | SARS-CoV-2 Ag (LumiraDx) |  |  | Pauci-symptomatic | 115 | 2 | 6 | 23 | 83.3 | 77.1,89.5 | 75 | 45.0,100 | 98.3 | 20.7 | PPV 95% CI (96,100) NPV 95%CI (6,35.4)
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-04 | SARS-CoV-2 Ag (LumiraDx) |  |  | Neurological symptoms | 1 | 0 | 1 | 0 | 100 | 100100 | 100 | 100100 | 100 | 100 | PPV 95% CI (100,100) NPV 95%CI (100,100)
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-05 | SARS-CoV-2 Ag (LumiraDx) |  |  | Female | 63 | 1 | 23 | 18 | 77.8 | 68.7,86.9 | 95.8 | 87.8,100 | 98.4 | 56.1 | PPV 95% CI (95.4,100) NPV 95%CI (40.9,71.3)
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-06 | SARS-CoV-2 Ag (LumiraDx) |  |  | Male | 111 | 2 | 28 | 36 | 75.5 | 68.5,82.5 | 93.3 | 84.4,100 | 98.2 | 43.8 | PPV 95% CI (95.8,100) NPV 95%CI (31.6,56)
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-07 | SARS-CoV-2 Ag (LumiraDx) |  |  | Age 0-39 | 12 | 1 | 18 | 6 | 66.7 | 44.9,88.5 | 94.7 | 84.7,100 | 92.3 | 75 | PPV 95% CI (77.8,100) NPV 95%CI (57.7,92.3)
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-08 | SARS-CoV-2 Ag (LumiraDx) |  |  | Age 40-59 | 54 | 1 | 18 | 15 | 78.3 | 68.6,88 | 94.7 | 84.7,100 | 98.2 | 54.4 | PPV 95% CI (94.7,100) NPV 95%CI (37.4,71.4)
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-09 | SARS-CoV-2 Ag (LumiraDx) |  |  | Age 60-79 | 79 | 1 | 8 | 27 | 74.5 | 66.2,82.8 | 88.9 | 68.4,100 | 98.8 | 22.9 | PPV 95% CI (96.4,100) NPV 95%CI (9,36.8)
Ellyn | Sherli | 38 | A. M. Cattelan; L. Sasset; F. Zabeo; A. Ferrari; L. Rossi; M. Mazzitelli; S. Cocchio; V. Baldo | Cattelan, et. al. Mar 2022 | Rapid Antigen Test LumiraDx(TM) vs. Real Time Polymerase Chain Reaction for the Diagnosis of SARS-CoV-2 Infection: A Retrospective Cohort Study | 2022-03-10 | SARS-CoV-2 Ag (LumiraDx) |  |  | Age 80+ | 29 | 0 | 7 | 6 | 82.9 | 70.4,95.4 | 100 | 100100 | 100 | 53.8 | PPV 95% CI (100,100) NPV 95%CI (26.7,80.9)
Yati | Sherli | 53 | Cento V. | Cento, et. al. May 2021 | Frontline Screening for SARS-CoV-2 Infection at Emergency Department Admission by Third Generation Rapid Antigen Test: Can We Spare RT-qPCR? | 2021-05-01 | SARS-CoV-2 Ag (LumiraDx) |  |  | Overall | 297 | 17 | 560 | 50 | 85.6 | 82,89 | 97.05 | 96,98 | 94.6 | 91.8 |  | 85.5908 | 97.0537 | 94.586 | 91.8033
Terry | Sherli | 54 | Cerutti F. | Cerutti, et. al. Aug 2021 | Urgent need of rapid tests for SARS CoV-2 antigen detection: Evaluation of the SD-Biosensor antigen test for SARS-CoV-2 | 2020-09-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 77 | 0 | 221 | 32 | 70.6 |  | 100 |  | 100 | 87.4 |  | 70.6422 | 100 | 100 | 87.3518
Yati | Sherli | 55 | Chaimayo C. | Chaimayo, et. al. Nov 2020 | Rapid SARS-CoV-2 antigen detection assay in comparison with real-time RT-PCR assay for laboratory diagnosis of COVID-19 in Thailand | 2020-11-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall |  |  |  |  | 98.33 | 91.06,99.6 | 98.73 | 97.06,99.59 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 43 | C. C. Cheng; C. C. Liu; T. F. Chiu; S. Shiou-Sheng Chen | Cheng, et. al. May 2022 | Evaluation of a Rapid Antigen Test for the Diagnosis of SARS-CoV-2 during the COVID-19 Pandemic | 2022-05-01 | Other |  | Total - All 3 tests |  | 289 | 60 | 5707 | 149 | 66 | 65.8,66.1 | 99 | 99.0,99.0 | 82.8 | 97.5 | PPV 95%CI 82.7,82.9 NPV 95% CI 97.4,97.5
Ellyn | Sherli | 43 | C. C. Cheng; C. C. Liu; T. F. Chiu; S. Shiou-Sheng Chen | Cheng, et. al. May 2022 | Evaluation of a Rapid Antigen Test for the Diagnosis of SARS-CoV-2 during the COVID-19 Pandemic | 2022-05-02 | VTRUST COVID-19 Antigen Rapid Test (Taidoc Technology Corporation) |  |  | YangMing | 44 | 13 | 718 | 12 | 78.6 | 78.2,78.9 | 98.2 | 98.2,98.3 | 77.2 | 98.4 | PPV 95%CI 76.8,77.5 NPV 95% CI 98.3,98.4
Ellyn | Sherli | 43 | C. C. Cheng; C. C. Liu; T. F. Chiu; S. Shiou-Sheng Chen | Cheng, et. al. May 2022 | Evaluation of a Rapid Antigen Test for the Diagnosis of SARS-CoV-2 during the COVID-19 Pandemic | 2022-05-03 | VTRUST COVID-19 Antigen Rapid Test (Taidoc Technology Corporation) |  |  | Zhongxiao | 52 | 5 | 533 | 34 | 60.5 | 60.1,60.8 | 99.1 | 99.0,99.1 | 91.2 | 94 | PPV 95%CI 91.0,91.5 NPV 95% CI 93.9,94.1
Ellyn | Sherli | 43 | C. C. Cheng; C. C. Liu; T. F. Chiu; S. Shiou-Sheng Chen | Cheng, et. al. May 2022 | Evaluation of a Rapid Antigen Test for the Diagnosis of SARS-CoV-2 during the COVID-19 Pandemic | 2022-05-04 | Enimmune Speedy COVID-19 AG Rapid Test |  |  | Zhongxing | 93 | 32 | 1885 | 51 | 64.6 | 64.3,64.8 | 98.3 | 98.3,98.3 | 74.4 | 97.4 | PPV 95%CI 74.2,74.6 NPV 95% CI 97.3,97.4
Ellyn | Sherli | 43 | C. C. Cheng; C. C. Liu; T. F. Chiu; S. Shiou-Sheng Chen | Cheng, et. al. May 2022 | Evaluation of a Rapid Antigen Test for the Diagnosis of SARS-CoV-2 during the COVID-19 Pandemic | 2022-05-05 | Enimmune Speedy COVID-19 AG Rapid Test |  |  | Heping | 56 | 9 | 1025 | 25 | 69.1 | 68.8,69.5 | 99.1 | 99.1,99.1 | 86.2 | 97.6 | PPV 95%CI 85.9,86.4 NPV 95% CI 97.6,97.6
Ellyn | Sherli | 43 | C. C. Cheng; C. C. Liu; T. F. Chiu; S. Shiou-Sheng Chen | Cheng, et. al. May 2022 | Evaluation of a Rapid Antigen Test for the Diagnosis of SARS-CoV-2 during the COVID-19 Pandemic | 2022-05-06 | PanBio (Abbott) |  |  | RenAi | 44 | 1 | 1546 | 27 | 62 | 61.6,62.3 | 99.9 | 99.9,99.9 | 97.8 | 98.3 | PPV 95%CI 97.6,97.9 NPV 95% CI 98.3,98.3
Chathu | Sherli | 46 | S. I. A. Choudhary | Choudhary, et. al. Apr 2022 | Validation of rapid SARS-COV-2 antigen detection test as a screening tool for detection of Covid-19 infection at district hospital in northern India | 2022-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 71 | 5 | 622 | 58 | 55.04 | 46.43,63.35 | 99.2 | 98.15,99.66 | 93.42 | 91.47 | PPV 95% CI (85.51,97.16) NPV 95%CI (89.13,93.34)
Brandi | Sherli | 51 | E. Cottone; F. Van Hoecke; G. A. Martens; E. De Laere; R. De Smedt; S. Vervaeke; M. Vanhee; D. De Smet | Cottone, et. al. May 2022 | Pitfalls of SARS-CoV-2 antigen testing at emergency department | 2022-05-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall | 46 | 7 | 356 | 55 | 45.5 | 35.6,55.8 | 98.1 | 96.1,99.2 | 86.8 | 86.6 | PPV and NPV calculated with prevalence of 21.8%. PPV 95%CI 74.7,94.5 NPV 95% CI 82.9,89.8
Brandi | Sherli | 51 | E. Cottone; F. Van Hoecke; G. A. Martens; E. De Laere; R. De Smedt; S. Vervaeke; M. Vanhee; D. De Smet | Cottone, et. al. May 2022 | Pitfalls of SARS-CoV-2 antigen testing at emergency department | 2022-05-02 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptomatic <=5 |  |  |  |  | 68.2 | 50.0,86.4
Brandi | Sherli | 51 | E. Cottone; F. Van Hoecke; G. A. Martens; E. De Laere; R. De Smedt; S. Vervaeke; M. Vanhee; D. De Smet | Cottone, et. al. May 2022 | Pitfalls of SARS-CoV-2 antigen testing at emergency department | 2022-05-03 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptomatic >5 |  |  |  |  | 40.9 | 27.3,54.5
Brandi | Sherli | 51 | E. Cottone; F. Van Hoecke; G. A. Martens; E. De Laere; R. De Smedt; S. Vervaeke; M. Vanhee; D. De Smet | Cottone, et. al. May 2022 | Pitfalls of SARS-CoV-2 antigen testing at emergency department | 2022-05-04 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptomatic <=7 |  |  |  |  | 67.7 | 51.6,83.9
Brandi | Sherli | 51 | E. Cottone; F. Van Hoecke; G. A. Martens; E. De Laere; R. De Smedt; S. Vervaeke; M. Vanhee; D. De Smet | Cottone, et. al. May 2022 | Pitfalls of SARS-CoV-2 antigen testing at emergency department | 2022-05-05 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptomatic >7 |  |  |  |  | 34.3 | 20.0,51.4
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-01 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Northumberland Overall | 74 | 4 | 157 | 12 | 86 | 76.9,92.6 | 97.5 | 93.8,99.3 | 94.9 | 92.9 | PPV CI95 [87.5-98.0] NPV CI95 [88.6-95.7]
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-02 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Overall Yorkshire | 11 | 0 | 40 | 2 | 84.6 | 54.6,98.1 | 100 | 91.2,100 | 100 | 95.2 | PPV CI95 [71.5-100] NPV CI95 [84.8-98.6]
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-03 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Northumberland CT<20 | 51 |  |  | 1 | 98.1 | 89.7,100.0
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-04 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Yorkshire CT<20 | 1 |  |  | 0 | 100 | 2.5,100
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-05 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Northumberland CT<25 | 70 |  |  | 5 | 93.3 | 85.1,97.8
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-06 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Yorkshire CT<25 | 5 |  |  | 0 | 100 | 47.8,100
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-07 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Northumberland CT<33 | 73 |  |  | 9 | 89 | 80.2,94.9
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-08 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Yorkshire CT<33 | 5 |  |  | 1 | 83.3 | 35.9,99.6
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-09 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Northumberland CT<45 | 74 |  |  | 12 | 86 | 76.9,92.6
Amy | Sherli | 277 | A. B. F. B. R. B. K. C. D. C. M. C. A. C. L. D. A. E. N. E. T. F. D. F. T. G. E. H.-W. Cubas-Atienzar | Cubas-Atienzar, et. al. May 2021 | Accuracy of the Mologic COVID-19 rapid antigen test: a prospective multi-centre analytical and clinical evaluation | 2021-05-10 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Yorkshire CT<45 | 6 |  |  | 1 | 85.7 | 42.1,99.6
Ellyn | Sherli | 75 | Dierks S. | Dierks, et. al. May 2021 | Diagnosing SARS-CoV-2 with Antigen Testing, Transcription-Mediated Amplification and Real-Time PCR | 2021-05-01 | SARS-CoV-2 Ag (LumiraDx) |  |  | Overall |  |  |  |  | 45.45 | 20.22,73.26 | 99.54 | 98.17,99.88 | 71.02 | 98.63 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 75 | Dierks S. | Dierks, et. al. May 2021 | Diagnosing SARS-CoV-2 with Antigen Testing, Transcription-Mediated Amplification and Real-Time PCR | 2021-05-01 | NADAL COVID-19 Antigen Rapid Test (New Art Laboratories/nal von minden) |  |  | Overall |  |  |  |  | 14.29 | 1.94,58.35 | 76.44 | 70.16,81.74 | 9.66 | 97.68 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-01 | PanBio (Abbott) |  |  | COVID-19 Suspected Cases; Overall |  |  |  |  | 71.1 | 55.6,83.6 |  |  |  | 92.2 | NPV 95CI [87.1 - 95.8]
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-02 | SGTI-Flex |  |  | COVID-19 Suspected Cases; Overall |  |  |  |  | 68.8 | 55.7,80 |  |  |  | 88 | NPV 95CI [81.8-92.6]
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-03 | NovaGen |  |  | COVID-19 Suspected Cases; Overall |  |  |  |  | 84.6 | 72.0,93.1 |  |  |  | 95 | NPV 95CI [90.1-97.8]
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-04 | PanBio (Abbott) |  |  | Close Asymptomatic Contacts; Overall |  |  |  |  | 33.3 | 11.8,61.6 |  |  |  | 89.5 | NPV 95CI [81.4-94.8]
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-05 | SGTI-Flex |  |  | Close Asymptomatic Contacts; Overall |  |  |  |  | 84.6 | 54.5,98.1 |  |  |  | 97.7 | NPV 95CI [92.1-99.7]
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-06 | NovaGen |  |  | Close Asymptomatic Contacts; Overall |  |  |  |  | 55.5 | 30.7,78.4 |  |  |  | 91.1 | NPV 95CI [83.2-96]
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-07 | PanBio (Abbott) |  |  | COVID-19 Suspected Cases; CT=<20 |  |  |  |  | 100 | 100.0,100.0
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-08 | SGTI-Flex |  |  | COVID-19 Suspected Cases; CT=<20 |  |  |  |  | 87.8 | 73.8,96
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-09 | NovaGen |  |  | COVID-19 Suspected Cases; CT=<20 |  |  |  |  | 93.7 | 79.2,99.2
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-10 | PanBio (Abbott) |  |  | Close Asymptomatic Contacts; CT=<20 |  |  |  |  | 100 | 100.0,100.0
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-11 | SGTI-Flex |  |  | Close Asymptomatic Contacts; CT=<20 |  |  |  |  | 100 | 100.0,100.0
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-12 | NovaGen |  |  | Close Asymptomatic Contacts; CT=<20 |  |  |  |  | 83.3 | 35.9,99.6
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-13 | PanBio (Abbott) |  |  | COVID-19 Suspected Cases; CT>20 |  |  |  |  | 45.8 | 25.5,67.2
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-14 | SGTI-Flex |  |  | COVID-19 Suspected Cases; CT>20 |  |  |  |  | 30 | 11.9,54.3
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-15 | NovaGen |  |  | COVID-19 Suspected Cases; CT>20 |  |  |  |  | 70 | 45.7,88.1
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-16 | PanBio (Abbott) |  |  | Close Asymptomatic Contacts; CT>20 |  |  |  |  | 16.7 | 2.1,48.4
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-17 | SGTI-Flex |  |  | Close Asymptomatic Contacts; CT>20 |  |  |  |  | 60 | 14.6,94.7
Sabrina/Amy | Sherli | 53 | P. Escribano; A. E. Sánchez-Pulido; J. González-Leiva; I. Valero-López; P. Catalán; P. Muñoz; J. Guinea | Escribano, et. al. Feb 2022 | Different performance of three point-of-care SARS-CoV-2 antigen detection devices in symptomatic patients and close asymptomatic contacts: a real-life study | 2022-02-18 | NovaGen |  |  | Close Asymptomatic Contacts; CT>20 |  |  |  |  | 41.7 | 15.2,72.2
Chathu | Sherli | 54 | B. F. Escrivá; M. D. O. Mochón; R. M. González; C. S. García; A. T. Pla; A. S. Ricart; M. M. García; I. T. Aranda; F. G. García; C. G. Cardona | Escrivá, et. al. Aug 2021 | The effectiveness of rapid antigen test-based for SARS-CoV-2 detection in nursing homes in Valencia, Spain"" | 2021-08-01 | PanBio (Abbott) |  |  | Overall | 99 | 0 | 331 | 18 | 85 | 90-99 | 100 | 100100 | 100 | 95 | PPV 95% CI (100,100) NPV 95%CI (97,99.6)
Terry | Sherli | 55 | K. S. Faíco-Filho; F. E. F. Júnior; L. V. L. Moreira; P. R. G. Lins; A. F. O. Justo; N. Bellei | Faíco-Filho, et. al. Mar 2022 | Evaluation of the Panbio™ COVID-19 ag rapid test at an emergency room in a hospital in São Paulo, Brazil | 2022-03-01 | PanBio (Abbott) |  |  | Ct<=25 |  |  |  |  | 97.3 |  |  |  | 60 | 98.5
Terry | Sherli | 55 | K. S. Faíco-Filho; F. E. F. Júnior; L. V. L. Moreira; P. R. G. Lins; A. F. O. Justo; N. Bellei | Faíco-Filho, et. al. Mar 2022 | Evaluation of the Panbio™ COVID-19 ag rapid test at an emergency room in a hospital in São Paulo, Brazil | 2022-03-02 | PanBio (Abbott) |  |  | Ct<=30 |  |  |  |  | 88.1 |  |  |  | 86.7 | 89.5
Terry | Sherli | 55 | K. S. Faíco-Filho; F. E. F. Júnior; L. V. L. Moreira; P. R. G. Lins; A. F. O. Justo; N. Bellei | Faíco-Filho, et. al. Mar 2022 | Evaluation of the Panbio™ COVID-19 ag rapid test at an emergency room in a hospital in São Paulo, Brazil | 2022-03-03 | PanBio (Abbott) |  |  | Ct<=35 |  |  |  |  | 83.8 |  |  |  | 95 | 83.6
Terry | Sherli | 55 | K. S. Faíco-Filho; F. E. F. Júnior; L. V. L. Moreira; P. R. G. Lins; A. F. O. Justo; N. Bellei | Faíco-Filho, et. al. Mar 2022 | Evaluation of the Panbio™ COVID-19 ag rapid test at an emergency room in a hospital in São Paulo, Brazil | 2022-03-04 | PanBio (Abbott) |  |  | Ct<=40 |  |  |  |  | 84.3 |  |  |  | 98.3 | 83.6
Terry | Sherli | 56 | E. Farfour; C. Amiel; E. Jolly; S. Zia-Chahabi; L. Mazaux; P. Simonin; I. Berkane; M. Cornily; T. Pascreau; V. M. Sars-Co | Farfour, et. al. Mar 2021 | The Panbio COVID-19 Ag Rapid Test: Which Performances are for COVID-19 Diagnosis? | 2021-03-01 | PanBio (Abbott) |  |  | Prospective study Symptomatic |  |  |  |  | 47.2% |  | 100% |  | 100% | 84.3%
Terry | Sherli | 56 | E. Farfour; C. Amiel; E. Jolly; S. Zia-Chahabi; L. Mazaux; P. Simonin; I. Berkane; M. Cornily; T. Pascreau; V. M. Sars-Co | Farfour, et. al. Mar 2021 | The Panbio COVID-19 Ag Rapid Test: Which Performances are for COVID-19 Diagnosis? | 2021-03-02 | PanBio (Abbott) |  |  | Prospective study </= 4 days symptoms appear |  |  |  |  | 55.2% |  | 100% |  | 100% | 85.6%
Terry | Sherli | 56 | E. Farfour; C. Amiel; E. Jolly; S. Zia-Chahabi; L. Mazaux; P. Simonin; I. Berkane; M. Cornily; T. Pascreau; V. M. Sars-Co | Farfour, et. al. Mar 2021 | The Panbio COVID-19 Ag Rapid Test: Which Performances are for COVID-19 Diagnosis? | 2021-03-03 | PanBio (Abbott) |  |  | Prospective study > 4 days symptoms appear |  |  |  |  | 14.3% |  | 100% |  | 100% | 80%
Terry | Sherli | 56 | E. Farfour; C. Amiel; E. Jolly; S. Zia-Chahabi; L. Mazaux; P. Simonin; I. Berkane; M. Cornily; T. Pascreau; V. M. Sars-Co | Farfour, et. al. Mar 2021 | The Panbio COVID-19 Ag Rapid Test: Which Performances are for COVID-19 Diagnosis? | 2021-03-04 | PanBio (Abbott) |  |  | Prospective study </= 35 years old |  |  |  |  | 61.1% |  | 100% |  | 100% | 88.5%
Terry | Sherli | 56 | E. Farfour; C. Amiel; E. Jolly; S. Zia-Chahabi; L. Mazaux; P. Simonin; I. Berkane; M. Cornily; T. Pascreau; V. M. Sars-Co | Farfour, et. al. Mar 2021 | The Panbio COVID-19 Ag Rapid Test: Which Performances are for COVID-19 Diagnosis? | 2021-03-05 | PanBio (Abbott) |  |  | Prospective study > 35 years old |  |  |  |  | 33.3% |  | 100% |  | 100% | 78.9%
Yati | Sherli | 87 | Fernandez-Montero A. | Fernandez-Montero, et. al. July 2021 | Validation of a rapid antigen test as a screening tool for SARS-CoV-2 infection in asymptomatic populations. Sensitivity, specificity and predictive values | 2021-07-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall | 35 | 8 | 2486 | 14 | 71.43 | 56.74,83.42 | 99.68 | 99.37,99.86 | 81.4 | 99.44 |  | 71.4286 | 99.6792 | 81.3953 | 99.44
Yati | Sherli | 87 | Fernandez-Montero A. | Fernandez-Montero, et. al. July 2021 | Validation of a rapid antigen test as a screening tool for SARS-CoV-2 infection in asymptomatic populations. Sensitivity, specificity and predictive values | 2021-07-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Ct<25 |  |  |  |  | 93.75 | 71.96,98.93 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Yati | Sherli | 87 | Fernandez-Montero A. | Fernandez-Montero, et. al. July 2021 | Validation of a rapid antigen test as a screening tool for SARS-CoV-2 infection in asymptomatic populations. Sensitivity, specificity and predictive values | 2021-07-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Ct>25 |  |  |  |  | 29.41 | 10.31,55.96 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 89 | Ferté T. | Ferté, et. al. Jun 2021 | Accuracy of COVID-19 rapid antigenic tests compared to RT-PCR in a student population: The StudyCov study | 2021-06-01 | PanBio (Abbott) |  |  | Overall |  |  |  |  | 63.5 | 49.0,76.4 | 100 | 99.4,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 89 | Ferté T. | Ferté, et. al. Jun 2021 | Accuracy of COVID-19 rapid antigenic tests compared to RT-PCR in a student population: The StudyCov study | 2021-06-01 | PanBio (Abbott) |  |  | Asymptomatic |  |  |  |  | 35 | 15.4,59.2 | 100 | 99.3,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 58 | F. Fitoussi; S. Tonen-Wolyec; N. Awaida; R. Dupont; L. Bélec | Fitoussi, et. al. Oct 2021 | Analytical performance of the point-of-care BIOSYNEX COVID-19 Ag BSS for the detection of SARS‐CoV‐2 nucleocapsid protein in nasopharyngeal swabs: a prospective field evaluation during the COVID-19 third wave in France | 2021-10-01 | BIOSYNEX Ag-RDT |  |  | Overall | 121 | 3 | 816 | 27 | 81.8% | 79.2,84.1 | 99.6% | 98.9,99.8 | 96.6 | 97.5 | PPV 95% CI 95.3, 97.6 NPV 95% CI 97.2, 99.8: PPV and NPV was calculated using prevalence
Terry | Sherli | 58 | F. Fitoussi; S. Tonen-Wolyec; N. Awaida; R. Dupont; L. Bélec | Fitoussi, et. al. Oct 2021 | Analytical performance of the point-of-care BIOSYNEX COVID-19 Ag BSS for the detection of SARS‐CoV‐2 nucleocapsid protein in nasopharyngeal swabs: a prospective field evaluation during the COVID-19 third wave in France | 2021-10-02 | BIOSYNEX Ag-RDT |  |  | Asymptomatic | 45 | 2 | 652 | 23 | 79.4% | 76.3,82.2 | 99.7% | 98.9,99.9 | 97.3 | 97.2 | PPV 95% CI 95.8, 98.2 NPV 95% CI 95.7, 98.2: PPV and NPV was calculated using prevalence. Sensitivity given in paper is different from the calculations
Terry | Sherli | 58 | F. Fitoussi; S. Tonen-Wolyec; N. Awaida; R. Dupont; L. Bélec | Fitoussi, et. al. Oct 2021 | Analytical performance of the point-of-care BIOSYNEX COVID-19 Ag BSS for the detection of SARS‐CoV‐2 nucleocapsid protein in nasopharyngeal swabs: a prospective field evaluation during the COVID-19 third wave in France | 2021-10-03 | BIOSYNEX Ag-RDT |  |  | Symptomatic | 76 | 1 | 164 | 4 | 95.6% | 92.2,97.5 | 99.3% | 97.2,99.8 | 95.6 | 99.3 | PPV 95% CI 92.2, 97.5 NPV 95% CI 97.2, 99.8: PPV and NPV was calculated using prevalence
Terry | Sherli | 58 | F. Fitoussi; S. Tonen-Wolyec; N. Awaida; R. Dupont; L. Bélec | Fitoussi, et. al. Oct 2021 | Analytical performance of the point-of-care BIOSYNEX COVID-19 Ag BSS for the detection of SARS‐CoV‐2 nucleocapsid protein in nasopharyngeal swabs: a prospective field evaluation during the COVID-19 third wave in France | 2021-10-04 | BIOSYNEX Ag-RDT |  |  | Symptomatic (</= 7 days from onset) | 72 | 1 | 153 | 3 | 96% | 92.6,97.9 | 99.4% | 97.3,99.9 | 95.7 | 99.4 | PPV 95% CI 92.2, 97.7 NPV 95% CI 97.3, 99.9: PPV and NPV was calculated using prevalence
Terry | Sherli | 60 | L. Ford; M. J. Whaley; M. M. Shah; P. P. Salvatore; H. E. Segaloff; A. Delaney; D. W. Currie; L. Boyle-Estheimer; M. O'Hegarty; C. N. Morgan; J. Meece; L. Ivacic; N. J. Thornburg; A. Tamin; J. L. Harcourt; J. M. Folster; M. Medrzycki; S. Jain; P. Wong; K. Goffard; D. Gieryn; J. Kahrs; K. Langolf; T. Zochert; J. E. Tate; C. H. Hsu; H. L. Kirking | Ford, et. al. Sep 2021 | Antigen Test Performance Among Children and Adults at a SARS-CoV-2 Community Testing Site | 2021-09-01 | BinaxNOW (Abbott) |  |  | Overall |  |  |  |  | 80.8% |  | 99.9% |  | 99.2% | 96.5%
Terry | Sherli | 60 | L. Ford; M. J. Whaley; M. M. Shah; P. P. Salvatore; H. E. Segaloff; A. Delaney; D. W. Currie; L. Boyle-Estheimer; M. O'Hegarty; C. N. Morgan; J. Meece; L. Ivacic; N. J. Thornburg; A. Tamin; J. L. Harcourt; J. M. Folster; M. Medrzycki; S. Jain; P. Wong; K. Goffard; D. Gieryn; J. Kahrs; K. Langolf; T. Zochert; J. E. Tate; C. H. Hsu; H. L. Kirking | Ford, et. al. Sep 2021 | Antigen Test Performance Among Children and Adults at a SARS-CoV-2 Community Testing Site | 2021-09-02 | BinaxNOW (Abbott) |  |  | Symptomatic |  |  |  |  | 83.3% |  | 99.9% |  | 99.5% | 95.1%
Terry | Sherli | 60 | L. Ford; M. J. Whaley; M. M. Shah; P. P. Salvatore; H. E. Segaloff; A. Delaney; D. W. Currie; L. Boyle-Estheimer; M. O'Hegarty; C. N. Morgan; J. Meece; L. Ivacic; N. J. Thornburg; A. Tamin; J. L. Harcourt; J. M. Folster; M. Medrzycki; S. Jain; P. Wong; K. Goffard; D. Gieryn; J. Kahrs; K. Langolf; T. Zochert; J. E. Tate; C. H. Hsu; H. L. Kirking | Ford, et. al. Sep 2021 | Antigen Test Performance Among Children and Adults at a SARS-CoV-2 Community Testing Site | 2021-09-03 | BinaxNOW (Abbott) |  |  | Asymptomatic |  |  |  |  | 65.9% |  | 99.9% |  | 96.4% | 98.1%
Terry | Sherli | 60 | L. Ford; M. J. Whaley; M. M. Shah; P. P. Salvatore; H. E. Segaloff; A. Delaney; D. W. Currie; L. Boyle-Estheimer; M. O'Hegarty; C. N. Morgan; J. Meece; L. Ivacic; N. J. Thornburg; A. Tamin; J. L. Harcourt; J. M. Folster; M. Medrzycki; S. Jain; P. Wong; K. Goffard; D. Gieryn; J. Kahrs; K. Langolf; T. Zochert; J. E. Tate; C. H. Hsu; H. L. Kirking | Ford, et. al. Sep 2021 | Antigen Test Performance Among Children and Adults at a SARS-CoV-2 Community Testing Site | 2021-09-04 | BinaxNOW (Abbott) |  |  | Exposed |  |  |  |  | 79.8% |  | 99.7% |  | 98.5% | 94.9%
Terry | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Freire, et. al. Jun 2022 | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil | 2022-06-01 | COVID-19 Ag ECO Teste (Eco Diagnostica) |  |  | Overall | 27 | 3 | 15 | 36 | 42.9% | 30.5,56.0% | 83.3% | 58.6,96.4%
Terry | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Freire, et. al. Jun 2022 | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil | 2022-06-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 35 | 2 | 13 | 31 | 53% | 40.3,65.4% | 86.7% | 59.5,98.3%
Terry | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Freire, et. al. Jun 2022 | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil | 2022-06-03 | CORIS Bioconcept1 Ag-RDT (Nanosens) |  |  | Overall | 6 | 0 | 15 | 55 | 9.8% | 3.7,20.2% | 100% | 78.2,100%
Terry | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Freire, et. al. Jun 2022 | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil | 2022-06-04 | CELLER WONDFO SARSCOV2 Ag-RDT |  |  | Overall | 25 | 0 | 10 | 28 | 47.2% | 33.3,61.4% | 100% | 69.2,100%
Terry | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Freire, et. al. Jun 2022 | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil | 2022-06-05 | NowCheck COVID-19 Ag test (Bionote) |  |  | Overall | 33 | 0 | 9 | 22 | 60% | 45.9,73.0% | 100% | 66.4,100%
Terry | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Freire, et. al. Jun 2022 | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil | 2022-06-06 | Ag-RDT COVID-19 (Acro Biotech) |  |  | Overall | 43 | 2 | 11 | 10 | 81.1% | 68.0,90.6% | 84.6% | 54.5,98.1%
Terry | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Freire, et. al. Jun 2022 | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil | 2022-06-07 | PanBio (Abbott) |  |  | NP Swab | 33 | 0 | 10 | 22 | 60% | 45.9,73.0% | 100% | 69.2,100%
Terry | Sherli | 65 | M. L. Freire; L. L. Alves; C. S. de Souza; J. W. Saliba; V. Faria; M. J. Pedras; N. O. Carvalho; G. Q. Andrade; A. Rabello; D. M. Avelar; G. Cota | Freire, et. al. Jun 2022 | Performance differences among commercially available antigen rapid tests for COVID-19 in Brazil | 2022-06-08 | PanBio (Abbott) |  |  | NS | 32 | 0 | 10 | 23 | 58.2% | 44.1,71.4% | 100% | 69.2,100%
Terry | Sherli | 68 | R. M. Galliez; L. Bomfim; D. Mariani; I. C. Leitão; A. C. P. Castiñeiras; C. C. A. Gonçalves; B. Ortiz da Silva; P. H. Cardoso; M. B. Arruda; P. Alvarez; R. Brindeiro; V. A. Ota; D. G. M. Rodrigues; L. J. da Costa; O. D. C. Ferreira, Jr.; T. Castiñeiras; D. S. Faffe; A. Tanuri | Galliez, et. al. Jun 2022 | Evaluation of the Panbio COVID-19 Antigen Rapid Diagnostic Test in Subjects Infected with Omicron Using Different Specimens | 2022-06-01 | PanBio (Abbott) |  |  | Nasal specimen | 113 | 0 | 65 | 14 | 89 | 82.4,93.3 | 100 | 94.4,100 | 100 | 82.3 | PPV 95% CI 96.7, 100.0 NPV 95% CI 72.4, 89.1
Terry | Sherli | 68 | R. M. Galliez; L. Bomfim; D. Mariani; I. C. Leitão; A. C. P. Castiñeiras; C. C. A. Gonçalves; B. Ortiz da Silva; P. H. Cardoso; M. B. Arruda; P. Alvarez; R. Brindeiro; V. A. Ota; D. G. M. Rodrigues; L. J. da Costa; O. D. C. Ferreira, Jr.; T. Castiñeiras; D. S. Faffe; A. Tanuri | Galliez, et. al. Jun 2022 | Evaluation of the Panbio COVID-19 Antigen Rapid Diagnostic Test in Subjects Infected with Omicron Using Different Specimens | 2022-06-02 | PanBio (Abbott) |  |  | Oral Specimen | 16 | 0 | 65 | 111 | 12.6 | 7.9,19.5 | 100 | 94.4,100 | 100 | 36.9 | PPV 95% CI 80.6, 100.0 NPV 95% CI 30.2,44.3
Terry | Sherli | 63 | Garcia-Cardenas, Francisco; F. Alba; C. Ricardo; B. Jenny; V. Rafael; P. Fernando; F.-J. Emmanuel; C.-T. Alberto; M.-V. Alfredo; P. R.-G. Juan; H.-M. Alfredo; A. H. Luis | Garcia-Cardenas, et. al. May 2022a | Analytical performances of the COVISTIX and Panbio antigen rapid tests for SARS-CoV-2 detection in an unselected population (all commers) | 2022-05-01 | PanBio (Abbott) |  |  | Overall | 275 | 4 | 1755 | 168 | 62% | 58,64 | 99% | 99100 | 98.6 | 91.3
Terry | Sherli | 63 | Garcia-Cardenas, Francisco; F. Alba; C. Ricardo; B. Jenny; V. Rafael; P. Fernando; F.-J. Emmanuel; C.-T. Alberto; M.-V. Alfredo; P. R.-G. Juan; H.-M. Alfredo; A. H. Luis | Garcia-Cardenas, et. al. May 2022a | Analytical performances of the COVISTIX and Panbio antigen rapid tests for SARS-CoV-2 detection in an unselected population (all commers) | 2022-05-02 | COVISTIX Covid 19 Antigen Rapid Test Device |  |  | Overall | 205 | 21 | 508 | 49 | 81 | 76,85 | 96 | 94,98 | 90.7 | 91.2
Terry | Sherli | 63 | Garcia-Cardenas, Francisco; F. Alba; C. Ricardo; B. Jenny; V. Rafael; P. Fernando; F.-J. Emmanuel; C.-T. Alberto; M.-V. Alfredo; P. R.-G. Juan; H.-M. Alfredo; A. H. Luis | Garcia-Cardenas, et. al. May 2022a | Analytical performances of the COVISTIX and Panbio antigen rapid tests for SARS-CoV-2 detection in an unselected population (all commers) | 2022-05-03 | COVISTIX Covid 19 Antigen Rapid Test Device |  |  | Symptomatic <=7 days | 103 | 12 | 103 | 22 | 82 | 76,89 | 89.6 |  | 89.6 | 82.4
Terry | Sherli | 63 | Garcia-Cardenas, Francisco; F. Alba; C. Ricardo; B. Jenny; V. Rafael; P. Fernando; F.-J. Emmanuel; C.-T. Alberto; M.-V. Alfredo; P. R.-G. Juan; H.-M. Alfredo; A. H. Luis | Garcia-Cardenas, et. al. May 2022a | Analytical performances of the COVISTIX and Panbio antigen rapid tests for SARS-CoV-2 detection in an unselected population (all commers) | 2022-05-04 | COVISTIX Covid 19 Antigen Rapid Test Device |  |  | Symptomatic >7 days | 56 | 4 | 23 | 12 | 82 | 73,91 | 85.2 |  | 93.3 | 65.7
Terry | Sherli | 63 | Garcia-Cardenas, Francisco; F. Alba; C. Ricardo; B. Jenny; V. Rafael; P. Fernando; F.-J. Emmanuel; C.-T. Alberto; M.-V. Alfredo; P. R.-G. Juan; H.-M. Alfredo; A. H. Luis | Garcia-Cardenas, et. al. May 2022a | Analytical performances of the COVISTIX and Panbio antigen rapid tests for SARS-CoV-2 detection in an unselected population (all commers) | 2022-05-05 | COVISTIX Covid 19 Antigen Rapid Test Device |  |  | Asymptomatic | 46 | 5 | 382 | 15 | 75 | 65,86 | 98.7 |  | 90.2 | 96.2
Ellyn | Sherli | 69 | F. Garcia-Cardenas; F. Peñaloza; J. Bertin-Montoya; R. Valdéz-Vázquez; A. Franco; R. Cortés; E. Frias-Jimenez; A. Cedro-Tanda; A. Mendoza-Vargas; J. P. Reyes-Grajeda; A. Hidalgo-Miranda; L. A. Herrera | Garcia-Cardenas, et. al. May 2022b | Analytical Performances of the COVISTIX™ Antigen Rapid Test for SARS-CoV-2 Detection in an Unselected Population (All-Comers) | 2022-05-01 | COVISTIX Covid 19 Antigen Rapid Test Device |  | First cohort | Overall | 205 | 21 | 508 | 49 | 81 | 75.0,85.0 | 96 | 94.0,98.0
Ellyn | Sherli | 69 | F. Garcia-Cardenas; F. Peñaloza; J. Bertin-Montoya; R. Valdéz-Vázquez; A. Franco; R. Cortés; E. Frias-Jimenez; A. Cedro-Tanda; A. Mendoza-Vargas; J. P. Reyes-Grajeda; A. Hidalgo-Miranda; L. A. Herrera | Garcia-Cardenas, et. al. May 2022b | Analytical Performances of the COVISTIX™ Antigen Rapid Test for SARS-CoV-2 Detection in an Unselected Population (All-Comers) | 2022-05-02 | COVISTIX Covid 19 Antigen Rapid Test Device |  | First cohort | Ct≥34 | 8 |  |  | 17 | 32
Ellyn | Sherli | 69 | F. Garcia-Cardenas; F. Peñaloza; J. Bertin-Montoya; R. Valdéz-Vázquez; A. Franco; R. Cortés; E. Frias-Jimenez; A. Cedro-Tanda; A. Mendoza-Vargas; J. P. Reyes-Grajeda; A. Hidalgo-Miranda; L. A. Herrera | Garcia-Cardenas, et. al. May 2022b | Analytical Performances of the COVISTIX™ Antigen Rapid Test for SARS-CoV-2 Detection in an Unselected Population (All-Comers) | 2022-05-03 | COVISTIX Covid 19 Antigen Rapid Test Device |  | First cohort | 30≥Ct<34 | 21 |  |  | 12 | 64
Ellyn | Sherli | 69 | F. Garcia-Cardenas; F. Peñaloza; J. Bertin-Montoya; R. Valdéz-Vázquez; A. Franco; R. Cortés; E. Frias-Jimenez; A. Cedro-Tanda; A. Mendoza-Vargas; J. P. Reyes-Grajeda; A. Hidalgo-Miranda; L. A. Herrera | Garcia-Cardenas, et. al. May 2022b | Analytical Performances of the COVISTIX™ Antigen Rapid Test for SARS-CoV-2 Detection in an Unselected Population (All-Comers) | 2022-05-04 | COVISTIX Covid 19 Antigen Rapid Test Device |  | First cohort | 25≥Ct<30 | 53 |  |  | 13 | 80
Ellyn | Sherli | 69 | F. Garcia-Cardenas; F. Peñaloza; J. Bertin-Montoya; R. Valdéz-Vázquez; A. Franco; R. Cortés; E. Frias-Jimenez; A. Cedro-Tanda; A. Mendoza-Vargas; J. P. Reyes-Grajeda; A. Hidalgo-Miranda; L. A. Herrera | Garcia-Cardenas, et. al. May 2022b | Analytical Performances of the COVISTIX™ Antigen Rapid Test for SARS-CoV-2 Detection in an Unselected Population (All-Comers) | 2022-05-05 | COVISTIX Covid 19 Antigen Rapid Test Device |  | First cohort | 20≥Ct<25 | 70 |  |  | 5 | 93
Ellyn | Sherli | 69 | F. Garcia-Cardenas; F. Peñaloza; J. Bertin-Montoya; R. Valdéz-Vázquez; A. Franco; R. Cortés; E. Frias-Jimenez; A. Cedro-Tanda; A. Mendoza-Vargas; J. P. Reyes-Grajeda; A. Hidalgo-Miranda; L. A. Herrera | Garcia-Cardenas, et. al. May 2022b | Analytical Performances of the COVISTIX™ Antigen Rapid Test for SARS-CoV-2 Detection in an Unselected Population (All-Comers) | 2022-05-06 | COVISTIX Covid 19 Antigen Rapid Test Device |  | First cohort | <20 | 53 |  |  | 2 | 96
Ellyn | Sherli | 69 | F. Garcia-Cardenas; F. Peñaloza; J. Bertin-Montoya; R. Valdéz-Vázquez; A. Franco; R. Cortés; E. Frias-Jimenez; A. Cedro-Tanda; A. Mendoza-Vargas; J. P. Reyes-Grajeda; A. Hidalgo-Miranda; L. A. Herrera | Garcia-Cardenas, et. al. May 2022b | Analytical Performances of the COVISTIX™ Antigen Rapid Test for SARS-CoV-2 Detection in an Unselected Population (All-Comers) | 2022-05-07 | COVISTIX Covid 19 Antigen Rapid Test Device |  | Second cohort | Overall | 85 | 15 | 845 | 6 | 93 | 88,98 | 98 | 97,99
Terry | Sherli | 71 | S. García-Fernández; D. Pablo-Marcos; S. V. de la Fuente; M. J. R. Rodríguez; M. Gozalo; J. Rodríguez-Lozano; J. M. Méndez-Legaza; J. Calvo | García-Fernández, et. al. Mar 2022 | Evaluation of the rapid antigen detection test STANDARD F COVID-19 Ag FIA for diagnosing SARS-CoV-2: experience from an Emergency Department | 2022-03-01 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Overall |  |  |  |  | 84 | 76.1,89.7 | 99.6 | 98.5,99.9 | 98.1 | 96.4 | PPV 95% CI 92.7,99.7 NPV 95% CI 94.4,97.7
Terry | Sherli | 71 | S. García-Fernández; D. Pablo-Marcos; S. V. de la Fuente; M. J. R. Rodríguez; M. Gozalo; J. Rodríguez-Lozano; J. M. Méndez-Legaza; J. Calvo | García-Fernández, et. al. Mar 2022 | Evaluation of the rapid antigen detection test STANDARD F COVID-19 Ag FIA for diagnosing SARS-CoV-2: experience from an Emergency Department | 2022-03-02 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Ct </= 15 |  |  |  |  | 100
Terry | Sherli | 71 | S. García-Fernández; D. Pablo-Marcos; S. V. de la Fuente; M. J. R. Rodríguez; M. Gozalo; J. Rodríguez-Lozano; J. M. Méndez-Legaza; J. Calvo | García-Fernández, et. al. Mar 2022 | Evaluation of the rapid antigen detection test STANDARD F COVID-19 Ag FIA for diagnosing SARS-CoV-2: experience from an Emergency Department | 2022-03-03 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Ct 16-20 |  |  |  |  | 100
Terry | Sherli | 71 | S. García-Fernández; D. Pablo-Marcos; S. V. de la Fuente; M. J. R. Rodríguez; M. Gozalo; J. Rodríguez-Lozano; J. M. Méndez-Legaza; J. Calvo | García-Fernández, et. al. Mar 2022 | Evaluation of the rapid antigen detection test STANDARD F COVID-19 Ag FIA for diagnosing SARS-CoV-2: experience from an Emergency Department | 2022-03-04 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Ct 21-25 |  |  |  |  | 92.7
Terry | Sherli | 71 | S. García-Fernández; D. Pablo-Marcos; S. V. de la Fuente; M. J. R. Rodríguez; M. Gozalo; J. Rodríguez-Lozano; J. M. Méndez-Legaza; J. Calvo | García-Fernández, et. al. Mar 2022 | Evaluation of the rapid antigen detection test STANDARD F COVID-19 Ag FIA for diagnosing SARS-CoV-2: experience from an Emergency Department | 2022-03-05 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Ct 26-30 |  |  |  |  | 72.2
Terry | Sherli | 71 | S. García-Fernández; D. Pablo-Marcos; S. V. de la Fuente; M. J. R. Rodríguez; M. Gozalo; J. Rodríguez-Lozano; J. M. Méndez-Legaza; J. Calvo | García-Fernández, et. al. Mar 2022 | Evaluation of the rapid antigen detection test STANDARD F COVID-19 Ag FIA for diagnosing SARS-CoV-2: experience from an Emergency Department | 2022-03-06 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Ct >/=31 |  |  |  |  | 7.7
Terry | Sherli | 93 | García-Fiñana MartaHughes | García-Fiñana, et. al. Jul 2021 | Performance of the Innova SARS-CoV-2 antigen rapid lateral flow test in the Liverpool asymptomatic testing pilot: population based cohort study | 2021-06-01 | Innova (recalled 06/2021) |  |  | Overall | 28 | 3 | 5431 | 42 | 40 | 28.5,52.4 | 99.9 | 99.8,99.99 | 90.3 | 99.2 |  | 40 | 99.9448 | 90.3226 | 99.2326
Amy | Sherli | 72 | A. Goga; E. S. Mayne; K. Woeber; S. Takuva; D. Nsibande; M. Lekalakala; S. Jaumdally; P. Mutevedzi; H. Vreede; B. Daniels; C. Kufe; K. Dheda; K. Chetty; G. E. Gray | Goga, et. al. Mar 2022 | UTILITY of COVID-19 POINT-OF-CARE ANTIGEN TESTS in LOW-MIDDLE INCOME SETTINGS | 2022-03-01 | RapiGen (BioCredit) |  |  | Overall |  |  |  |  | 34.8 | 26.1,44.2 | 97.6 | 93.9,99.3 | 90.9 | 68.1 | PPV CI95 [78.3-97.5] NPV CI95 [61.7-74.0]
Amy | Sherli | 72 | A. Goga; E. S. Mayne; K. Woeber; S. Takuva; D. Nsibande; M. Lekalakala; S. Jaumdally; P. Mutevedzi; H. Vreede; B. Daniels; C. Kufe; K. Dheda; K. Chetty; G. E. Gray | Goga, et. al. Mar 2022 | UTILITY of COVID-19 POINT-OF-CARE ANTIGEN TESTS in LOW-MIDDLE INCOME SETTINGS | 2022-03-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall |  |  |  |  | 49.1 | 43.3,55.0 | 95.7 | 93.5,97.3 | 87.3 | 75.9 | PPV CI95 [81.2-91.9] NPV CI95 [72.4-79.2]
Amy | Sherli | 72 | A. Goga; E. S. Mayne; K. Woeber; S. Takuva; D. Nsibande; M. Lekalakala; S. Jaumdally; P. Mutevedzi; H. Vreede; B. Daniels; C. Kufe; K. Dheda; K. Chetty; G. E. Gray | Goga, et. al. Mar 2022 | UTILITY of COVID-19 POINT-OF-CARE ANTIGEN TESTS in LOW-MIDDLE INCOME SETTINGS | 2022-03-03 | SARS-CoV-2 Ag (LumiraDx) |  |  | Overall |  |  |  |  | 63.8 | 55.9,71.2 | 97 | 95.5,98.3 | 86 | 90.7 | PPV CI95 [78.5-91.6] NPV CI95 [88.2-92.9]
Chathu | Sherli | 73 | S. González-Fiallo; I. Mena-Rodríguez; P. Castro-Batista; V. M. Doeste-Hernández; V. Louit-Laborit | González-Fiallo, et. al. Apr 2022 | Evaluation of SARS-CoV-2 Rapid Antigen Tests in Use on the Isle of Youth, Cuba | 2022-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 70 | 27 | 658 | 23 | 75.3% | 66.0,84.6 | 96.1% | 94.5,97.6 | 72.2% | 96.6% | PPV 95%CI 62.7,81.6; NPV 95% CI 95.2,98.1
Terry | Sherli | 104 | Gupta A | Gupta, et. al. Feb 2021 | Rapid chromatographic immunoassay-based evaluation of COVID-19: A cross-sectional, diagnostic test accuracy study & its implications for COVID-19 management in India | 2021-02-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 63 | 1 | 252 | 14 | 81.8 | 71.3,89.6 | 99.6 | 97.8,99.9 | 98.4 | 94.7 |  | 81.8182 | 99.6047 | 98.4375 | 94.7368
Yati | Sherli | 109 | Harris David T. | Harris, et. al. May 2021 | SARS-CoV-2 Rapid Antigen Testing of Symptomatic and Asymptomatic Individuals on the University of Arizona Campus | 2021-05-01 | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | Asymptomatic | 6 | 7 | 1525 | 13 | 31.58 |  | 100 |  | 100 | 99.15 |  | 31.5789 | 99.5431 | 46.1538 | 99.1547
Yati | Sherli | 109 | Harris David T. | Harris, et. al. May 2021 | SARS-CoV-2 Rapid Antigen Testing of Symptomatic and Asymptomatic Individuals on the University of Arizona Campus | 2021-05-01 | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | Symptomatic | 251 | 7 | 573 | 54 | 82.23 |  | 98.79 |  | 97.29 | 91.39 |  | 82.2951 | 98.7931 | 97.2868 | 91.3876
Yati | Sherli | 109 | Harris David T. | Harris, et. al. May 2021 | SARS-CoV-2 Rapid Antigen Testing of Symptomatic and Asymptomatic Individuals on the University of Arizona Campus | 2021-05-01 | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | Overall | 257 | 14 | 2098 | 67 | 91.4 |  | 100 |  |  |  |  | 79.321 | 99.3371 | 94.8339 | 96.9053
Yati | Sherli | 114 | Holzner C. | Holzner, et. al. April 2021 | SARS-CoV-2 rapid antigen test: Fast-safe or dangerous? An analysis in the emergency department of an university hospital | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 379 | 8 | 1816 | 172 | 68.8 | 66.84,70.73 | 99.56 | 99.3,99.82 | 97.93 | 91.34 |  | 68.784 | 99.5614 | 97.9328 | 91.3481
Yati | Sherli | 114 | Holzner C. | Holzner, et. al. April 2021 | SARS-CoV-2 rapid antigen test: Fast-safe or dangerous? An analysis in the emergency department of an university hospital | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic |  |  |  |  | 62 | 61.68,62.32 | 97.63 | 96.6,98.66 | 97.63 | 97.63 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Yati | Sherli | 114 | Holzner C. | Holzner, et. al. April 2021 | SARS-CoV-2 rapid antigen test: Fast-safe or dangerous? An analysis in the emergency department of an university hospital | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct≤30 |  |  |  |  | 79.67 | 77.61,81.73 | 99.51 | 99.16,99.86 | 98.53 | 92.31 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Yati | Sherli | 114 | Holzner C. | Holzner, et. al. April 2021 | SARS-CoV-2 rapid antigen test: Fast-safe or dangerous? An analysis in the emergency department of an university hospital | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct≤30 |  |  |  |  | 90.9 | 88.94,92.86 | 99.61 |  | 90.9 | 99.61 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Yati | Sherli | 114 | Holzner C. | Holzner, et. al. April 2021 | SARS-CoV-2 rapid antigen test: Fast-safe or dangerous? An analysis in the emergency department of an university hospital | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptomatic |  |  |  |  | 69.46 | 67.16,71.76 | 99.51 | 99.17,99.85 | 98.58 | 87.09 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Yati | Sherli | 115 | Homza M. | Homza, et. al. Apr 2021 | Five Antigen Tests for SARS-CoV-2: Virus Viability Matters | 2021-04-01 | Ecotest (Assure Tech) |  |  | Overall |  | 7 |  | 26 | 75.7 | 66.5,83.5 | 96.7 | 93.3,98.7 | 92 | 88.7 |  | 0 | 0 | 0 | 0
Yati | Sherli | 115 | Homza M. | Homza, et. al. Apr 2021 | Five Antigen Tests for SARS-CoV-2: Virus Viability Matters | 2021-04-01 | Immupass VivaDiag (VivaChek Biotech) |  |  | Overall |  | 7 |  | 53 | 41.8 | 31.5,52.6 | 96 | 92.0,98.4 | 84.4 | 76.2 |  | 0 | 0 | 0 | 0
Yati | Sherli | 115 | Homza M. | Homza, et. al. Apr 2021 | Five Antigen Tests for SARS-CoV-2: Virus Viability Matters | 2021-04-01 | ND Covid (NDFOS) |  |  | Overall |  | 2 |  | 23 | 70.1 | 58.6,80 | 56.1 | 46.5,65.4 | 51.9 | 73.6 |  | 0 | 0 | 0 | 0
Yati | Sherli | 115 | Homza M. | Homza, et. al. Apr 2021 | Five Antigen Tests for SARS-CoV-2: Virus Viability Matters | 2021-04-01 | SARS-CoV-2 Antigen Rapid Test (JoysBio) |  |  | Overall |  | 2 |  | 38 | 57.8 | 46.9,68.1 | 98.5 | 94.8,99.8 | 96.3 | 77.8 |  | 0 | 0 | 0 | 0
Yati | Sherli | 115 | Homza M. | Homza, et. al. Apr 2021 | Five Antigen Tests for SARS-CoV-2: Virus Viability Matters | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall |  | 1 |  | 16 | 61.9 | 45.6,76.4 | 99 | 94.4,100 | 96.3 | 85.7 |  | 0 | 0 | 0 | 0
Chathu | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Hörber, et. al. Jun 2022 | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay | 2022-06-01 | CoV2Ag assay (Siemens Healthineers, Eschborn, Germany) |  | First cohort | Overall | 207 | 1 | 212 | 27 | 88.5% | 83.7,91.9 | 99.5% | 97.4,99.9
Chathu | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Hörber, et. al. Jun 2022 | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay | 2022-06-02 | CoV2Ag assay (Siemens Healthineers, Eschborn, Germany) |  | Second cohort | Ct<30 |  |  |  |  | 94.4 | 90.4,96.7
Chathu | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Hörber, et. al. Jun 2022 | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay | 2022-06-03 | CoV2Ag assay (Siemens Healthineers, Eschborn, Germany) |  | Second cohort | Ct<27 |  |  |  |  | 97.4 | 94.1,98.9
Chathu | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Hörber, et. al. Jun 2022 | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay | 2022-06-04 | CoV2Ag assay (Siemens Healthineers, Eschborn, Germany) |  | Second cohort | Ct<25 |  |  |  |  | 98.3 | 95.1,99.4
Chathu | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Hörber, et. al. Jun 2022 | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay | 2022-06-05 | CoV2Ag assay (Siemens Healthineers, Eschborn, Germany) |  | Second cohort | Ct<23 |  |  |  |  | 98.8 | 95.6,99.7
Chathu | Sherli | 83 | S. Hörber; C. Drees; T. Ganzenmueller; K. Schmauder; S. Peter; D. Biskup; A. Peter | Hörber, et. al. Jun 2022 | Evaluation of a laboratory-based high-throughput SARS-CoV-2 antigen assay | 2022-06-06 | CoV2Ag assay (Siemens Healthineers, Eschborn, Germany) |  | Second cohort | Ct<20 |  |  |  |  | 100 | 96.7,100
Yati - done | Sherli | 86 | M. Ifko; Ž. Tkalčić Švabek; I. Friščić; M. M. Kardum Paro; I. Prkačin; L. Đerek; A. Livun; M. Skvarč | Ifko, et. al. July 2021 | Diagnostic validation of two SARS-CoV-2 immunochromatographic tests in Slovenian and Croatian hospitals | 2021-06-01 | NADAL COVID-19 Antigen Rapid Test (New Art Laboratories/nal von minden) |  | Croatia cohort | Overall | 11 | 0 | 38 | 2 | 84.61 | 54.55,98.08 | 100 | 90.75,100 | 100 | 95 | NPV 95%CI 84.15,98.55
Terry - finished. Need some clarification from authors on data. | Sherli | 85 | M. Ifko; M. Skvarc | Ifko, et. al. July 2021 | Use of Immunochromatographic SARS-CoV-2 Antigen Testing in Eight Long-Term Care Facilities for the Elderly | 2021-07-01 | NADAL COVID-19 Antigen Rapid Test (New Art Laboratories/nal von minden) |  |  | Overall | 20 | 12 | 90 | 3 | 86.96 | 66.41,97.23 | 88.24 | 80.35,93.77 | 62.5 | 96.77 | PPV 95% CI: 48.91,74.37 NPV 95% CI 91.24,96.86
Ellyn | Sherli | 119 | Igloi Z. | Iglòi, et. al. May 2021 | Clinical Evaluation of Roche SD Biosensor Rapid Antigen Test for SARS-CoV-2 in Municipal Health Service Testing Site, the Netherlands | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 158 | 4 | 780 | 28 | 84.9 | 79.1,89.4 | 99.5 | 98.7,99.8 | 97.5 | 96.5 |  | 84.9462 | 99.4898 | 97.5309 | 96.5347
Ellyn | Sherli | 119 | Igloi Z. | Iglòi, et. al. May 2021 | Clinical Evaluation of Roche SD Biosensor Rapid Antigen Test for SARS-CoV-2 in Municipal Health Service Testing Site, the Netherlands | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct≤25 |  |  |  |  | 99.1 | 95.2,100 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 119 | Igloi Z. | Iglòi, et. al. May 2021 | Clinical Evaluation of Roche SD Biosensor Rapid Antigen Test for SARS-CoV-2 in Municipal Health Service Testing Site, the Netherlands | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct≤30 |  |  |  |  | 94.3 | 89.6,97.0 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 123 | Jakobsen K. | Jakobsen, et. al. June 2021 | Accuracy and cost description of rapid antigen test compared with reverse transcriptase-polymerase chain reaction for SARS-CoV-2 detection | 2021-06-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 154 | 23 | 4567 | 67 | 69.7 |  | 99.5 |  | 87 | 98.6 |  | 69.6833 | 99.4989 | 87.0056 | 98.5542
Ellyn | Sherli | 123 | Jakobsen K. | Jakobsen, et. al. June 2021 | Accuracy and cost description of rapid antigen test compared with reverse transcriptase-polymerase chain reaction for SARS-CoV-2 detection | 2021-06-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic | 29 | 11 | 2938 | 30 | 49.2 |  | 99.6 |  | 72.5 | 98.9 |  | 49.1525 | 99.627 | 72.5 | 98.9892
Ellyn | Sherli | 123 | Jakobsen K. | Jakobsen, et. al. June 2021 | Accuracy and cost description of rapid antigen test compared with reverse transcriptase-polymerase chain reaction for SARS-CoV-2 detection | 2021-06-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Syptomatic | 67 | 7 | 613 | 18 | 78.8 |  | 98.8 |  | 90.5 | 97.1 |  | 78.8235 | 98.871 | 90.5405 | 97.1474
Terry | Sherli | 92 | K. K. Jakobsen; J. S. Jensen; T. Todsen; N. Kirkby; F. Lippert; A. M. Vangsted; M. Klokker; C. von Buchwald | Jakobsen, et. al. Oct 2021 | Accuracy of anterior nasal swab rapid antigen tests compared with RT-PCR for massive SARS-CoV-2 screening in low prevalence population | 2021-10-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 32 | 0 | 7008 | 34 | 48.5 |  | 100 |  | 100 | 99.5
Terry | Sherli | 92 | K. K. Jakobsen; J. S. Jensen; T. Todsen; N. Kirkby; F. Lippert; A. M. Vangsted; M. Klokker; C. von Buchwald | Jakobsen, et. al. Oct 2021 | Accuracy of anterior nasal swab rapid antigen tests compared with RT-PCR for massive SARS-CoV-2 screening in low prevalence population | 2021-10-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct<30 | 23 | 0 | 7008 | 13 | 63.9 |  | 100 |  | 100 | 99.8
Terry | Sherli | 92 | K. K. Jakobsen; J. S. Jensen; T. Todsen; N. Kirkby; F. Lippert; A. M. Vangsted; M. Klokker; C. von Buchwald | Jakobsen, et. al. Oct 2021 | Accuracy of anterior nasal swab rapid antigen tests compared with RT-PCR for massive SARS-CoV-2 screening in low prevalence population | 2021-10-03 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct<33 | 27 | 0 | 7008 | 21 | 56.2 |  | 100 |  | 100 | 99.7
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Jeewandara, et. al. Mar 2022 | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection | 2022-03-01 | Standard Q COVID-19 Ag (SD Biosensor) |  | 4786 participants | Overall | 316 | 94 | 3755 | 556 | 36.24 | 33.1,39.5 | 97.6 | 97,98 | 77 | 87.1
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Jeewandara, et. al. Mar 2022 | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection | 2022-03-02 | PanBio (Abbott) |  | 3325 participants | Overall | 151 | 13 | 3275 | 136 | 52.6 | 46.7,58.5 | 99.6 | 99.2,99.8 | 92.1 | 95.7
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Jeewandara, et. al. Mar 2022 | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection | 2022-03-03 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT<25 | 242 |  |  | 73
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Jeewandara, et. al. Mar 2022 | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection | 2022-03-04 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT 26-30 | 61 |  |  | 197
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Jeewandara, et. al. Mar 2022 | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection | 2022-03-05 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT 31-38 | 14 |  |  | 277
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Jeewandara, et. al. Mar 2022 | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection | 2022-03-06 | PanBio (Abbott) |  |  | CT <25 | 129 |  |  | 29
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Jeewandara, et. al. Mar 2022 | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection | 2022-03-07 | PanBio (Abbott) |  |  | CT 26-30 | 18 |  |  | 76
Amy | Sherli | 93 | C. Jeewandara; D. Guruge; P. D. Pushpakumara; D. Madhusanka; T. T. Jayadas; I. P. Chaturanga; I. S. Aberathna; S. Danasekara; T. Pathmanathan; D. Jayathilaka; G. Somathilaka; H. Kuruppu; L. Gomes; V. Gunasekara; R. Wijayamuni; G. S. Ogg; G. N. Malavige | Jeewandara, et. al. Mar 2022 | Sensitivity and specificity of two WHO approved SARS-CoV2 antigen assays in detecting patients with SARS-CoV2 infection | 2022-03-08 | PanBio (Abbott) |  |  | CT 31-38 | 3 |  |  | 31
Ellyn | Sherli | 126 | Jegerlehner Sabrina | Jegerlehner, et. al. July 2021 | Diagnostic accuracy of a SARS-CoV-2 rapid antigen test in real-life clinical settings | 2021-07-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 92 | 2 | 1319 | 49 | 65.3 | 56.8,73.1 | 99.9 | 99.5,100 | 97.9 | 96.4 |  | 65.2482 | 99.8486 | 97.8723 | 96.4181
Ellyn | Sherli | 126 | Jegerlehner Sabrina | Jegerlehner, et. al. July 2021 | Diagnostic accuracy of a SARS-CoV-2 rapid antigen test in real-life clinical settings | 2021-07-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic |  |  |  |  | 44 | 24.4,65.1 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Jegerlehner, et. al. Mar 2022 | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-01 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | Overall | 16 | 6 | 317 | 37 | 30.2 | 18.3,44.3 | 98.1 | 96.0,99.3 | 72.7 | 89.5
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Jegerlehner, et. al. Mar 2022 | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-02 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT 40 |  |  |  |  | 30.2 |  | 69.8
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Jegerlehner, et. al. Mar 2022 | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-03 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT 38 |  |  |  |  | 30.2 |  | 69.8
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Jegerlehner, et. al. Mar 2022 | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-04 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT 36 |  |  |  |  | 30.8 |  | 69.2
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Jegerlehner, et. al. Mar 2022 | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-05 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT 34 |  |  |  |  | 30.8 |  | 69.2
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Jegerlehner, et. al. Mar 2022 | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-06 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT 32 |  |  |  |  | 32 |  | 68
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Jegerlehner, et. al. Mar 2022 | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-07 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT 30 |  |  |  |  | 33.3 |  | 66.7
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Jegerlehner, et. al. Mar 2022 | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-08 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT 28 |  |  |  |  | 33.3 |  | 66.7
Tim | Sherli | 95 | S. Jegerlehner; F. Suter-Riniker; P. Jent; P. Bittel; M. Nagler | Jegerlehner, et. al. Mar 2022 | Diagnostic accuracy of SARS-CoV-2 saliva antigen testing in a real-life clinical setting | 2022-03-09 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | CT 26 |  |  |  |  | 33.3 |  | 66.7
Terry | Sherli | 98 | S. Jirungda; P. Khamphrom; K. Klaprajun; T. Prachakul; J. Kraisin | Jirungda, et. al. May 2022 | Clinical performance of the standard F COVID-19 AG FIA for the detection of SARS-COV-2 infection | 2022-05-01 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Overall |  |  |  |  | 98.8 |  | 89.7 |  | 94.6 | 97.7
Terry | Sherli | 102 | M. Kahn; L. Schuierer; C. Bartenschlager; S. Zellmer; R. Frey; M. Freitag; C. Dhillon; M. Heier; A. Ebigbo; C. Denzel; S. Temizel; H. Messmann; M. Wehler; R. Hoffmann; E. Kling; C. Römmele | Kahn, et. al. Aug 2021 | Performance of antigen testing for diagnosis of COVID-19: a direct comparison of a lateral flow device to nucleic acid amplification based tests | 2021 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Overall | 57 | 31 | 2983 | 39 | 59.4 |  | 99 |  | 64.8 | 98.7
Terry | Sherli | 102 | M. Kahn; L. Schuierer; C. Bartenschlager; S. Zellmer; R. Frey; M. Freitag; C. Dhillon; M. Heier; A. Ebigbo; C. Denzel; S. Temizel; H. Messmann; M. Wehler; R. Hoffmann; E. Kling; C. Römmele | Kahn, et. al. Aug 2021 | Performance of antigen testing for diagnosis of COVID-19: a direct comparison of a lateral flow device to nucleic acid amplification based tests | 2021 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Asymptomatic | 20 | 25 | 2273 | 22 | 47.6 |  | 98.9 |  | 44.4 | 99
Terry | Sherli | 102 | M. Kahn; L. Schuierer; C. Bartenschlager; S. Zellmer; R. Frey; M. Freitag; C. Dhillon; M. Heier; A. Ebigbo; C. Denzel; S. Temizel; H. Messmann; M. Wehler; R. Hoffmann; E. Kling; C. Römmele | Kahn, et. al. Aug 2021 | Performance of antigen testing for diagnosis of COVID-19: a direct comparison of a lateral flow device to nucleic acid amplification based tests | 2021 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Symptomatic | 34 | 6 | 192 | 17 | 66.7 |  | 97 |  | 85 | 91.9
Terry/Amy | Sherli | 106 | H. H. Kessler; F. Prüller; M. Hardt; E. Stelzl; E. Föderl-Höbenreich; S. Pailer; A. Lueger; P. Kreuzer; K. Zatloukal; M. Herrmann | Kessler, et. al. Mar 2022 | Identification of contagious SARS-CoV-2 infected individuals by Roche's Rapid Antigen Test | 2022-03-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Central Lab | 12 | 0 | 40 | 15
Ellyn | Sherli | 136 | Kim D.Lee | Kim, et. al. Apr 2021 | Development and Clinical Evaluation of an Immunochromatography-Based Rapid Antigen Test (GenBody™ COVAG025) for COVID-19 Diagnosis | 2021-04-01 | GenBody COVAG025 (GenBody) |  |  | Overall | [11 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 136 | Kim D.Lee | Kim, et. al. Apr 2021 | Development and Clinical Evaluation of an Immunochromatography-Based Rapid Antigen Test (GenBody™ COVAG025) for COVID-19 Diagnosis | 2021-04-01 | GenBody COVAG025 (GenBody) |  |  | Prospective | 94 | 0 | 100 | 6 | 94 | 87.4,97.77 | 100 | 96.38,100 | 100 | 94.3 |  | 94 | 100 | 100 | 94.3396
Ellyn | Sherli | 136 | Kim D.Lee | Kim, et. al. Apr 2021 | Development and Clinical Evaluation of an Immunochromatography-Based Rapid Antigen Test (GenBody™ COVAG025) for COVID-19 Diagnosis | 2021-04-01 | GenBody COVAG025 (GenBody) |  |  | Retrospective | 27 | 2 | 98 | 3 | 90 | 73.47,97.89 | 98 | 92.96,99.76 | 93.1 | 97.03 |  | 90 | 98 | 93.1034 | 97.0297
Terry | Sherli | 109 | C. W. E. L.-d. H. M. S. R. F. C. P. R. v. K. S. King | King, et. al. Sep 2021 | Validation of the Panbio™ COVID-19 Antigen Rapid Test (Abbott) to screen for SARS-CoV-2 infection in Sint Maarten: a diagnostic accuracy study | Pre-print Sep 2021 | PanBio (Abbott) |  |  | Overall | 100 | 1 | 1291 | 19 | 84 | 76.2,90.1 | 99.9 | 99.6,100 | 99.9 | 98.6 | PPV 95%CI 93.4, 99.9; NPV 95%CI 97.8,99
Terry | Sherli | 109 | C. W. E. L.-d. H. M. S. R. F. C. P. R. v. K. S. King | King, et. al. Sep 2021 | Validation of the Panbio™ COVID-19 Antigen Rapid Test (Abbott) to screen for SARS-CoV-2 infection in Sint Maarten: a diagnostic accuracy study | Pre-print Sep 2021 | PanBio (Abbott) |  |  | Symptomatic |  |  |  |  | 90.5 | 82.1,95.8 | 100 | 99.2,100 | 100 | 98.3 | NPV 95%CI 96.8,99.1
Terry | Sherli | 109 | C. W. E. L.-d. H. M. S. R. F. C. P. R. v. K. S. King | King, et. al. Sep 2021 | Validation of the Panbio™ COVID-19 Antigen Rapid Test (Abbott) to screen for SARS-CoV-2 infection in Sint Maarten: a diagnostic accuracy study | Pre-print Sep 2021 | PanBio (Abbott) |  |  | Asymptomatic |  |  |  |  | 67.7 | 49.5,82.6 | 99.9 | 99.3,100 | 95.8 | 98.7 | PPV 95% CI 76.2,99.4; NPV 95% CI 97.8,99.2
Terry | Sherli | 109 | C. W. E. L.-d. H. M. S. R. F. C. P. R. v. K. S. King | King, et. al. Sep 2021 | Validation of the Panbio™ COVID-19 Antigen Rapid Test (Abbott) to screen for SARS-CoV-2 infection in Sint Maarten: a diagnostic accuracy study | Pre-print Sep 2021 | PanBio (Abbott) |  |  | Ct< 35 |  |  |  |  | 73.5 | 65.3,80.7 | 99.9 | 99.6,100
Terry | Sherli | 109 | C. W. E. L.-d. H. M. S. R. F. C. P. R. v. K. S. King | King, et. al. Sep 2021 | Validation of the Panbio™ COVID-19 Antigen Rapid Test (Abbott) to screen for SARS-CoV-2 infection in Sint Maarten: a diagnostic accuracy study | Pre-print Sep 2021 | PanBio (Abbott) |  |  | Ct<30 |  |  |  |  | 90.7 | 83.6,95.5 | 99.8 | 99.3,100
Terry | Sherli | 109 | C. W. E. L.-d. H. M. S. R. F. C. P. R. v. K. S. King | King, et. al. Sep 2021 | Validation of the Panbio™ COVID-19 Antigen Rapid Test (Abbott) to screen for SARS-CoV-2 infection in Sint Maarten: a diagnostic accuracy study | Pre-print Sep 2021 | PanBio (Abbott) |  |  | Ct< 25 |  |  |  |  | 97.7 | 91.9,99.7 | 98.7 | 98,99.3
Terry | Sherli | 110 | Y. Kiyasu; Y. Takeuchi; Y. Akashi; D. Kato; M. Kuwahara; S. Muramatsu; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kiyasu, et. al. Jul 2021 | Prospective analytical performance evaluation of the QuickNavi™-COVID19 Ag for asymptomatic individuals | 2021-07-01 | QuickNavi-COVID19 Ag |  |  | asymptomatic | 51 | 0 | 997 | 25 | 67.1 | 55.4,77.5 | 100 | 99.4,100 | 100 | 97.6 | PPV 95% CI 89.7,100; NPV 95% CI 96.4,98.4
Terry | Sherli | 110 | Y. Kiyasu; Y. Takeuchi; Y. Akashi; D. Kato; M. Kuwahara; S. Muramatsu; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kiyasu, et. al. Jul 2021 | Prospective analytical performance evaluation of the QuickNavi™-COVID19 Ag for asymptomatic individuals | 2021-07-02 | QuickNavi-COVID19 Ag |  |  | Symptomatic | 100 | 0 | 749 | 12 | 89.3 | 82,94.3 | 100 | 99.3,100 | 100 | 98.4 | PPV 95% CI 94.6,100; NPV 95% CI 97.3,99.2
Terry | Sherli | 110 | Y. Kiyasu; Y. Takeuchi; Y. Akashi; D. Kato; M. Kuwahara; S. Muramatsu; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kiyasu, et. al. Jul 2021 | Prospective analytical performance evaluation of the QuickNavi™-COVID19 Ag for asymptomatic individuals | 2021-07-03 | QuickNavi-COVID19 Ag |  |  | Overall | 151 | 0 | 1746 | 37 | 80.3 | 73.9,85.7 | 100 | 99.7,100
Terry | Sherli | 111 | A. Klajmon; A. Olechowska-Jarząb; D. Salamon; A. Sroka-Oleksiak; M. Brzychczy-Włoch; T. Gosiewski | Klajmon, et. al. Dec 2021 | Comparison of antigen tests and qpcr in rapid diagnostics of infections caused by sars-cov-2 virus | 2021-12-01 | Humasis COVID-19 Ag Test kit (Humasis Co., Ltd.) |  |  | Overall | 43 | 2 | 140 | 4 | 91.49 | 79.62,97.63 | 97.9 | 93.99,99.57 | 93.48 | 97.22 | PPV 95% CI 82.10,98.63; NPV 95% CI 93.04,99.24
Ellyn | Sherli | 144 | Klein J. | Klein, et. al. May 2021 | Head-to-head performance comparison of self-collected nasal versus professional-collected nasopharyngeal swab for a WHO-listed SARS-CoV-2 antigen-detecting rapid diagnostic test | 2021-05-01 | PanBio (Abbott) |  |  | NMT |  |  |  |  | 84.4 | 71.2,92.3 | 99.2 | 97.1,99.8 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 144 | Klein J. | Klein, et. al. May 2021 | Head-to-head performance comparison of self-collected nasal versus professional-collected nasopharyngeal swab for a WHO-listed SARS-CoV-2 antigen-detecting rapid diagnostic test | 2021-05-01 | PanBio (Abbott) |  |  | NP |  |  |  |  | 88.9 | 76.5,95.5 | 99.2 | 97.1,99.8 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 148 | Kohmer N. | Kohmer, et. al. Jan 2021 | Article the comparative clinical performance of four SARS-CoV-2 rapid antigen tests and their correlation to infectivity in vitro | 2021-01-01 | SARS-CoV-2 Ag (LumiraDx) |  |  | Overall |  |  |  |  | 50 | 38.1,61.9 | 100 | 86.8,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 148 | Kohmer N. | Kohmer, et. al. Jan 2021 | Article the comparative clinical performance of four SARS-CoV-2 rapid antigen tests and their correlation to infectivity in vitro | 2021-01-01 | NADAL COVID-19 Antigen Rapid Test (New Art Laboratories/nal von minden) |  |  | Overall |  |  |  |  | 24.3 | 15.1,35.7 | 100 | 86.8,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 148 | Kohmer N. | Kohmer, et. al. Jan 2021 | Article the comparative clinical performance of four SARS-CoV-2 rapid antigen tests and their correlation to infectivity in vitro | 2021-01-01 | Rida Quick SARS-CoV-2 (R-Biopharm) |  |  | Overall |  |  |  |  | 39.2 | 28,51.2 | 96.2 | 80.4,99.9 | PPV - not provided; NPV - not provided |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 148 | Kohmer N. | Kohmer, et. al. Jan 2021 | Article the comparative clinical performance of four SARS-CoV-2 rapid antigen tests and their correlation to infectivity in vitro | 2021-01-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall |  |  |  |  | 43.2 | 37.8,55.3 | 100 | 86.8,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 152 | Korenkov M. | Korenkov, et. al. Aug 2021 | Evaluation of a rapid antigen test to detect SARS-CoV-2 infection and identify potentially infectious individuals | 2021-08-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 29 | 25 | 64 | 0 | 100 | 88.3,100 | 71.91 | 61.82,80.20 | 53.7 | 100 |  | 100 | 71.9101 | 53.7037 | 100
Yati | Sherli | 113 | M. P. N. M. M. V. K. E. R. W. M. F. I. D. F. G. L. L. C. F. G. G. H. P. N. H. E. K. F. Korenkov | Korenkov, et. al. May 2021 | Assessment of SARS-CoV-2 infectivity by a Rapid Antigen Detection Test | Preprint May 2021 | Standard Q COVID-19 Ag (SD Biosensor) |  | From table 1 | Overall | 90 | 2 | 1816 | 120 | 42.86 |  |  | 99.89
Yati | Sherli | 113 | M. P. N. M. M. V. K. E. R. W. M. F. I. D. F. G. L. L. C. F. G. G. H. P. N. H. E. K. F. Korenkov | Korenkov, et. al. May 2021 | Assessment of SARS-CoV-2 infectivity by a Rapid Antigen Detection Test | Preprint May 2021 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct<20 | 14 | 0 | 0 | 0 | 100
Yati | Sherli | 113 | M. P. N. M. M. V. K. E. R. W. M. F. I. D. F. G. L. L. C. F. G. G. H. P. N. H. E. K. F. Korenkov | Korenkov, et. al. May 2021 | Assessment of SARS-CoV-2 infectivity by a Rapid Antigen Detection Test | Preprint May 2021 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct<25 | 56 | 0 | 0 | 1 | 98.25
Yati | Sherli | 113 | M. P. N. M. M. V. K. E. R. W. M. F. I. D. F. G. L. L. C. F. G. G. H. P. N. H. E. K. F. Korenkov | Korenkov, et. al. May 2021 | Assessment of SARS-CoV-2 infectivity by a Rapid Antigen Detection Test | Preprint May 2021 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct<30 | 78 | 0 | 0 | 10 | 88.64
Yati | Sherli | 113 | M. P. N. M. M. V. K. E. R. W. M. F. I. D. F. G. L. L. C. F. G. G. H. P. N. H. E. K. F. Korenkov | Korenkov, et. al. May 2021 | Assessment of SARS-CoV-2 infectivity by a Rapid Antigen Detection Test | Preprint May 2021 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct<35 | 89 | 0 | 0 | 87 | 50.57
Yati | Sherli | 113 | M. P. N. M. M. V. K. E. R. W. M. F. I. D. F. G. L. L. C. F. G. G. H. P. N. H. E. K. F. Korenkov | Korenkov, et. al. May 2021 | Assessment of SARS-CoV-2 infectivity by a Rapid Antigen Detection Test | Preprint May 2021 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptomatic | 72 | 1 | 735 | 58 | 55.38
Yati | Sherli | 113 | M. P. N. M. M. V. K. E. R. W. M. F. I. D. F. G. L. L. C. F. G. G. H. P. N. H. E. K. F. Korenkov | Korenkov, et. al. May 2021 | Assessment of SARS-CoV-2 infectivity by a Rapid Antigen Detection Test | Preprint May 2021 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic | 18 | 1 | 729 | 62 | 22.5
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-09 | SARS-CoV-2 Ag (LumiraDx) |  |  | Ct>=30 | 10 |  |  | 14 | 41.7 | 24.5,61.2
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-10 | SARS-CoV-2 Ag (LumiraDx) |  |  | Ct<25 | 88 |  |  | 7 | 92.6 | 85.6,96.4
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-11 | SARS-CoV-2 Ag (LumiraDx) |  |  | Ct>=25 | 32 |  |  | 19 | 62.7 | 49,74.7
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-01 | SARS-CoV-2 Ag (LumiraDx) |  | n=761 | Overall | 120 | 4 | 611 | 26 | 82.2 | 75.2,87.5 | 99.3 | 97.9,99.7
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-02 | SARS-CoV-2 Ag (LumiraDx) |  | n=488 | Heidelberg | 55 | 3 | 420 | 10 | 84.6 | 7.9,91.4 | 99.3 | 97.9,99.7
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-03 | SARS-CoV-2 Ag (LumiraDx) |  | n=273 | Berlin | 65 | 1 | 191 | 16 | 80.2 | 70.3,87.5 | 99.5 | 97.1,100
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-04 | SARS-CoV-2 Ag (LumiraDx) |  |  | 0-7 days | 102 | 2 | 303 | 16 | 86.4 | 79.1,91.5
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-05 | SARS-CoV-2 Ag (LumiraDx) |  |  | 8-14 days | 7 | 0 | 26 | 6 | 53.8 | 29.1,76.8
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-06 | SARS-CoV-2 Ag (LumiraDx) |  |  | Symptomatic | 113 | 3 | 346 | 24 | 82.5 | 75.3,87.9
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-07 | SARS-CoV-2 Ag (LumiraDx) |  |  | Asymptomatic | 7 | 1 | 262 | 2 | 77.8 | 45.3,93.7
Chathu | Sherli | 116 | L. J. Krüger; J. A. F. Klein; F. Tobian; M. Gaeddert; F. Lainati; S. Klemm; P. Schnitzler; R. Bartenschlager; B. Cerikan; C. J. Neufeldt; O. Nikolai; A. K. Lindner; F. P. Mockenhaupt; J. Seybold; T. C. Jones; V. M. Corman; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. A. Sacks; C. M. Denkinger | Krüger, et. al. Aug 2021 | Evaluation of accuracy, exclusivity, limit-of-detection and ease-of-use of LumiraDx™: An antigen-detecting point-of-care device for SARS-CoV-2 | 2021-08-08 | SARS-CoV-2 Ag (LumiraDx) |  |  | Ct<30 | 110 |  |  | 12 | 90.2 | 83.6,94.3
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-01 | RapiGen (BioCredit) |  |  | Overall | 100 | 4 | 1569 | 42 | 70.4 | 62.4,77.3 | 99.7 | 99.3,99.9 |  | 93
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-02 | RapiGen (BioCredit) |  |  | Brazil | 87 | 4 | 355 | 30 | 74.4 | 65.8,81.4 | 98.9 | 97.2,99.6
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-03 | RapiGen (BioCredit) |  |  | Germany | 13 | 0 | 1214 | 12 | 52 | 33.5,70 | 100 | 99.7,100
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-04 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Overall | 120 | 27 | 943 | 39 | 75.5 | 68.2,81.5 | 97.2 | 96.0,98.1
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-05 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Brazil | 93 | 7 | 326 | 27 | 77.5 | 69.2,84.1 | 97.9 | 95.7,99
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-06 | STANDARD F COVID-19 Ag FIA (SD Biosensor Inc.) |  |  | Germany | 27 | 20 | 617 | 12 | 69.2 | 53.6,81.4 | 96.9 | 95.2,98
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-07 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 190 | 18 | 1860 | 42 | 81.9 | 76.4,86.3 | 99 | 98.5,99.4
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-08 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Brazil | 94 | 7 | 287 | 12 | 88.7 | 81.2,93.4 | 97.6 | 95.2,98.8
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-09 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Germany | 96 | 11 | 1573 | 30 | 76.2 | 68.0,82.8 | 99.3 | 98.8,99.6
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-10 | Fluorecare (Colloidal Gold/Fluorescent) SARS-CoV-2 Spike Protein Test kit (Shenzen Microprofit) |  | Bioeasy 2019-nCoV Ag Fluorescence Rapid Test Kit (Shenzhen Bioeasy Biotechnology Co. LTD) | Germany | 10 | 49 | 663 | 5 | 66.7 | 41.7,84.8 | 93.1 | 91.0,94.8
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-11 | Espline SARS-CoV-2 (Fujirebio) |  |  | Germany | 89 | 0 | 611 | 23 | 79.5 | 71.1,85.9 | 100 | 99.4,100
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-12 | Mologic Covid-19 Rapid Antigen Test (Mologic Ltd. United Kingdom) |  |  | Germany | 173 | 0 | 457 | 19 | 90.1 | 85.1,93.6 | 100 | 99.2,100
Terry | Sherli | 117 | L. J. Krüger; A. Tanuri; A. K. Lindner; M. Gaeddert; L. Köppel; F. Tobian; L. E. Brümmer; J. A. F. Klein; F. Lainati; P. Schnitzler; O. Nikolai; F. P. Mockenhaupt; J. Seybold; V. M. Corman; T. C. Jones; C. Drosten; C. Gottschalk; S. F. Weber; S. Weber; O. C. Ferreira; D. Mariani; E. R. Dos Santos Nascimento; T. M. Pereira Pinto Castineiras; R. M. Galliez; D. S. Faffe; I. C. Leitão; C. Dos Santos Rodrigues; T. S. Frauches; K. Nocchi; N. M. Feitosa; S. S. Ribeiro; N. R. Pollock; B. Knorr; A. Welker; M. de Vos; J. Sacks; S. Ongarello; C. M. Denkinger | Krüger, et. al. Dec 2021 | Accuracy and ease-of-use of seven point-of-care SARS-CoV-2 antigen-detecting tests: A multi-centre clinical evaluation | 2021-12-13 | NowCheck COVID-19 Ag test (Bionote) |  |  | Brazil | 91 | 8 | 290 | 11 | 89.2 | 81.7,93.9 | 97.3 | 94.8,98.6
Yati | Sherli | 118 | L. J. L. A. K. Krüger; M. T. F. K. J. S. S. L. F. S. P. N. O. M. F. P. Gaeddert; J. C. V. M. Seybold; T. C. Jones; N. R. Pollock; B. W. A. W. S. S. N. S. J. S. H. P. A. T. M. P. L. Knorr; M. O. S. S. J. A. de Vos; C. D. C. Escadafal | Krüger, et. al. Dec 2022 | A Multi-Center Clinical Diagnostic Accuracy Study of Surestatus - an Affordable, WHO Emergency-Use-Listed, Rapid, Point-of-Care, Antigen-Detecting Diagnostic Test for SARS-CoV-2 (preprint) | Preprint 2022 | SureStatus |  | No other info provided | Overall |  |  |  |  | 82·4 | 76·6,87·1 | 98·5 | 97·4,99·1
Terry | Sherli | 159 | Krüger L. | Krüger, et. al. May 2021 | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | PanBio (Abbott) |  |  | Asymptomatic | 12 | 0 | 374 | 2 | 85.7 | 60.1,96.0 | 100 | 99.0,100 | 100 | 99.5 |  | 85.7143 | 100 | 100 | 99.4681
Terry | Sherli | 159 | Krüger L. | Krüger, et. al. May 2021 | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | PanBio (Abbott) |  |  | Ct<25 |  |  |  |  | 95.8 | 88.5,98.6 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 159 | Krüger L. | Krüger, et. al. May 2021 | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | PanBio (Abbott) |  |  | Ct<30 |  |  |  |  | 93.5 | 86.6,97.0 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 159 | Krüger L. | Krüger, et. al. May 2021 | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | PanBio (Abbott) |  |  | Ct>25 |  |  |  |  | 66.7 | 49.6,80.2 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 159 | Krüger L. | Krüger, et. al. May 2021 | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | PanBio (Abbott) |  |  | Ct>30 |  |  |  |  | 33 | 13.8,60.1 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 159 | Krüger L. | Krüger, et. al. May 2021 | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | PanBio (Abbott) |  |  | Symptomatic | 79 | 1 | 620 | 12 | 86.8 | 78.2,92.3 | 99.8 | 99.1,100 | 98.8 | 98.1 |  | 86.8132 | 99.839 | 98.75 | 98.1013
Terry | Sherli | 159 | Krüger L. | Krüger, et. al. May 2021 | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | PanBio (Abbott) |  |  | Symptomatic <7 days | 69 | 1 | 533 | 7 | 90.8 | 82.2,95.5 | 99.6 | 98.9,100 | 98.6 | 98.7 |  | 90.7895 | 99.8127 | 98.5714 | 98.7037
Terry | Sherli | 159 | Krüger L. | Krüger, et. al. May 2021 | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | PanBio (Abbott) |  |  | Symptomatic >7 days | 8 | 0 | 57 | 5 | 61.5 | 35.5,82.3 | 100 | 93.7,100 | 100 | 91.1 |  | 61.5385 | 100 | 100 | 91.9355
Terry | Sherli | 159 | Krüger L. | Krüger, et. al. May 2021 | The Abbott PanBio WHO emergency use listed, rapid, antigen-detecting point-of-care diagnostic test for SARS-CoV-2-Evaluation of the accuracy and ease-of-use | 2021-05-01 | PanBio (Abbott) |  |  | Overall | 92 | 1 | 1001 | 14 | 86.8 | 79.0,92.0 | 99.9 | 99.4,100 | 98.9 | 98.6 |  | 86.7925 | 99.9002 | 98.9247 | 98.6207
Karen | Sherli | 119 | Y. Kurihara; Y. Kiyasu; Y. Akashi; Y. Takeuchi; K. Narahara; S. Mori; T. Takeshige; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kurihara, et. al. Jul 2021 | The evaluation of a novel digital immunochromatographic assay with silver amplification to detect SARS-CoV-2 | 2021-07-01 | Custom/Novel/In-house |  | QuickChaser® Auto SARS-CoV-2 and QuickChaser Immuno Reader II. Limits of Detection: QuickChaser® Auto SARS-CoV-2, 34–120 copies/test; Espline SARS®-CoV-2, 481–549 copies/test; QuickNaviTM- COVID19 Ag, 4394 copies/test; PanbioTM COVID-19 Ag Rapid Test De- vice, 1098–1924 copies/test; and SARS-CoV-2 Rapid Antigen Test, 549–1924 copies/test. Sn and Sp data only provided for QuickChaser test. | Overall | 62 | 2 | 1316 | 21 | 74.7 | 64.0,83.6 | 99.8 | 99.5,100 | 96.9 | 98.4 | PPV 95%CI 89.2,99.6; NPV 95% CI 97.6,99
Karen | Sherli | 119 | Y. Kurihara; Y. Kiyasu; Y. Akashi; Y. Takeuchi; K. Narahara; S. Mori; T. Takeshige; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kurihara, et. al. Jul 2021 | The evaluation of a novel digital immunochromatographic assay with silver amplification to detect SARS-CoV-2 | 2021-07-01 | Custom/Novel/In-house |  | QuickChaser® Auto SARS-CoV-2 and QuickChaser Immuno Reader II. | Symptomatic | 32 | 0 | 825 | 4 | 88.9 | 73.9,96.9 | 100 | 99.3,100 | 100 | 99.5 | PPV 95%CI 84.2,100; NPV 95% CI 98.8,99.9
Karen | Sherli | 119 | Y. Kurihara; Y. Kiyasu; Y. Akashi; Y. Takeuchi; K. Narahara; S. Mori; T. Takeshige; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kurihara, et. al. Jul 2021 | The evaluation of a novel digital immunochromatographic assay with silver amplification to detect SARS-CoV-2 | 2021-07-02 | Custom/Novel/In-house |  | QuickChaser® Auto SARS-CoV-2 and QuickChaser Immuno Reader II. | Asymptomatic | 30 | 2 | 491 | 17 | 63.8 | 48.5,77.3 | 99.6 | 98.5,100 | 93.8 | 96.7 | PPV 95%CI 79.2,99.2; NPV 95% CI 94.7,98
Karen | Sherli | 119 | Y. Kurihara; Y. Kiyasu; Y. Akashi; Y. Takeuchi; K. Narahara; S. Mori; T. Takeshige; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kurihara, et. al. Jul 2021 | The evaluation of a novel digital immunochromatographic assay with silver amplification to detect SARS-CoV-2 | 2021-07-03 | PanBio (Abbott) |  |  | Overall
Karen | Sherli | 119 | Y. Kurihara; Y. Kiyasu; Y. Akashi; Y. Takeuchi; K. Narahara; S. Mori; T. Takeshige; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kurihara, et. al. Jul 2021 | The evaluation of a novel digital immunochromatographic assay with silver amplification to detect SARS-CoV-2 | 2021-07-04 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall
Karen | Sherli | 119 | Y. Kurihara; Y. Kiyasu; Y. Akashi; Y. Takeuchi; K. Narahara; S. Mori; T. Takeshige; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kurihara, et. al. Jul 2021 | The evaluation of a novel digital immunochromatographic assay with silver amplification to detect SARS-CoV-2 | 2021-07-05 | Espline SARS-CoV-2 (Fujirebio) |  |  | Overall
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | E gene: Ct <20 |  |  |  |  | 91.1 | 78.7,97.0
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | E gene: Ct 20-24 |  |  |  |  | 78.3 | 64.2,87.9
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-03 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | E gene: Ct 25-29 |  |  |  |  | 39.1 | 26.4,53.6
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-04 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | E gene: Ct 30-34 |  |  |  |  | 10 | 4.9,18.8
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-05 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | E gene: Ct 35-38 |  |  |  |  | 5 | 0,25.4
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-06 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | E gene: Overall |  |  |  |  | 43.9 | 37.7,50.3
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-07 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | RdRp gene: Ct <20 |  |  |  |  | 92.7 | 79.9,09.2
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-08 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | RdRp gene: Ct 20-24 |  |  |  |  | 82.6 | 69.0,91.2
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-09 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | RdRp gene: Ct 25-29 |  |  |  |  | 37.8 | 25.1,52.4
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-10 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | RdRp gene: Ct 30-34 |  |  |  |  | 12.8 | 6.9,22.2
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kweon, et. al. May 2022 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-11 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | RdRp gene: Ct 35-38 |  |  |  |  | 3.7 | 0,19.8
Karen | Sherli | 119 | Y. Kurihara; Y. Kiyasu; Y. Akashi; Y. Takeuchi; K. Narahara; S. Mori; T. Takeshige; S. Notake; A. Ueda; K. Nakamura; H. Ishikawa; H. Suzuki | Kweon, et. al. May 2022 | The evaluation of a novel digital immunochromatographic assay with silver amplification to detect SARS-CoV-2 | 2021-07-06 | QuickNavi-COVID19 Ag |  |  | Overall
Terry | Sherli | 120 | O. J. Kweon; J. H. Lee; Y. S. Choi; B. S. Kim; Y. K. Lim; M. K. Lee; J. H. Park; J. Y. Park; S. H. Kim | Kyritsi, et. al. Aug 2021 | Positivity of Rapid Antigen Testing for SARS-CoV-2 With Serial Followed-up Nasopharyngeal Swabs in Hospitalized Patients due to COVID-19 | 2022-05-12 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | RdRp gene: Overall |  |  |  |  | 43.9 | 37.7,50.3
Terry | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Kyritsi, et. al. Aug 2021 | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-01 | Rapid Test Ag 2019-nCov (PROGNOSIS, BIOTECH, Greece) |  |  | Overall | 141 | 1 | 458 | 24 | 85.5 | 79.1,90.5 | 99.8 | 98.8,100
Terry | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Kyritsi, et. al. Aug 2021 | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-02 | Rapid Test Ag 2019-nCov (PROGNOSIS, BIOTECH, Greece) |  |  | Ct <30 | 141 |  |  | 15 | 90.4 | 84.6,94.5
Terry | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Kyritsi, et. al. Aug 2021 | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-03 | Rapid Test Ag 2019-nCov (PROGNOSIS, BIOTECH, Greece) |  |  | Overall: 1 part/thousand prevalence | 141 | 1 | 458 | 24 | 85.5 | 79.1,90.5 | 99.8 | 98.8,100 | 28.2% | 99.9% | The paper specified the PPV, NPV bsed on the prevalence formulas
Terry | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Kyritsi, et. al. Aug 2021 | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-04 | Rapid Test Ag 2019-nCov (PROGNOSIS, BIOTECH, Greece) |  |  | Overall: 1 percent prevalence | 141 | 1 | 458 | 24 | 85.5 | 79.1,90.5 | 99.8 | 98.8,100 | 79.9% | 99.9% | The paper specified the PPV, NPV bsed on the prevalence formulas
Terry | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Kyritsi, et. al. Aug 2021 | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-05 | Rapid Test Ag 2019-nCov (PROGNOSIS, BIOTECH, Greece) |  |  | Overall: 5 percent prevalence | 141 | 1 | 458 | 24 | 85.5 | 79.1,90.5 | 99.8 | 98.8,100 | 95.4% | 99.2% | The paper specified the PPV, NPV bsed on the prevalence formulas
Terry | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Kyritsi, et. al. Aug 2021 | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-06 | Rapid Test Ag 2019-nCov (PROGNOSIS, BIOTECH, Greece) |  |  | Ct <30: 1 part/thousand prevalence | 141 |  |  | 15 | 90.4 | 84.6,94.5 |  |  | 28.6% | 99.9% | The paper specified the PPV, NPV bsed on the prevalence formulas
Terry | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Kyritsi, et. al. Aug 2021 | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-07 | Rapid Test Ag 2019-nCov (PROGNOSIS, BIOTECH, Greece) |  |  | Ct <30: 1 percent prevalence | 141 |  |  | 15 | 90.4 | 84.6,94.5 |  |  | 80.7% | 99.9% | The paper specified the PPV, NPV bsed on the prevalence formulas
Terry | Sherli | 124 | M. Kyritsi; A. Vontas; I. Voulgaridi; A. Matziri; A. Komnos; D. Babalis; A. Papadogoulas; A. Oikonomou; V. A. Mouchtouri; M. Speletas; C. Hadjichristodoulou | Kyritsi, et. al. Aug 2021 | Rapid Test Ag 2019-nCoV (PROGNOSIS, BIOTECH, Larissa, Greece); Performance Evaluation in Hospital Setting with Real Time RT-PCR | 2021-08-08 | Rapid Test Ag 2019-nCov (PROGNOSIS, BIOTECH, Greece) |  |  | Ct <30: 5 percent prevalence | 141 |  |  | 15 | 90.4 | 84.6,94.5 |  |  | 95.6% | 99.5% | The paper specified the PPV, NPV bsed on the prevalence formulas
Terry | Sherli | 125 | L. Landaverde; J. Turcinovic; L. Doucette-Stamm; K. Gonzales; J. Platt; J. H. Connor; C. Klapperich | Landaverde, et. al. Mar 2022 | Comparison of BinaxNOW TM and SARS-CoV-2 qRT-PCR detection of the Omicron Variant from Matched Anterior Nares Swabs (preprint) | 2022-03-01 | BinaxNOW (Abbott) |  | Data from Table 1 Combined table | Overall | 55 | 0 | 217 | 47 | 53.9 |  | 100 |  | 100 | 82.2
Terry | Sherli | 126 | E. Layer; S. Hoehl; M. Widera; D. Bojkova; T. Westphal; R. Gottschalk; B. Boeddinghaus; J. Schork; S. Ciesek; U. Goetsch | Layer, et. al. Feb 2022 | SARS-CoV-2 screening strategies for returning international travellers: Evaluation of a rapid antigen test approach | 2022-02-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall | 26 | 0 | 1444 | 18 | 59 |  | 100
Karen/Amy | Sherli | 128 | J. LeGoff; S. Kernéis; C. Elie; S. Mercier-Delarue; N. Gastli; L. Choupeaux; J. Fourgeaud; M. L. Alby; P. Quentin; J. Pavie; P. Brazille; M. L. Néré; M. Minier; A. Gabassi; C. Leroy; B. Parfait; J. M. Tréluyer; C. Delaugerre | LeGoff, et. al. Oct 2021 | Evaluation of a saliva molecular point of care for the detection of SARS-CoV-2 in ambulatory care | 2021-10-01 | EasyCov (SkillCell-Alcen, France) |  | 1718 samples | Overall |  |  |  |  | 34 | 26,44 | 97 | 96,98
Karen/Amy | Sherli | 128 | J. LeGoff; S. Kernéis; C. Elie; S. Mercier-Delarue; N. Gastli; L. Choupeaux; J. Fourgeaud; M. L. Alby; P. Quentin; J. Pavie; P. Brazille; M. L. Néré; M. Minier; A. Gabassi; C. Leroy; B. Parfait; J. M. Tréluyer; C. Delaugerre | LeGoff, et. al. Oct 2021 | Evaluation of a saliva molecular point of care for the detection of SARS-CoV-2 in ambulatory care | 2021-10-02 | EasyCov (SkillCell-Alcen, France) |  | 691 samples | Symptomatic |  |  |  |  | 32 | 22,44 | 96 | 95,98
Karen/Amy | Sherli | 128 | J. LeGoff; S. Kernéis; C. Elie; S. Mercier-Delarue; N. Gastli; L. Choupeaux; J. Fourgeaud; M. L. Alby; P. Quentin; J. Pavie; P. Brazille; M. L. Néré; M. Minier; A. Gabassi; C. Leroy; B. Parfait; J. M. Tréluyer; C. Delaugerre | LeGoff, et. al. Oct 2021 | Evaluation of a saliva molecular point of care for the detection of SARS-CoV-2 in ambulatory care | 2021-10-03 | EasyCov (SkillCell-Alcen, France) |  | 1027 samples | Asymptomatic |  |  |  |  | 38 | 23,55 | 97 | 96,98
Karen/Amy | Sherli | 128 | J. LeGoff; S. Kernéis; C. Elie; S. Mercier-Delarue; N. Gastli; L. Choupeaux; J. Fourgeaud; M. L. Alby; P. Quentin; J. Pavie; P. Brazille; M. L. Néré; M. Minier; A. Gabassi; C. Leroy; B. Parfait; J. M. Tréluyer; C. Delaugerre | LeGoff, et. al. Oct 2021 | Evaluation of a saliva molecular point of care for the detection of SARS-CoV-2 in ambulatory care | 2021-10-04 | Standard Q COVID-19 Ag (SD Biosensor) |  | 1474 samples | Overall |  |  |  |  | 85 | 77,91 | 99 | 98,99
Karen/Amy | Sherli | 128 | J. LeGoff; S. Kernéis; C. Elie; S. Mercier-Delarue; N. Gastli; L. Choupeaux; J. Fourgeaud; M. L. Alby; P. Quentin; J. Pavie; P. Brazille; M. L. Néré; M. Minier; A. Gabassi; C. Leroy; B. Parfait; J. M. Tréluyer; C. Delaugerre | LeGoff, et. al. Oct 2021 | Evaluation of a saliva molecular point of care for the detection of SARS-CoV-2 in ambulatory care | 2021-10-05 | Standard Q COVID-19 Ag (SD Biosensor) |  | 652 samples | Symptomatic |  |  |  |  | 90 | 81,96 | 98 | 96,99
Karen/Amy | Sherli | 128 | J. LeGoff; S. Kernéis; C. Elie; S. Mercier-Delarue; N. Gastli; L. Choupeaux; J. Fourgeaud; M. L. Alby; P. Quentin; J. Pavie; P. Brazille; M. L. Néré; M. Minier; A. Gabassi; C. Leroy; B. Parfait; J. M. Tréluyer; C. Delaugerre | LeGoff, et. al. Oct 2021 | Evaluation of a saliva molecular point of care for the detection of SARS-CoV-2 in ambulatory care | 2021-10-06 | Standard Q COVID-19 Ag (SD Biosensor) |  | 822 samples | Asymptomatic |  |  |  |  | 74 | 57,88 | 99 | 98.0,100
Ellyn | Sherli | 175 | Leixner G. | Leixner, et. al. Jul 2021 | Evaluation of the AMP SARS-CoV-2 rapid antigen test in a hospital setting | 2021-07-01 | AMP Rapid Test SARS-CoV-2 Ag (AMP Diagnostics) |  |  | Overall | 65 | 1 | 297 | 29 | 69.15 | 58.8,78.3 | 99.66 | 98.1,100 | 98.48 | 91.1 |  | 69.1489 | 99.6644 | 98.4848 | 91.1043
Ellyn | Sherli | 175 | Leixner G. | Leixner, et. al. Jul 2021 | Evaluation of the AMP SARS-CoV-2 rapid antigen test in a hospital setting | 2021-07-01 | AMP Rapid Test SARS-CoV-2 Ag (AMP Diagnostics) |  |  | Ct<25 |  |  |  |  | 100 | 89.4,100 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 175 | Leixner G. | Leixner, et. al. Jul 2021 | Evaluation of the AMP SARS-CoV-2 rapid antigen test in a hospital setting | 2021-07-01 | AMP Rapid Test SARS-CoV-2 Ag (AMP Diagnostics) |  |  | Ct<30 |  |  |  |  | 91.8 | 81.9,97.3 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 181 | Linares M. | Linares, et. al. Oct 2020 | Panbio antigen rapid test is reliable to diagnose SARS-CoV-2 infection in the first 7 days after the onset of symptoms | 2020-10-01 | PanBio (Abbott) |  |  | Overall |  |  |  |  | 73.3 | 62.2,83.8 | 100 |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 181 | Linares M. | Linares, et. al. Oct 2020 | Panbio antigen rapid test is reliable to diagnose SARS-CoV-2 infection in the first 7 days after the onset of symptoms | 2020-10-01 | PanBio (Abbott) |  |  | Symptomatic <5 days | 29 | 0 | 102 | 5 | 85.3 | 73.4,97.2 | 100 |  | 100 | 95.3 |  | 85.2941 | 100 | 100 | 95.3271
Ellyn | Sherli | 181 | Linares M. | Linares, et. al. Oct 2020 | Panbio antigen rapid test is reliable to diagnose SARS-CoV-2 infection in the first 7 days after the onset of symptoms | 2020-10-01 | PanBio (Abbott) |  |  | Symptomatic <7 days | 32 | 0 | 104 | 5 | 86.5 | 75.5,97.5 | 100 |  | 100 | 95.4 |  | 86.4865 | 100 | 100 | 95.4128
Ellyn | Sherli | 181 | Linares M. | Linares, et. al. Oct 2020 | Panbio antigen rapid test is reliable to diagnose SARS-CoV-2 infection in the first 7 days after the onset of symptoms | 2020-10-01 | PanBio (Abbott) |  |  | Symptomatic ≥7 days | 7 | 0 | 29 | 6 | 53.8 | 26.7,80.9 | 100 |  | 100 | 82.8 |  | 53.8462 | 100 | 100 | 82.8571
Ellyn | Sherli | 183 | Lindner A. | Lindner, et. al. Apr 2021 | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with professional-collected nasal versus nasopharyngeal swab | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | [11 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 183 | Lindner A. | Lindner, et. al. Apr 2021 | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with professional-collected nasal versus nasopharyngeal swab | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | NMT |  |  |  |  | 80.5 | 66.0,89.8 | 98.6 | 94.9,99.6 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 183 | Lindner A. | Lindner, et. al. Apr 2021 | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with professional-collected nasal versus nasopharyngeal swab | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | NP |  |  |  |  | 73.2 | 58.1,84.3 | 99.3 | 96.0,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 182 | Lindner A. | Lindner, et. al. Apr 2021 | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with self-collected nasal swab versus professional-collected nasopharyngeal swab | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | NMT |  |  |  |  | 74.4 | 58.9,85.4 | 99.2 | 97.1,99.8 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 182 | Lindner A. | Lindner, et. al. Apr 2021 | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with self-collected nasal swab versus professional-collected nasopharyngeal swab | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | NMT High viral load |  |  |  |  | 95.7 | 79.0,99.8 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 182 | Lindner A. | Lindner, et. al. Apr 2021 | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with self-collected nasal swab versus professional-collected nasopharyngeal swab | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | NP |  |  |  |  | 79.5 | 64.5,89.2 | 99.6 | 97.8,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 182 | Lindner A. | Lindner, et. al. Apr 2021 | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with self-collected nasal swab versus professional-collected nasopharyngeal swab | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | NP High viral load |  |  |  |  | 100 | 85.7,100 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 182 | Lindner A. | Lindner, et. al. Apr 2021 | Head-to-head comparison of SARS-CoV-2 antigen-detecting rapid test with self-collected nasal swab versus professional-collected nasopharyngeal swab | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | [11 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 184 | Lindner A. | Lindner, et. al. May 2021 | Diagnostic accuracy and feasibility of patient self-testing with a SARS-CoV-2 antigen-detecting rapid test | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | High viral load |  |  |  |  | 96.6 | 82.8,99.8 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Ellyn | Sherli | 184 | Lindner A. | Lindner, et. al. May 2021 | Diagnostic accuracy and feasibility of patient self-testing with a SARS-CoV-2 antigen-detecting rapid test | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Professional | 34 | 1 | 105 | 6 | 85 | 70.9,92.9 | 99.1 | 94.8,99.5 | 97.1 | 94.6 |  | 85 | 99.0566 | 97.1429 | 94.5946
Ellyn | Sherli | 184 | Lindner A. | Lindner, et. al. May 2021 | Diagnostic accuracy and feasibility of patient self-testing with a SARS-CoV-2 antigen-detecting rapid test | 2021-05-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Self testing | 33 | 0 | 104 | 7 | 82.5 | 68.1,91.3 | 100 | 96.5,100 | 100 | 93.7 |  | 82.5 | 100 | 100 | 93.6937
Amy | Sherli | 141 | D. K. Mandal; B. R. Bhattarai; S. Pokhrel; M. Chhusyabaga; P. Bhandari; M. P. Bhatt; S. B. Marhattha | Mandal, et. al. May 2022 | Diagnostic Performance of SARS-CoV-2 Rapid Antigen Test in relation to RT-PCR CqValue | 2022-05-01 | Espline SARS-CoV-2 (Fujirebio) |  |  | Overall | 75 | 2 | 93 | 43 | 63.6% | 54.7,71.9 | 97.9% | 93.6,99.6 | 97.4% | 68.4% | PPV CI 92.2–99.6% NPV CI 60.3–75.8%
Amy | Sherli | 141 | D. K. Mandal; B. R. Bhattarai; S. Pokhrel; M. Chhusyabaga; P. Bhandari; M. P. Bhatt; S. B. Marhattha | Mandal, et. al. May 2022 | Diagnostic Performance of SARS-CoV-2 Rapid Antigen Test in relation to RT-PCR CqValue | 2022-05-02 | Espline SARS-CoV-2 (Fujirebio) |  |  | CT <20 | 44 |  |  | 6
Amy | Sherli | 141 | D. K. Mandal; B. R. Bhattarai; S. Pokhrel; M. Chhusyabaga; P. Bhandari; M. P. Bhatt; S. B. Marhattha | Mandal, et. al. May 2022 | Diagnostic Performance of SARS-CoV-2 Rapid Antigen Test in relation to RT-PCR CqValue | 2022-05-03 | Espline SARS-CoV-2 (Fujirebio) |  |  | CT 20-<25 | 57 |  |  | 10
Amy | Sherli | 141 | D. K. Mandal; B. R. Bhattarai; S. Pokhrel; M. Chhusyabaga; P. Bhandari; M. P. Bhatt; S. B. Marhattha | Mandal, et. al. May 2022 | Diagnostic Performance of SARS-CoV-2 Rapid Antigen Test in relation to RT-PCR CqValue | 2022-05-04 | Espline SARS-CoV-2 (Fujirebio) |  |  | CT 25-<30 | 36 |  |  | 19
Amy | Sherli | 141 | D. K. Mandal; B. R. Bhattarai; S. Pokhrel; M. Chhusyabaga; P. Bhandari; M. P. Bhatt; S. B. Marhattha | Mandal, et. al. May 2022 | Diagnostic Performance of SARS-CoV-2 Rapid Antigen Test in relation to RT-PCR CqValue | 2022-05-05 | Espline SARS-CoV-2 (Fujirebio) |  |  | CT 30-<37 | 13 |  |  | 51
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-01 | PathoCatch (Accucare) |  |  | Symptomatic NP swabs | 81 |  |  | 40 | 87.3 | 77.9,93.8 | 100 | 96.1,100
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-02 | PathoCatch (Accucare) |  |  | Asymptomatic NP Swabs |  |  |  |  | 28.6 | 15.7,44.6 | 100 | 98.8,100.0
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-03 | PathoCatch (Accucare) |  |  | Overall oral swabs | 62 |  |  | 0
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-04 | PathoCatch (Accucare) |  |  | HCW collected Symptomatic Oral Swabs | 30 | 2 | 60 | 17 | 63.8 | 48.5,77.3 | 96.8 | 88.8,99.6 | 93.8 | 77.9 | PPV CI95 [79.1-98.4] NPV CI95 [70.7-83.8]
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-05 | PathoCatch (Accucare) |  |  | Self collected Symptomatic Oral Swabs | 20 | 1 | 30 | 12 | 62.5 | 43.7,78.9 | 96.7 | 82.8,99.99 | 90.9 | 70.7 | PPV CI95 [71.9-97.5] NPV CI95 [60.5-79.2]
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-06 | PathoCatch (Accucare) |  |  | HCW collected Asymptomatic Oral Swabs |  |  |  |  | 23.1 | 8.9,43.6 | 96.8 | 93.3,98.8 | 50 | 89.9 | PPV CI95 [25.8-74.2] NPV CI95 [87.8-91.7]
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-07 | PathoCatch (Accucare) |  |  | Self collected Asymptomatic Oral Swabs |  |  |  |  | 37.5 | 15.2,64.6 | 98.5 | 94.6,99.8 | 75 | 92.8 | PPV CI95 [39.8-93.2] NPV CI95 [89.82-94.9]
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-08 | PathoCatch (Accucare) |  |  | Oral Swabs CT =<25 | 57 |  |  | 23
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-09 | PathoCatch (Accucare) |  |  | Oral Swabs CT >25 | 5 |  |  | 36
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-10 | PathoCatch (Accucare) |  |  | Oral Swabs CT =<30 | 59 |  |  | 29
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-11 | PathoCatch (Accucare) |  |  | Oral Swabs CT >30 | 3 |  |  | 30
Amy | Sherli | 300 | A. Mane; S. Jain; A. Jain; M. Pereira; A. Sirsat; G. Pathak; V. Bhoi; S. Bhavsar; S. Panda | Mane, et. al. May 2022 | Diagnostic performance of oral swab specimen for SARS-CoV-2 detection with rapid point-of-care lateral flow antigen test | 2022-05-12 | PathoCatch (Accucare) |  |  | Overall | 62 |  |  | 59
Amy | Sherli | 301 | M. Maniscalco; P. Ambrosino; A. Ciullo; S. Fuschillo; V. Valente; C. Gaudiosi; D. Paris; R. Cobuccio; F. Stefanelli; A. Motta | Maniscalco, et. al. Aug 2021 | A Rapid Antigen Detection Test to Diagnose SARS-CoV-2 Infection Using Exhaled Breath Condensate by A Modified Inflammacheck(®) Device | 2021-08-01 | Inflammacheck (Exhalation Technology LTD) |  |  | Overall | 12 | 1 | 91 | 1 | 92.3 | 64.0,99.8 | 98.9 | 94.1,100.0 | 92.3 | 98.8 | PPV CI95 [62.9-98.8] NPV CI95 [93.3-99.8]
Ellyn | Sherli | 194 | Masiá M. | Masiá, et. al. Jan 2021 | Nasopharyngeal Panbio COVID-19 Antigen Performed at Point-of-Care Has a High Sensitivity in Symptomatic and Asymptomatic Patients with Higher Risk for Transmission and Older Age | 2021-01-01 | PanBio (Abbott) |  |  | Overall | 92 | 0 | 457 | 43 | 68.1 |  | 100 |  | 100 | 91 |  | 68.1481 | 100 | 100 | 91.4
Ellyn | Sherli | 194 | Masiá M. | Masiá, et. al. Jan 2021 | Nasopharyngeal Panbio COVID-19 Antigen Performed at Point-of-Care Has a High Sensitivity in Symptomatic and Asymptomatic Patients with Higher Risk for Transmission and Older Age | 2021-01-01 | PanBio (Abbott) |  |  | Asymptomatic | 13 | 0 | 253 | 26 | 33.3 | 19.6,50.3 | 100 |  | 100 | 90.7 |  | 33.3333 | 100 | 100 | 90.681
Ellyn | Sherli | 194 | Masiá M. | Masiá, et. al. Jan 2021 | Nasopharyngeal Panbio COVID-19 Antigen Performed at Point-of-Care Has a High Sensitivity in Symptomatic and Asymptomatic Patients with Higher Risk for Transmission and Older Age | 2021-01-01 | PanBio (Abbott) |  |  | Symptomatic | 105 | 0 | 456 | 51 | 67.3 | 59.3,74.5 | 100 |  | 100 | 89.9 |  | 67.3077 | 100 | 100 | 89.9408
Terry | Sherli | 155 | A. Mizrahi; J. C. Nguyen Van; N. El Helali; J. Lourtet-Hascoet; I. Jabnoune; M. L. Pacreau; Y. Talb; J. Fourgeaud; M. Leruez-Ville; B. Pilmis; V. Avettand-Fenoel; A. Le Monnier | Mizrahi, et. al. Nov 2021 | The Coris BioConcept COVID 19 Ag Respi-Strip, a field experience feedback | 2021-11-01 | Respi-Strip (Coris BioConcept) |  |  | Initial evaluation (15-min reading) (n = 52) | 8 | 0 | 41 | 3 | 72 |  | 100 |  | 100 | 93
Terry | Sherli | 155 | A. Mizrahi; J. C. Nguyen Van; N. El Helali; J. Lourtet-Hascoet; I. Jabnoune; M. L. Pacreau; Y. Talb; J. Fourgeaud; M. Leruez-Ville; B. Pilmis; V. Avettand-Fenoel; A. Le Monnier | Mizrahi, et. al. Nov 2021 | The Coris BioConcept COVID 19 Ag Respi-Strip, a field experience feedback | 2021-11-02 | Respi-Strip (Coris BioConcept) |  |  | Coris-Ag: 15-min reading (n = 294) | 16 | 0 | 252 | 26 | 38.1 |  | 100 |  | 100 | 90.6
Terry | Sherli | 155 | A. Mizrahi; J. C. Nguyen Van; N. El Helali; J. Lourtet-Hascoet; I. Jabnoune; M. L. Pacreau; Y. Talb; J. Fourgeaud; M. Leruez-Ville; B. Pilmis; V. Avettand-Fenoel; A. Le Monnier | Mizrahi, et. al. Nov 2021 | The Coris BioConcept COVID 19 Ag Respi-Strip, a field experience feedback | 2021-11-03 | Respi-Strip (Coris BioConcept) |  |  | Coris-Ag: 30-min reading (n = 294) | 19 | 0 | 252 | 23 | 45.2 |  | 100 |  | 100 | 91.6
Terry | Sherli | 155 | A. Mizrahi; J. C. Nguyen Van; N. El Helali; J. Lourtet-Hascoet; I. Jabnoune; M. L. Pacreau; Y. Talb; J. Fourgeaud; M. Leruez-Ville; B. Pilmis; V. Avettand-Fenoel; A. Le Monnier | Mizrahi, et. al. Nov 2021 | The Coris BioConcept COVID 19 Ag Respi-Strip, a field experience feedback | 2021-11-04 | Respi-Strip (Coris BioConcept) |  | Period 1 was based on 30 minute readings | Period 1 (n = 158) | 16 | 0 | 131 | 11 | 59.3 |  | 100 |  | 100 | 92.2
Terry | Sherli | 155 | A. Mizrahi; J. C. Nguyen Van; N. El Helali; J. Lourtet-Hascoet; I. Jabnoune; M. L. Pacreau; Y. Talb; J. Fourgeaud; M. Leruez-Ville; B. Pilmis; V. Avettand-Fenoel; A. Le Monnier | Mizrahi, et. al. Nov 2021 | The Coris BioConcept COVID 19 Ag Respi-Strip, a field experience feedback | 2021-11-05 | Respi-Strip (Coris BioConcept) |  | Period 2 was based on 30 min readings | Period 2 (n = 136) | 3 | 0 | 121 | 12 | 20 |  | 100 |  | 100 | 91
Terry | Sherli | 156 | I. J. B. Møller; A. R. Utke; U. K. Rysgaard; L. J. Østergaard; S. Jespersen | Møller, et. al. Jan 2022 | Diagnostic performance, user acceptability, and safety of unsupervised SARS-CoV-2 rapid antigen-detecting tests performed at home | 2022-01-01 | SARS-CoV-2 Antigen Rapid Test (Hangzhou Immuno Biotech Co Ltd, China). |  |  | Overall | 41 | 0 | 335 | 25 | 62.1 | 50.1,72.9 | 100 | 98.9,100
Terry | Sherli | 156 | I. J. B. Møller; A. R. Utke; U. K. Rysgaard; L. J. Østergaard; S. Jespersen | Møller, et. al. Jan 2022 | Diagnostic performance, user acceptability, and safety of unsupervised SARS-CoV-2 rapid antigen-detecting tests performed at home | 2022-01-02 | COVID-19 Antigen Detection Kit (DNA Diagnostic A/S, Denmark) |  |  | Overall | 23 | 0 | 388 | 12 | 65.7 | 49.2,79.2 | 100 | 99100
Terry | Sherli | 156 | I. J. B. Møller; A. R. Utke; U. K. Rysgaard; L. J. Østergaard; S. Jespersen | Møller, et. al. Jan 2022 | Diagnostic performance, user acceptability, and safety of unsupervised SARS-CoV-2 rapid antigen-detecting tests performed at home | 2022-01-03 | PanBio (Abbott) |  |  | Overall | 0 | 1 | 387 | 0 | Not estimable |  | 100 | 95.6,100
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Prof.-sampling: All (N =36), Prof AN | 31 | 0 | 96 | 5 | 86.1 | 71.3,93.9 | 100 | 95.7,100
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Prof.-sampling: All (N =36), Prof NMT | 31 | 0 | 96 | 5 | 86.1 | 71.3,93.9 | 100 | 95.7,100
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-03 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Prof.-sampling: 7 log10 (N=29), Prof AN | 28 |  |  | 1 | 96.6 | 82.8,99.8
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-04 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Prof.-sampling: 7 log10 (N=29), Prof NMT | 28 |  |  | 1 | 96.6 | 82.8,99.9
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-05 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Prof.-sampling: <7 log10 (N=7), Prof AN | 3 |  |  | 4 | 42.9 | 15.8,75
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-06 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Prof.-sampling: <7 log10 (N=7), Prof NMT | 3 |  |  | 4 | 42.9 | 15.8,76
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-07 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Self-sampling: All (N=34), Self NMT | 31 | 1 | 61 | 3 | 91.2 | 77.0,97 | 98.4 | 91.4,99.9
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-08 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Self-sampling: All (N=34), Prof. NP | 31 | 0 | 62 | 3 | 91.2 | 77,97 | 100 | 94.2,100
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-09 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Self-sampling: 7 log10 (N=25), Self NMT | 25 |  |  | 0 | 100 | 86.7,100
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-10 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Self-sampling: 7 log10 (N=25), Prof. NP | 25 |  |  | 0 | 100 | 86.7,101
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2021 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-11 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Self-sampling: <7 log10 (N=9), Self NMT | 6 |  |  | 3 | 66.7 | 35.4,87.9
Terry | Sherli | 162 | O. Nikolai; C. Rohardt; F. Tobian; A. Junge; V. M. Corman; T. C. Jones; M. Gaeddert; F. Lainati; J. A. Sacks; J. Seybold; F. P. Mockenhaupt; C. M. Denkinger; A. K. Lindner | Nikolai, et. al. Aug 2022 | Anterior nasal versus nasal mid-turbinate sampling for a SARS-CoV-2 antigen-detecting rapid test: does localisation or professional collection matter? | 2021-08-12 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Self-sampling: <7 log10 (N=9), Prof. NP | 6 |  |  | 3 | 66.7 | 35.4,87.10
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-01 | PanBio (Abbott) |  |  | Overall | 83 | 1 | 337 | 74
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-02 | CoV2Ag assay (Siemens Healthineers, Eschborn, Germany) |  |  | Overall | 22 | 1 | 45 | 23
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-03 | GenBody COVAG025 (GenBody) |  |  | Overall | 23 | 6 | 43 | 15
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-04 | GENEDIA W COVID-19 Ag Test (Green Cross Medical Science Corp.) |  |  | Overall | 20 | 1 | 51 | 16
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-05 | Humasis COVID-19 Ag Test kit (Humasis Co., Ltd.) |  |  | Overall | 18 | 11 | 46 | 15
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-06 | Immupass VivaDiag (VivaChek Biotech) |  |  | Overall | 13 | 0 | 71 | 18
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-07 | Helix i-SARS-CoV-2 Ag Rapid Test (Cellex Biotech Co.) |  |  | Overall | 14 | 0 | 15 | 7
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-08 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall | 43 | 0 | 104 | 22
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-09 | Rapid COVID-19 Antigen Test (Healgen Scientific) |  |  | Overall | 18 | 0 | 50 | 12
Karen/Amy | Sherli | 165 | M. Nóra; D. Déri; D. S. Veres; Z. Kis; E. Barcsay; B. Pályi | Nóra, et. al. Feb 2022 | Evaluating the field performance of multiple SARS-Cov-2 antigen rapid tests using nasopharyngeal swab samples | 2022-02-10 | Severe Acute Respiratory Syndrome Coronavirus 2 (SARS-CoV-2) Antigen Detection Kit (Colloidal Gold-Based) Nanjing Vazyme Medical Technology Co |  |  | Overall | 34 | 0 | 63 | 11
Terry | Sherli | 167 | G. A. Okoye; H. I. Kamara; M. Strobeck; T. A. Mellman; J. Kwagyan; A. Sullivan; A. S. Byrd; B. Shokrani; H. E. Mighty | Okoye, et. al. Oct 2021 | Diagnostic accuracy of a rapid diagnostic test for the early detection of COVID-19 | 2021-10-01 | BinaxNOW (Abbott) |  |  | Overall | 45 | 2 | 3759 | 4 | 91.84 | 80.40,97.73 | 99.95 | 99.81,99.99 | 95.74 | 99.89
Terry | Sherli | 168 | S. N. Onsongo; K. Otieno; S. van Duijn; E. Adams; M. Omollo; I. A. Odero; A. K'Oloo; N. Houben; E. Milimo; R. Aroka; H. C. Barsosio; F. Oluoch; A. Odhiambo; S. Kariuki; T. F. R. de Wit | Onsongo, et. al. Nov 2021 | Performance of a rapid antigen test for SARS-CoV-2 in Kenya | 2021-11-01 | NowCheck COVID-19 Ag test (Bionote) |  |  | Ct </= 25 |  |  |  |  | 87.5 | 71.0,96.5 | 89.5 | 87.4,91.4 | 21.7 (18.1−25.8) | 99.5 (98.9−99.8)
Terry | Sherli | 168 | S. N. Onsongo; K. Otieno; S. van Duijn; E. Adams; M. Omollo; I. A. Odero; A. K'Oloo; N. Houben; E. Milimo; R. Aroka; H. C. Barsosio; F. Oluoch; A. Odhiambo; S. Kariuki; T. F. R. de Wit | Onsongo, et. al. Nov 2021 | Performance of a rapid antigen test for SARS-CoV-2 in Kenya | 2021-11-02 | NowCheck COVID-19 Ag test (Bionote) |  |  | Ct </= 30 |  |  |  |  | 87.7 | 77.2,94.5 | 92.3 | 90.4,93.9 | 44.2 (38.4−50.2) | 99.1 (98.3−99.5)
Terry | Sherli | 168 | S. N. Onsongo; K. Otieno; S. van Duijn; E. Adams; M. Omollo; I. A. Odero; A. K'Oloo; N. Houben; E. Milimo; R. Aroka; H. C. Barsosio; F. Oluoch; A. Odhiambo; S. Kariuki; T. F. R. de Wit | Onsongo, et. al. Nov 2021 | Performance of a rapid antigen test for SARS-CoV-2 in Kenya | 2021-11-03 | NowCheck COVID-19 Ag test (Bionote) |  |  | Ct </= 35 |  |  |  |  | 84.5 | 76.0 ,90.85 | 95.3 | 93.7 ,96.59 | 67.4 (60.4 −73.79) | 98.16 (97.1 − 98.8)
Terry | Sherli | 168 | S. N. Onsongo; K. Otieno; S. van Duijn; E. Adams; M. Omollo; I. A. Odero; A. K'Oloo; N. Houben; E. Milimo; R. Aroka; H. C. Barsosio; F. Oluoch; A. Odhiambo; S. Kariuki; T. F. R. de Wit | Onsongo, et. al. Nov 2021 | Performance of a rapid antigen test for SARS-CoV-2 in Kenya | 2021-11-04 | NowCheck COVID-19 Ag test (Bionote) |  |  | Ct </= 40 |  |  |  |  | 71.5 | 63.2,78.6 | 97.5 | 96.2,98.5 | 83.7 (76.9−88.1) | 95.1 (93.7−96.1)
Yati | Sherli | 225 | Osmanodja B. | Osmanodja, et. al. May 2021 | Accuracy of a novel sars-cov-2 antigen-detecting rapid diagnostic test from standardized self-collected anterior nasal swabs | 2021-05-01 | Custom/Novel/In-house |  | Drager Antigen Test | Overall |  |  |  |  | 88.6 | 78.7,94.9 | 99.7 | 98.2,100 | 98.4 | 97.5 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Yati | Sherli | 225 | Osmanodja B. | Osmanodja, et. al. May 2021 | Accuracy of a novel sars-cov-2 antigen-detecting rapid diagnostic test from standardized self-collected anterior nasal swabs | 2021-05-01 | Custom/Novel/In-house |  | Drager Antigen Test | Ct≤32 |  |  |  |  | 92.5 | 83.4,97.5 | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Terry | Sherli | 171 | K. C. Paap; A. M. van Loon; F. M. Koene; L. W. van Buul; S. Jurriaans; M. Smalbrugge; M. D. de Jong; C. Hertogh | Paap, et. al. Nov 2021 | Clinical evaluation of single-swab sampling for rapid COVID-19 detection in outbreak settings in Dutch nursing homes | 2021-11-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall | 27 | 45 | 363 | 26 | 50.9 |  | 89
Terry | Sherli | 171 | K. C. Paap; A. M. van Loon; F. M. Koene; L. W. van Buul; S. Jurriaans; M. Smalbrugge; M. D. de Jong; C. Hertogh | Paap, et. al. Nov 2021 | Clinical evaluation of single-swab sampling for rapid COVID-19 detection in outbreak settings in Dutch nursing homes | 2021-11-02 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Ct < 20 |  |  |  |  | 100
Terry | Sherli | 171 | K. C. Paap; A. M. van Loon; F. M. Koene; L. W. van Buul; S. Jurriaans; M. Smalbrugge; M. D. de Jong; C. Hertogh | Paap, et. al. Nov 2021 | Clinical evaluation of single-swab sampling for rapid COVID-19 detection in outbreak settings in Dutch nursing homes | 2021-11-03 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Ct 20-25 |  |  |  |  | 90
Terry | Sherli | 171 | K. C. Paap; A. M. van Loon; F. M. Koene; L. W. van Buul; S. Jurriaans; M. Smalbrugge; M. D. de Jong; C. Hertogh | Paap, et. al. Nov 2021 | Clinical evaluation of single-swab sampling for rapid COVID-19 detection in outbreak settings in Dutch nursing homes | 2021-11-04 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Ct 25-30 |  |  |  |  | 22.2
Terry | Sherli | 171 | K. C. Paap; A. M. van Loon; F. M. Koene; L. W. van Buul; S. Jurriaans; M. Smalbrugge; M. D. de Jong; C. Hertogh | Paap, et. al. Nov 2022 | Clinical evaluation of single-swab sampling for rapid COVID-19 detection in outbreak settings in Dutch nursing homes | 2021-11-05 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Ct > 30 |  |  |  |  | 14.3
Terry | Sherli | 172 | A. K. Pandey; A. Mohanty; V. Hada; R. S. Rath; S. Kumar; S. Kishore; R. Kant | Pandey, et. al. Aug 2021 | Comparison of the Rapid Antigen Testing Method With RT-qPCR for the Diagnosis of COVID-19 | 2021-08-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 30 | 7 | 255 | 26 | 53.6 | 39.7,67.0 | 97.3 | 94.6,98.9 | 81.1 (64.8-92.0) | 90.7 (86.7-93.9)
Terry | Sherli | 172 | A. K. Pandey; A. Mohanty; V. Hada; R. S. Rath; S. Kumar; S. Kishore; R. Kant | Pandey, et. al. Aug 2021 | Comparison of the Rapid Antigen Testing Method With RT-qPCR for the Diagnosis of COVID-19 | 2021-08-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptomatic | 25 | 5 | 85 | 16 | 61 | 44.5,75.8 | 94.4 | 87.5,98.2 | 83.3 (65.3-94.4) | 84.2 (75.6-90.7)
Terry | Sherli | 172 | A. K. Pandey; A. Mohanty; V. Hada; R. S. Rath; S. Kumar; S. Kishore; R. Kant | Pandey, et. al. Aug 2022 | Comparison of the Rapid Antigen Testing Method With RT-qPCR for the Diagnosis of COVID-19 | 2021-08-03 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic | 5 | 2 | 170 | 10 | 33.3 | 11.8,61.6 | 98.8 | 95.9,99.9 | 71.4(29.0-96.3) | 94.4 (90.0-97.3)
Amy | Sherli | 174 | J. Y. Park; J. H. Lee; B. K. Cha; B. S. Kim; H. J. Lee; G. H. Kim; K. T. Kang; Y. S. Lee; S. K. Ahn; S. H. Kim | Park, et. al. Feb 2022 | Analysis of the Efficacy of Universal Screening of Coronavirus Disease with Antigen-Detecting Rapid Diagnostic Tests at Point-or-Care Settings and Sharing the Experience of Admission Protocol—A Pilot Study | 2022-02-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 28 | 3 | 602 | 13 | 68.3 |  | 99.5 |  | 90.3 | 97.9
Amy | Sherli | 174 | J. Y. Park; J. H. Lee; B. K. Cha; B. S. Kim; H. J. Lee; G. H. Kim; K. T. Kang; Y. S. Lee; S. K. Ahn; S. H. Kim | Park, et. al. Feb 2022 | Analysis of the Efficacy of Universal Screening of Coronavirus Disease with Antigen-Detecting Rapid Diagnostic Tests at Point-or-Care Settings and Sharing the Experience of Admission Protocol—A Pilot Study | 2022-02-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptomatic |  |  |  |  | 73 |  | 98.8 |  | 87.1 | 96.9
Amy | Sherli | 177 | W. F. Peacock; K. M. Soto-Ruiz; S. L. House; C. M. Cannon; G. Headden; B. Tiffany; S. Motov; K. Merchant-Borna; A. M. Chang; C. Pearson; B. W. Patterson; A. E. Jones; J. Miller; J. Varon; A. Bastani; C. Clark; Z. Rafique; B. Kea; J. Eppensteiner; J. M. Williams; S. A. Mahler; B. E. Driver; P. Hendry; E. Quackenbush; D. Robinson; J. W. Schrock; J. P. D'Etienne; C. J. Hogan; A. Osborne; R. Riviello; S. Young | Peacock, et. al. Jan 2022 | Utility of COVID-19 antigen testing in the emergency department | 2022-01-01 | BinaxNOW (Abbott) |  |  | Overall |  |  |  |  | 76.9 | 69.9,82.9 | 98.6 | 97.2,99.4 | 94.3 | 93.3 | PPV CI95 [89.1-97.5] NPV CI95 [90.9-95.1]
Amy | Sherli | 177 | W. F. Peacock; K. M. Soto-Ruiz; S. L. House; C. M. Cannon; G. Headden; B. Tiffany; S. Motov; K. Merchant-Borna; A. M. Chang; C. Pearson; B. W. Patterson; A. E. Jones; J. Miller; J. Varon; A. Bastani; C. Clark; Z. Rafique; B. Kea; J. Eppensteiner; J. M. Williams; S. A. Mahler; B. E. Driver; P. Hendry; E. Quackenbush; D. Robinson; J. W. Schrock; J. P. D'Etienne; C. J. Hogan; A. Osborne; R. Riviello; S. Young | Peacock, et. al. Jan 2022 | Utility of COVID-19 antigen testing in the emergency department | 2022-01-02 | BinaxNOW (Abbott) |  |  | <8 days |  |  |  |  | 84.6 | 76.8,90.6 | 98.5 | 96.6,99.5 | 95.2 | 94.9 | PPV CI95 [89.1-98.4] NPV CI95 [92.1-97.0]
Amy | Sherli | 177 | W. F. Peacock; K. M. Soto-Ruiz; S. L. House; C. M. Cannon; G. Headden; B. Tiffany; S. Motov; K. Merchant-Borna; A. M. Chang; C. Pearson; B. W. Patterson; A. E. Jones; J. Miller; J. Varon; A. Bastani; C. Clark; Z. Rafique; B. Kea; J. Eppensteiner; J. M. Williams; S. A. Mahler; B. E. Driver; P. Hendry; E. Quackenbush; D. Robinson; J. W. Schrock; J. P. D'Etienne; C. J. Hogan; A. Osborne; R. Riviello; S. Young | Peacock, et. al. Jan 2022 | Utility of COVID-19 antigen testing in the emergency department | 2022-01-03 | BinaxNOW (Abbott) |  |  | 8-14 days |  |  |  |  | 61.4 | 45.5,75.6 | 98.4 | 91.2,100 | 96.4 | 77.9 | PPV CI95 [81.7-99.9] NPV CI95 [67.0-86.6]
Amy | Sherli | 177 | W. F. Peacock; K. M. Soto-Ruiz; S. L. House; C. M. Cannon; G. Headden; B. Tiffany; S. Motov; K. Merchant-Borna; A. M. Chang; C. Pearson; B. W. Patterson; A. E. Jones; J. Miller; J. Varon; A. Bastani; C. Clark; Z. Rafique; B. Kea; J. Eppensteiner; J. M. Williams; S. A. Mahler; B. E. Driver; P. Hendry; E. Quackenbush; D. Robinson; J. W. Schrock; J. P. D'Etienne; C. J. Hogan; A. Osborne; R. Riviello; S. Young | Peacock, et. al. Jan 2022 | Utility of COVID-19 antigen testing in the emergency department | 2022-01-04 | BinaxNOW (Abbott) |  |  | >14 days |  |  |  |  | 50 | 11.8,88.2 | 98 | 89.4,99.9 | 75 | 94.2 | PPV CI95 [19.4-99.4] NPV CI95 [84.1-98.8]
Amy | Sherli | 177 | W. F. Peacock; K. M. Soto-Ruiz; S. L. House; C. M. Cannon; G. Headden; B. Tiffany; S. Motov; K. Merchant-Borna; A. M. Chang; C. Pearson; B. W. Patterson; A. E. Jones; J. Miller; J. Varon; A. Bastani; C. Clark; Z. Rafique; B. Kea; J. Eppensteiner; J. M. Williams; S. A. Mahler; B. E. Driver; P. Hendry; E. Quackenbush; D. Robinson; J. W. Schrock; J. P. D'Etienne; C. J. Hogan; A. Osborne; R. Riviello; S. Young | Peacock, et. al. Jan 2022 | Utility of COVID-19 antigen testing in the emergency department | 2022-01-05 | BinaxNOW (Abbott) |  |  | Asymptomatic |  |  |  |  | 66.7 | 22.3,95.7 | 99.1 | 94.9,100 | 80 | 98.1 | PPV CI95 [28.4-99.5] NPV CI95 [93.4-99.8]
Ellyn | Sherli | 237 | Peña M. | Peña, et. al. Apr 2021 | Performance of SARS-CoV-2 rapid antigen test compared with real-time RT-PCR in asymptomatic individuals | 2021-04-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 51 | 3 | 766 | 22 | 69.86 | 58.56,9.18 TYPO IN PAPER | 99.61 | 98.86,99.87 | 94.44 | 97.2 |  | 69.863 | 99.6099 | 94.4444 | 97.2081
Ellyn | Sherli | 238 | Peña-Rodríguez M. | Peña-Rodriguez, et. al. Feb 2021 | Performance evaluation of a lateral flow assay for nasopharyngeal antigen detection for SARS-CoV-2 diagnosis | 2021-02-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 79 | 0 | 265 | 25 | 75.9 | 66.5,83.8% | 100 | 98.6,100 | 100 | 91.4 |  | 75.9615 | 100 | 100 | 91.3793
Amy | Sherli | 180 | C. Peronace; R. Tallerico; M. Colosimo; V. Sacco; R. Talarico; M. De Fazio; F. Pasceri; I. Talotta; G. Panduri; J.-H. Kim; E. Cione; P. Minchella | Peronace, et. al. May 2022 | Validation of GeneFinder COVID-19 Ag Plus Rapid Test and Its Potential Utility to Slowing Infection Waves: A Single-Center Laboratory Evaluation Study | 2022-05-01 | GeneFinder COVID-19 Ag Plus Rapid Test |  | detects nucleocapsid protein | Overall | 145 | 1 | 451 | 6 | 96.03 | 91.55,98.53 | 99.78 | 98.77,99.99 | 99.32 | 98.69 | PPV CI95 [95.43-99.90] NPV CI95 [97.17-99.40]
Ellyn | Sherli | 241 | Pilarowski GenayLebel | Pilarowski, et. al. Jan 2021 | Performance characteristics of a rapid SARS-CoV-2 antigen detection assay at a public plaza testing site in San Francisco | 2021-01-01 | BinaxNOW (Abbott) |  |  | Overall | 15 | 0 | 845 | 11 | 57.7 | 36.9,76.6 | 100 | 99.6,100 | 100 | 98.7 |  | 57.6923 | 100 | 100 | 98.715
Ellyn | Sherli | 241 | Pilarowski GenayLebel | Pilarowski, et. al. Jan 2021 | Performance characteristics of a rapid SARS-CoV-2 antigen detection assay at a public plaza testing site in San Francisco | 2021-01-01 | BinaxNOW (Abbott) |  |  | Ct<30 | 14 | 1 | 855 | 1 | 93.3 | 68.1,99.8 | 99.9 | 99.4,99.9 | 93.3 | 99.9 |  | 93.3333 | 99.8832 | 93.3333 | 99.8832
Ellyn | Sherli | 241 | Pilarowski GenayLebel | Pilarowski, et. al. Jan 2021 | Performance characteristics of a rapid SARS-CoV-2 antigen detection assay at a public plaza testing site in San Francisco | 2021-01-01 | BinaxNOW (Abbott) |  |  | Ct<33 | 15 | 0 | 855 | 1 | 93.8 | 69.8,99.8 | 100 | 99.6,100 | 100 | 99.9 |  | 93.75 | 100 | 100 | 99.8832
Amy | Sherli | 244 | Pollock N. | Pollock, et. al. Apr 2021 | Performance and Implementation Evaluation of the Abbott BinaxNOW Rapid Antigen Test in a High-Throughput Drive-Through Community Testing Site in Massachusetts | 2021-04-01 | BinaxNOW (Abbott) |  |  | Overall | 132 | 5 | 1219 | 25 | 84.1 | 77.4,89.4 | 99.6 | 99.1,99.9 | 96.4 | 98 |  | 84.0764 | 99.5915 | 96.3504 | 97.9904
Amy | Sherli | 244 | Pollock N. | Pollock, et. al. Apr 2021 | Performance and Implementation Evaluation of the Abbott BinaxNOW Rapid Antigen Test in a High-Throughput Drive-Through Community Testing Site in Massachusetts | 2021-04-01 | BinaxNOW (Abbott) |  |  | Asymptomatic | 40 | 4 | 913 | 17 | 70.2 | 56.6,81.6 | 99.6 | 98.9,99.9 | 90.9 | 98.1 |  | 70.1754 | 99.5638 | 90.9091 | 98.172
Amy | Sherli | 244 | Pollock N. | Pollock, et. al. Apr 2021 | Performance and Implementation Evaluation of the Abbott BinaxNOW Rapid Antigen Test in a High-Throughput Drive-Through Community Testing Site in Massachusetts | 2021-04-01 | BinaxNOW (Abbott) |  |  | Symptomatic >7 days | 10 | 1 | 35 | 5 | 66.7 | 38.4,88.2 | 97.2 | 85.5,99.9 | 90 | 88.2 |  | 66.6667 | 97.2222 | 90.9091 | 87.5
Amy | Sherli | 244 | Pollock N. | Pollock, et. al. Apr 2021 | Performance and Implementation Evaluation of the Abbott BinaxNOW Rapid Antigen Test in a High-Throughput Drive-Through Community Testing Site in Massachusetts | 2021-04-01 | BinaxNOW (Abbott) |  |  | Symptomatic ≤7 days | 82 | 0 | 270 | 3 | 96.5 | 90.0,99.3 | 100 | 98.6,100 | 100 | 98.9 |  | 96.4706 | 100 | 100 | 98.9011
Chathu | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Poopalasingam, et. al. Feb 2022 | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 43 | 2 | 610 | 32 | 57.3 | 46.1,67.9 | 99.7 | 98.8,99.9 | 100 | 94.6 | PPV 95% CI = 85.2, 99.2; NPV 95% CI = 93.1, 96.5
Chathu | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Poopalasingam, et. al. Feb 2022 | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptomatic |  |  |  |  | 71.1 | 56.6,82.3 | 100 | 98.3,100 | 100 | 94.6 | PPV 95% CI = 89.3,100.0); NPV, 95% CI = 90.9, 96.8
Chathu | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Poopalasingam, et. al. Feb 2022 | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-03 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic |  |  |  |  | 36.7 | 21.9,54.5 | 99.5 | 98.1,99.9 | 84.6 | 95.3 | PPV 95% CI = 57.8,97.3 NPV 95% CI = 92.8, 97.0
Chathu | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Poopalasingam, et. al. Feb 2022 | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-04 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct ≤ 20 |  |  |  |  | 100 | 82.4,100
Chathu | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Poopalasingam, et. al. Feb 2022 | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-05 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct ≤ 25 |  |  |  |  | 94.4 | 81.9,99
Chathu | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Poopalasingam, et. al. Feb 2022 | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-06 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct ≤ 30 |  |  |  |  | 81.1 | 68.6,89.4
Chathu | Sherli | 190 | N. Poopalasingam; M. Korenkov; A. Ashurov; J. Strobel; I. Fish; M. Hellmich; H. Gruell; C. Lehmann; E. Heger; F. Klein | Poopalasingam, et. al. Feb 2023 | Determining the reliability of rapid SARS-CoV-2 antigen detection in fully vaccinated individuals | 2022-02-07 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Ct > 30 |  |  |  |  | 0 | 0,14.9
Chathu | Sherli | 193 | M. Prost | Prost, et. al. Dec 2021 | Evaluation of a Rapid in Vitro Diagnostic Test Device for Detection of SARS-CoV-2 Antigen in Nasal Swabs | 2021-12-01 | SARS-CoV-2 Antigen Rapid Test (Hangzhou Immuno Biotech Co Ltd, China). |  |  | Overall | 214 | 1 | 199 | 6 | 97.3 | 94.2,99.0 | 99.5 | 97.3,100
Amy | Sherli | 195 | M. M. Rahman; A. F. Hoque; Y. Karim; Z. Kawser; A. B. Siddik; M. K. Sumiya; A. Siddika; M. S. Alam; A. N. Alam; M. Alam; M. E. Hossain; S. Banu; F. Qadri; T. Shirin; M. Rahman; M. Z. Rahman | Rahman, et. al. Nov 2021 | Clinical evaluation of SARS-CoV-2 antigen-based rapid diagnostic test kit for detection of COVID-19 cases in Bangladesh | 2021-11-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Adults | 253 | 0 | 566 | 42 | 85.76 | 81.25,89.54
Lupe | Sherli | 196 | A. F. A. S. N. Y. M. T. A. Raja Kamran | Raja Kamran, et. al. Sep 2021 | DIAGNOSTIC ACCURACY OF PANBIO COVID-19 RAPID ANTIGEN METHOD FOR SCREENING IN EMERGENCY CASES | 2021-09-01 | PanBio (Abbott) |  |  | Overall | 19 | 0 | 1518 | 2 | 90.47 |  | 100 |  | 100 | 99.8
Lupe | Sherli | 197 | N. Rana; P. Thakur; S. Saxena; V. Manchanda; R. Arora | Rana, et. al. Sept 2021 | Evaluation of the currently used antigen-based Rapid Diagnostic Test for the detection of SARS CoV-2 virus in respiratory specimens | 2021-09-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall |  |  |  |  | 37.5 |  | 99.79 |  | 96.4 | 91.6
Amy | Sherli | 262 | Rastawicki W. | Rastawicki, et. al. Jan 2021 | Evaluation of PCL rapid point of care antigen test for detection of SARS-CoV-2 in nasopharyngeal swabs | 2021-01-01 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | Overall | 14 | 1 | 71 | 22 | 38.9 |  | 83.3 |  | 96 | 50 |  | 38.8889 | 98.6111 | 93.3333 | 76.3441
Amy | Sherli | 262 | Rastawicki W. | Rastawicki, et. al. Jan 2021 | Evaluation of PCL rapid point of care antigen test for detection of SARS-CoV-2 in nasopharyngeal swabs | 2021-01-01 | PCL COVID19 Ag Rapid FIA Antigen Test (PCL) |  |  | Ct<25 |  |  |  |  | 92.2 | [6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 266 | Reza SoleimaniCorentin | Reza, et. al. May 2021 | Rapid COVID-19 antigenic tests: Usefulness of a modified method for diagnosis | 2021-05-01 | PanBio (Abbott) |  |  | Ct<25 |  |  |  |  | 96.8 | 92.7,99 |  |  |  | 97.1 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 266 | Reza SoleimaniCorentin | Reza, et. al. May 2021 | Rapid COVID-19 antigenic tests: Usefulness of a modified method for diagnosis | 2021-05-01 | PanBio (Abbott) |  |  | Ct<30 |  |  |  |  | 88.8 | 83.5,92.8 |  |  |  | 88.5 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 266 | Reza SoleimaniCorentin | Reza, et. al. May 2021 | Rapid COVID-19 antigenic tests: Usefulness of a modified method for diagnosis | 2021-05-01 | PanBio (Abbott) |  |  | Overall |  |  |  |  | 75 | 68.9,80.4 |  |  |  | 74.5 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 266 | Reza SoleimaniCorentin | Reza, et. al. May 2021 | Rapid COVID-19 antigenic tests: Usefulness of a modified method for diagnosis | 2021-05-01 | COVID19-Speed/Biospeedia COVID19 Antigen test (Biospeedia) |  |  | Overall |  |  |  |  | 65.5 | 59.0,71.6 | 100 |  | 100 | 67.9 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 266 | Reza SoleimaniCorentin | Reza, et. al. May 2021 | Rapid COVID-19 antigenic tests: Usefulness of a modified method for diagnosis | 2021-05-01 | COVID19-Speed/Biospeedia COVID19 Antigen test (Biospeedia) |  |  | Ct<25 |  |  |  |  | 90.4 | 84.6,94.5 |  |  |  | 91.9 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 266 | Reza SoleimaniCorentin | Reza, et. al. May 2021 | Rapid COVID-19 antigenic tests: Usefulness of a modified method for diagnosis | 2021-05-01 | COVID19-Speed/Biospeedia COVID19 Antigen test (Biospeedia) |  |  | Ct<30 |  |  |  |  | 77.6 | 71.1,83.2 |  |  |  | 79.3 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Lupe | Sherli | 203 | J. Rohde; W. Himmel; C. Hofinger; T. T. Lâm; H. Schrader; J. Wallstabe; O. Kurzai; I. Gágyor | Rohde, et. al. Feb 2022 | Diagnostic accuracy and feasibility of a rapid SARS-CoV-2 antigen test in general practice - a prospective multicenter validation and implementation study | 2022-02-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  | 1450 | Overall | 108 | 7 | 1305 | 30 | 78.3 | 70.9,84.6 | 99.5 | 99,99.8 | 93.9 | 97.8 | PPV CI: 88.6,97.3 NPV CI: 96.9,98.5
Lupe | Sherli | 203 | J. Rohde; W. Himmel; C. Hofinger; T. T. Lâm; H. Schrader; J. Wallstabe; O. Kurzai; I. Gágyor | Rohde, et. al. Feb 2022 | Diagnostic accuracy and feasibility of a rapid SARS-CoV-2 antigen test in general practice - a prospective multicenter validation and implementation study | 2022-02-02 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  | 22 | Ct < 20
Lupe | Sherli | 203 | J. Rohde; W. Himmel; C. Hofinger; T. T. Lâm; H. Schrader; J. Wallstabe; O. Kurzai; I. Gágyor | Rohde, et. al. Feb 2022 | Diagnostic accuracy and feasibility of a rapid SARS-CoV-2 antigen test in general practice - a prospective multicenter validation and implementation study | 2022-02-03 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  | 34 | Ct = 20-25
Lupe | Sherli | 203 | J. Rohde; W. Himmel; C. Hofinger; T. T. Lâm; H. Schrader; J. Wallstabe; O. Kurzai; I. Gágyor | Rohde, et. al. Feb 2022 | Diagnostic accuracy and feasibility of a rapid SARS-CoV-2 antigen test in general practice - a prospective multicenter validation and implementation study | 2022-02-04 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  | 20 | Ct = 25-30
Lupe | Sherli | 203 | J. Rohde; W. Himmel; C. Hofinger; T. T. Lâm; H. Schrader; J. Wallstabe; O. Kurzai; I. Gágyor | Rohde, et. al. Feb 2022 | Diagnostic accuracy and feasibility of a rapid SARS-CoV-2 antigen test in general practice - a prospective multicenter validation and implementation study | 2022-02-05 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  | 12 | CT > 30 |  |  |  |  | 25 | 7.7,57.2
Amy | Sherli | 212 | N. Salcedo; B. F. Sena; X. Qu; B. B. Herrera | Salcedo, et. al. Feb 2022 | Comparative Evaluation of Rapid Isothermal Amplification and Antigen Assays for Screening Testing of SARS-CoV-2 | 2022-02-01 | Custom/Novel/In-house |  | E25Bio, Inc | Asymptomatic | 29 | 1 | 75 | 6 | 82.86 | 66.35,93.44 | 98.68 | 92.89,99.97 | 96.67 | 92.59 | PPV CI95 [80.45-99.51] NPV CI95 [85.78-96.28]
Amy | Sherli | 212 | N. Salcedo; B. F. Sena; X. Qu; B. B. Herrera | Salcedo, et. al. Feb 2022 | Comparative Evaluation of Rapid Isothermal Amplification and Antigen Assays for Screening Testing of SARS-CoV-2 | 2022-02-02 | Custom/Novel/In-house |  | E25Bio, Inc | Symptomatic | 22 | 0 | 38 | 2 | 91.67 | 73.0,98.97 | 100 | 90.75,100.0 | 100 | 95 | NPV CI95 [83.45-98.62] PPV CI95 not given
Amy | Sherli | 212 | N. Salcedo; B. F. Sena; X. Qu; B. B. Herrera | Salcedo, et. al. Feb 2022 | Comparative Evaluation of Rapid Isothermal Amplification and Antigen Assays for Screening Testing of SARS-CoV-2 | 2022-02-03 | Custom/Novel/In-house |  | E25Bio, Inc | Overall (added above #) | 51 | 1 | 113 | 8
Amy | Sherli | 274 | Salvagno G. | Salvagno, et. al. Jan 2021 | Clinical assessment of the Roche SARS-CoV-2 rapid antigen test | 2021-01-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall |  |  |  |  | 72.5 | 64.6,79.5 | 99.4 | 96.8,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Lupe | Sherli | 213 | G. L. Salvagno; G. Gianfilippi; L. Pighi; S. De Nitto; B. M. Henry; G. Lippi | Salvagno, et. al. May 2021 | Real-world assessment of Fluorecare SARS-CoV-2 Spike Protein Test Kit | 2021-05-01 | Fluorecare (Colloidal Gold/Fluorescent) SARS-CoV-2 Spike Protein Test kit (Shenzen Microprofit) |  | 354 | Overall |  |  |  |  | 27.5 | 21.8,33.7 | 99.2 | 95.5,100 | 98.5 | 41.5 | PPV CI: 90,99.8 NPV CI: 39.6,43.5
Lupe | Sherli | 213 | G. L. Salvagno; G. Gianfilippi; L. Pighi; S. De Nitto; B. M. Henry; G. Lippi | Salvagno, et. al. May 2021 | Real-world assessment of Fluorecare SARS-CoV-2 Spike Protein Test Kit | 2021-05-02 | Fluorecare (Colloidal Gold/Fluorescent) SARS-CoV-2 Spike Protein Test kit (Shenzen Microprofit) |  | 96 | CT<=30 |  |  |  |  | 66.7 | 56.3,76
Lupe | Sherli | 213 | G. L. Salvagno; G. Gianfilippi; L. Pighi; S. De Nitto; B. M. Henry; G. Lippi | Salvagno, et. al. May 2021 | Real-world assessment of Fluorecare SARS-CoV-2 Spike Protein Test Kit | 2021-05-03 | Fluorecare (Colloidal Gold/Fluorescent) SARS-CoV-2 Spike Protein Test kit (Shenzen Microprofit) |  | 63 | CT<=25 |  |  |  |  | 90.5 | 80.4,96.4
Lupe | Sherli | 213 | G. L. Salvagno; G. Gianfilippi; L. Pighi; S. De Nitto; B. M. Henry; G. Lippi | Salvagno, et. al. May 2021 | Real-world assessment of Fluorecare SARS-CoV-2 Spike Protein Test Kit | 2021-05-04 | Fluorecare (Colloidal Gold/Fluorescent) SARS-CoV-2 Spike Protein Test kit (Shenzen Microprofit) |  | 36 | CT<=20 |  |  |  |  | 100 | 90.3,100
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-01 | Covios COVID-19 Antigen Rapid Dianostic test |  | Can treat this line as an overall | Self-taken swab | 67 | 1 | 173 | 7 | 90.5 | 83.9,97.2 | 99.4 | 98.3,100.0 | 98.5 | 96.1 | PPV CI95 [95.7-100.0] NPV CI95 [93.3-98.9]
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-02 | Covios COVID-19 Antigen Rapid Dianostic test |  | Can treat this line as an overall | Health-care worker taken swab | 58 | 2 | 172 | 16 | 78.4 | 69.0,87.8 | 98.9 | 97.3,100.0 | 96.7 | 91.5 | PPV CI95 [92.1-100.0] NPV CI95 [87.5-95.5]
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-03 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | Self-taken CT<20 | 32 |  |  | 0 | 100 | 89.1,100.0
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-04 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | Self-taken CT 20-24.9 | 22 |  |  | 1 | 91.7 | 73.0,99.0
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-05 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | Self-taken CT 25-29.9 | 12 |  |  | 3 | 80 | 51.9,95.7
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-06 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | Self-taken CT 30-34.9 | 1 |  |  | 2 | 33.3 | 0.8,90.6
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-07 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | Self-taken CT >=35 | 0 |  |  | 1 | 0 | 0.0,97.5
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-08 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | HCW-taken CT<20 | 32 |  |  | 0 | 100 | 89.1,100.0
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-09 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | HCW-taken CT 20-24.9 | 20 |  |  | 4 | 83.3 | 62.6,95.3
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-10 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | HCW-taken CT 25-29.9 | 6 |  |  | 9 | 40 | 16.3,67.7
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-11 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | HCW-taken CT 30-34.9 | 1 |  |  | 2 | 33.3 | 0.8,90.6
Amy | Sherli | 218 | H. R. Savage; L. Finch; R. Body; R. L. Watkins; G. Hayward; E. Cook; A. I. Cubas-Atienzar; L. E. Cuevas; P. MacPherson; E. R. Adams | Savage, et. al. Jun 2022 | A prospective diagnostic evaluation of accuracy of self-taken and healthcare worker-taken swabs for rapid COVID-19 testing | 2022-06-12 | Covios COVID-19 Antigen Rapid Dianostic test |  |  | HCW-taken CT >=35 | 0 |  |  | 1 | 0 | 0.0,97.5
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | PanBio (Abbott) |  |  | Asymptomatic |  |  |  |  | 38.5 | 18,64 | 71.4 | 45,88 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | PanBio (Abbott) |  |  | BAL |  |  |  |  | 72.7 | 43,90 | 100 | 34.0,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | PanBio (Abbott) |  |  | Symptomatic |  |  |  |  | 40 | 17,69 | 84.6 | 58,96 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | PanBio (Abbott) |  |  | Overall |  |  |  |  | 50 | 35,64 | 77.4 | 60,89 |  | 53.3 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | RapiGen (BioCredit) |  |  | Overall |  |  |  |  | 33.3 | 21,48 | 87.1 | 71,95 | 77.8 | 49.1 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | RapiGen (BioCredit) |  |  | Asymptomatic |  |  |  |  | 30.8 | 13,58 | 92.9 | 69,99 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | RapiGen (BioCredit) |  |  | BAL |  |  |  |  | 54.6 | 28,78 | 100 | 34.0,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | RapiGen (BioCredit) |  |  | Symptomatic |  |  |  |  | 30 | 11,60 | 76.9 | 50,92 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Asymptomatic |  |  |  |  | 84.6 | 57,96 | 14.3 | 4.0,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | BAL |  |  |  |  | 72.7 | 52,95 | 100 | 34.0,100 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptomatic |  |  |  |  | 100 | 72100 | 7.7 | 1,33 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 278 | Schildgen V. | Schildgen, et. al. Jan 2021 | Limits and opportunities of sars-cov-2 antigen rapid tests: An experienced-based perspective | 2021-01-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall |  |  |  |  | 88.1 | 75,95 | 19.4 | 9,36 | 59.7 | 54.6 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Amy | Sherli | 227 | R. A. P. Selvabai; L. V. Koshy; P. Shanmugam | Selvebai, et. al. Apr 2022 | Diagnostic Efficacy of COVID-19 Rapid Antigen Detection Card in Diagnosis of SARS-CoV-2 | 2022-04-01 | Athenese-DX COVID-19 RAT kit |  |  | Overall | 230 | 0 | 289 | 80 | 74.19 |  | 100 |  | 100 | 78.3
Amy | Sherli | 227 | R. A. P. Selvabai; L. V. Koshy; P. Shanmugam | Selvebai, et. al. Apr 2022 | Diagnostic Efficacy of COVID-19 Rapid Antigen Detection Card in Diagnosis of SARS-CoV-2 | 2022-04-02 | Athenese-DX COVID-19 RAT kit |  |  | CT between 10-20 | 64 |  |  | 2
Amy | Sherli | 227 | R. A. P. Selvabai; L. V. Koshy; P. Shanmugam | Selvebai, et. al. Apr 2022 | Diagnostic Efficacy of COVID-19 Rapid Antigen Detection Card in Diagnosis of SARS-CoV-2 | 2022-04-03 | Athenese-DX COVID-19 RAT kit |  |  | CT between 20-30 | 116 |  |  | 33
Amy | Sherli | 227 | R. A. P. Selvabai; L. V. Koshy; P. Shanmugam | Selvebai, et. al. Apr 2022 | Diagnostic Efficacy of COVID-19 Rapid Antigen Detection Card in Diagnosis of SARS-CoV-2 | 2022-04-04 | Athenese-DX COVID-19 RAT kit |  |  | CT >30 | 14 |  |  | 35
Amy | Sherli | 227 | R. A. P. Selvabai; L. V. Koshy; P. Shanmugam | Selvebai, et. al. Apr 2022 | Diagnostic Efficacy of COVID-19 Rapid Antigen Detection Card in Diagnosis of SARS-CoV-2 | 2022-04-05 | Athenese-DX COVID-19 RAT kit |  |  | Symptomatic |  |  |  |  | 73.7 |  | 100
Amy | Sherli | 227 | R. A. P. Selvabai; L. V. Koshy; P. Shanmugam | Selvebai, et. al. Apr 2022 | Diagnostic Efficacy of COVID-19 Rapid Antigen Detection Card in Diagnosis of SARS-CoV-2 | 2022-04-06 | Athenese-DX COVID-19 RAT kit |  |  | Asymptomatic |  |  |  |  | 79 |  | 100
Amy | Sherli | 230 | J. L. V. Shaw; V. Deslandes; J. Smith; M. Desjardins | Shaw, et. al. Jul 2021 | Evaluation of the Abbott Panbio(TM) COVID-19 Ag rapid antigen test for the detection of SARS-CoV-2 in asymptomatic Canadians | 2021-07-01 | PanBio (Abbott) |  |  | Asymptomatic | 6 | 0 | 2996 | 5 | 54.5 |  |  |  | 100 | 99.8
Amy | Sherli | 233 | Z. K. Siddiqui; M. Chaudhary; M. L. Robinson; A. B. McCall; R. Peralta; R. Esteve; C. W. Callahan; Y. C. Manabe; J. D. Campbell; J. K. Johnson; M. Elhabashy; M. Kantsiper; J. R. Ficke | Siddiqui, et. al. Dec 2021 | Implementation and Accuracy of BinaxNOW Rapid Antigen COVID-19 Test in Asymptomatic and Symptomatic Populations in a High-Volume Self-Referred Testing Site | 2021-12-01 | BinaxNOW (Abbott) |  |  | Overall | 179 | 13 | 5826 | 43 | 81 | 75,86 | 99.8 | 100.0,100.0 | 93 | 99 | PPV CI95 [89,96] NPV CI95 [99-100]
Amy | Sherli | 233 | Z. K. Siddiqui; M. Chaudhary; M. L. Robinson; A. B. McCall; R. Peralta; R. Esteve; C. W. Callahan; Y. C. Manabe; J. D. Campbell; J. K. Johnson; M. Elhabashy; M. Kantsiper; J. R. Ficke | Siddiqui, et. al. Dec 2021 | Implementation and Accuracy of BinaxNOW Rapid Antigen COVID-19 Test in Asymptomatic and Symptomatic Populations in a High-Volume Self-Referred Testing Site | 2021-12-02 | BinaxNOW (Abbott) |  |  | Symptomatic | 117 | 6 | 501 | 18 | 87 | 80,91 | 99 |  | 95 |  | PPV CI95 [90,98]
Amy | Sherli | 233 | Z. K. Siddiqui; M. Chaudhary; M. L. Robinson; A. B. McCall; R. Peralta; R. Esteve; C. W. Callahan; Y. C. Manabe; J. D. Campbell; J. K. Johnson; M. Elhabashy; M. Kantsiper; J. R. Ficke | Siddiqui, et. al. Dec 2021 | Implementation and Accuracy of BinaxNOW Rapid Antigen COVID-19 Test in Asymptomatic and Symptomatic Populations in a High-Volume Self-Referred Testing Site | 2021-12-03 | BinaxNOW (Abbott) |  |  | Asymptomatic | 62 | 7 | 5324 | 25 | 71 | 61,80 | 99 |  | 90 |  | PPV CI95 [80-96]
Amy | Sherli | 233 | Z. K. Siddiqui; M. Chaudhary; M. L. Robinson; A. B. McCall; R. Peralta; R. Esteve; C. W. Callahan; Y. C. Manabe; J. D. Campbell; J. K. Johnson; M. Elhabashy; M. Kantsiper; J. R. Ficke | Siddiqui, et. al. Dec 2021 | Implementation and Accuracy of BinaxNOW Rapid Antigen COVID-19 Test in Asymptomatic and Symptomatic Populations in a High-Volume Self-Referred Testing Site | 2021-12-04 | BinaxNOW (Abbott) |  |  | Symptomatic =<7 days | 112 | 4 | 442 | 15 | 88
Amy | Sherli | 233 | Z. K. Siddiqui; M. Chaudhary; M. L. Robinson; A. B. McCall; R. Peralta; R. Esteve; C. W. Callahan; Y. C. Manabe; J. D. Campbell; J. K. Johnson; M. Elhabashy; M. Kantsiper; J. R. Ficke | Siddiqui, et. al. Dec 2021 | Implementation and Accuracy of BinaxNOW Rapid Antigen COVID-19 Test in Asymptomatic and Symptomatic Populations in a High-Volume Self-Referred Testing Site | 2021-12-05 | BinaxNOW (Abbott) |  |  | Symptomatic >7 days | 5 | 2 | 55 | 3 | 63
Amy | Sherli | 233 | Z. K. Siddiqui; M. Chaudhary; M. L. Robinson; A. B. McCall; R. Peralta; R. Esteve; C. W. Callahan; Y. C. Manabe; J. D. Campbell; J. K. Johnson; M. Elhabashy; M. Kantsiper; J. R. Ficke | Siddiqui, et. al. Dec 2021 | Implementation and Accuracy of BinaxNOW Rapid Antigen COVID-19 Test in Asymptomatic and Symptomatic Populations in a High-Volume Self-Referred Testing Site | 2021-12-06 | BinaxNOW (Abbott) |  |  | CT <25 | 113 | 0 | 0 | 4 | 97
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-01 | PanBio (Abbott) |  | 547 participants | Overall | 86 | 6 | 333 | 122 | 41.3 | 34.6,48.4 | 98.2 | 96.2,99.3 | 93.3 | 52.9 | PPV CI95 [88.8-97.8] NPV CI95 [48.6-57.1]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-02 | Standard Q COVID-19 Ag (SD Biosensor) |  | 705 participants | Overall | 167 | 8 | 326 | 204 | 45 | 39.9,50.2 | 97.6 | 95.3,99.0 | 95.4 | 66.4 | PPV CI95 [92.3-98.5] NPV CI95 [62.3-70.4]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-03 | PanBio (Abbott) |  |  | Symptomatic |  |  |  |  | 43.2 | 36.0,50.5 | 97.9 | 95.5,99.2 | 93 | 55.3 | PPV CI95 [88.4-97.7] NPV CI95 [50.8-59.8]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-04 | PanBio (Abbott) |  |  | Asymptomatic |  |  |  |  | 22.2 | 6.4,47.6 | 100 | 93.2,N/A | 100 | 30.4 | PPV CI95 [100-NA] NPV CI95 [19.4-41.4]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-05 | PanBio (Abbott) |  |  | Symptoms >7days |  |  |  |  | 25.5 | 14.3,39.6 | 98.4 | 91.6,100 | 92.9 | 50.4 | PPV CI95 [79.4-100.0] NPV CI95 [41.0-59.8]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-06 | PanBio (Abbott) |  |  | symptoms =< 7days |  |  |  |  | 49.6 | 41.1-58.2 | 97.7 | 94.8,99.3 | 93.1 | 57.3 | PPV CI95 [88.1-98] NPV CI95 [52.2-62.4]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-07 | PanBio (Abbott) |  |  | symptoms =< 5 days |  |  |  |  | 53.3 | 42.6,63.7 | 97.4 | 93.5,99.3 | 91.3 | 56.5 | PPV CI95 [84.7-98.0] NPV CI95 [50.2-62.8]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-08 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptomatic |  |  |  |  | 49.4 | 43.8,55.0 | 97.6 | 94.8,99.1 | 96.4 | 71.5 | PPV CI95 [93.6-99.2] NPV CI95 [67.2-75.9]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-09 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic |  |  |  |  | 13.3 | 5.1,26.8 | 97.6 | 91.7,99.7 | 75 | 37.6 | PPV CI95 [45.0-100] NPV CI95 [29.0-46.3]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-10 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptoms >7days |  |  |  |  | 35.2 | 25.3,46.1 | 95.6 | 84.9,99.5 | 93.9 | 74.3 | PPV CI95 [85.8-100.0] NPV CI95 [65.7-82.8]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-11 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | symptoms =< 7days |  |  |  |  | 54.5 | 47.9,61.0 | 98 | 95.1,99.5 | 96.9 | 71.1 | PPV CI95 [94.0-99.9] NPV CI95 [66.0-76.2]
Amy | Sherli | 235 | N. Sitoe; J. Sambo; N. Nguenha; J. Chilaule; I. Chelene; O. Loquiha; C. Mudenyanga; S. Viegas; J. Cunningham; I. Jani | Sitoe, et. al. Feb 2022 | Performance Evaluation of the STANDARD(TM) Q COVID-19 and Panbio(TM) COVID-19 Antigen Tests in Detecting SARS-CoV-2 during High Transmission Period in Mozambique | 2022-02-12 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | symptoms =< 5 days |  |  |  |  | 61.9 | 53.9,69.4 | 97.9 | 94.1,99.6 | 97.1 | 73.8 | PPV CI95 [93.8-100.0] NPV CI95 [67.8-79.9]
Amy | Sherli | 236 | M. Skvarč | Skvarč Apr 2021 | Clinical validation of two immunochromatographic SARS-CoV-2 antigen tests in near hospital facilities | 2021-04-01 | Immupass VivaDiag (VivaChek Biotech) |  | 794 samples | Overall | 136 | 0 | 644 | 14 | 90.6 | 84.94,94.36 | 100 | 99.41,100.0
Amy | Sherli | 236 | M. Skvarč | Skvarč Apr 2022 | Clinical validation of two immunochromatographic SARS-CoV-2 antigen tests in near hospital facilities | 2021-04-02 | Alltest Covid19 Ag test |  | 333 samples | Overall | 134 | 0 | 312 | 8 | 94.37 | 89.20,97.54 | 100 | 98.83,100.0
Amy | Sherli | 288 | Smith R. | Smith, et. al. June 2021 | Clinical Evaluation of Sofia Rapid Antigen Assay for Detection of Severe Acute Respiratory Syndrome Coronavirus 2 among Emergency Department to Hospital Admissions | 2021-06-01 | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | Overall | 180 | 7 | 2645 | 55 | 76.6 | 71,82 | 99.7 | 99.0,100 | 96.3 | 97.9 |  | 76.5957 | 99.736 | 96.2567 | 97.963
Amy | Sherli | 288 | Smith R. | Smith, et. al. June 2021 | Clinical Evaluation of Sofia Rapid Antigen Assay for Detection of Severe Acute Respiratory Syndrome Coronavirus 2 among Emergency Department to Hospital Admissions | 2021-06-01 | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | Asymptomatic | 33 | 4 | 1160 | 9 | 78.6 | 67,91 | 99.7 | 99.0,100 | 89.2 | 99.2 |  | 78.5714 | 99.6564 | 89.1892 | 99.2301
Amy | Sherli | 288 | Smith R. | Smith, et. al. June 2021 | Clinical Evaluation of Sofia Rapid Antigen Assay for Detection of Severe Acute Respiratory Syndrome Coronavirus 2 among Emergency Department to Hospital Admissions | 2021-06-01 | Sofia SARS Rapid Antigen FIA/Sofia 2 (Quidel) |  |  | Symptomatic | 147 | 2 | 1480 | 46 | 76.2 | 70,82 | 99.9 | 99.0,100 | 98.7 | 89.2 |  | 76.1658 | 99.865 | 98.6577 | 96.9856
Amy | Sherli | 241 | J. Stohr; V. F. Zwart; G. Goderski; A. Meijer; C. R. S. Nagel-Imming; M. F. Q. Kluytmans-van den Bergh; S. D. Pas; F. van den Oetelaar; M. Hellwich; K. H. Gan; A. Rietveld; J. J. Verweij; J. L. Murk; W. van den Bijllaardt; J. Kluytmans | Stohr, et. al. Aug 2021 | Self-testing for the detection of SARS-CoV-2 infection with rapid antigen tests for people with suspected COVID-19 in the community | 2021-08-01 | BD Veritor COVID-19 Rapid Antigen Test (Becton-Dickinson) |  | 1595 | Overall | 86 | 2 | 1379 | 89 | 49.1 | 41.7,56.5 | 99.9 | 99.7,100.0
Amy | Sherli | 241 | J. Stohr; V. F. Zwart; G. Goderski; A. Meijer; C. R. S. Nagel-Imming; M. F. Q. Kluytmans-van den Bergh; S. D. Pas; F. van den Oetelaar; M. Hellwich; K. H. Gan; A. Rietveld; J. J. Verweij; J. L. Murk; W. van den Bijllaardt; J. Kluytmans | Stohr, et. al. Aug 2021 | Self-testing for the detection of SARS-CoV-2 infection with rapid antigen tests for people with suspected COVID-19 in the community | 2021-08-02 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  | 1606 | Overall | 118 | 4 | 1387 | 74 | 61.5 | 54.6,68.3 | 99.7 | 99.4,99.9
Amy | Sherli | 243 | K. Surasi; K. J. Cummings; C. Hanson; M. K. Morris; M. Salas; D. Seftel; L. Ortiz; R. Thilakaratne; C. Stainken; D. A. Wadford | Surasi, et. al. Nov 2021 | Effectiveness of Abbott BinaxNOW Rapid Antigen Test for Detection of SARS-CoV-2 Infections in Outbreak among Horse Racetrack Workers, California, USA | 2021-11-01 | BinaxNOW (Abbott) |  |  | Overall | 55 | 0 | 642 | 72 | 43.3 | 34.6,52.4 | 100 | 99.4,100.0 | 100 | 89.9 | PPV CI95 [93.5-100] NPV CI95 [87.5–92.0]
Amy | Sherli | 244 | H. Suzuki; Y. Akashi; D. Kato; Y. Takeuchi; Y. Kiyasu; N. Terada; Y. Kurihara; M. Kuwahara; S. Muramatsu; A. Ueda; S. Notake; K. Nakamura | Suzuki, et. al. May 2022 | Analytical performance of rapid antigen tests for the detection of SARS-CoV-2 during widespread circulation of the Omicron variant | 2022-05-01 | QuickNavi-COVID19 Ag |  |  | Overall | 393 | 3 | 653 | 24 | 94.2 | 91.6,96.3 | 99.5 | 98.7,99.9 | 99.2 | 96.5 | PPV CI95 [97.8-99.8] NPV CI95 [94.8-97.7]
Amy | Sherli | 244 | H. Suzuki; Y. Akashi; D. Kato; Y. Takeuchi; Y. Kiyasu; N. Terada; Y. Kurihara; M. Kuwahara; S. Muramatsu; A. Ueda; S. Notake; K. Nakamura | Suzuki, et. al. May 2022 | Analytical performance of rapid antigen tests for the detection of SARS-CoV-2 during widespread circulation of the Omicron variant | 2022-05-02 | QuickNavi-COVID19 Ag |  |  | Symptomatic | 366 | 1 | 530 | 22 | 94.3 | 91.5,96.4 | 99.8 | 99.0,100.0
Amy | Sherli | 244 | H. Suzuki; Y. Akashi; D. Kato; Y. Takeuchi; Y. Kiyasu; N. Terada; Y. Kurihara; M. Kuwahara; S. Muramatsu; A. Ueda; S. Notake; K. Nakamura | Suzuki, et. al. May 2022 | Analytical performance of rapid antigen tests for the detection of SARS-CoV-2 during widespread circulation of the Omicron variant | 2022-05-03 | QuickNavi-COVID19 Ag |  |  | Asymptomatic | 27 | 2 | 123 | 2 | 93.1 | 77.2,99.2 | 98.4 | 94.3,99.8
Amy | Sherli | 244 | H. Suzuki; Y. Akashi; D. Kato; Y. Takeuchi; Y. Kiyasu; N. Terada; Y. Kurihara; M. Kuwahara; S. Muramatsu; A. Ueda; S. Notake; K. Nakamura | Suzuki, et. al. May 2022 | Analytical performance of rapid antigen tests for the detection of SARS-CoV-2 during widespread circulation of the Omicron variant | 2022-05-04 | QuickNavi-COVID19 Ag |  |  | CT <20 | 181 |  |  | 2 | 98.9 | 96.1,99.9
Amy | Sherli | 244 | H. Suzuki; Y. Akashi; D. Kato; Y. Takeuchi; Y. Kiyasu; N. Terada; Y. Kurihara; M. Kuwahara; S. Muramatsu; A. Ueda; S. Notake; K. Nakamura | Suzuki, et. al. May 2022 | Analytical performance of rapid antigen tests for the detection of SARS-CoV-2 during widespread circulation of the Omicron variant | 2022-05-05 | QuickNavi-COVID19 Ag |  |  | CT 20-24 | 177 |  |  | 4 | 97.8 | 94.4,99.4
Amy | Sherli | 244 | H. Suzuki; Y. Akashi; D. Kato; Y. Takeuchi; Y. Kiyasu; N. Terada; Y. Kurihara; M. Kuwahara; S. Muramatsu; A. Ueda; S. Notake; K. Nakamura | Suzuki, et. al. May 2022 | Analytical performance of rapid antigen tests for the detection of SARS-CoV-2 during widespread circulation of the Omicron variant | 2022-05-06 | QuickNavi-COVID19 Ag |  |  | CT 25-29 | 24 |  |  | 4 | 85.7 | 67.3,96.0
Amy | Sherli | 244 | H. Suzuki; Y. Akashi; D. Kato; Y. Takeuchi; Y. Kiyasu; N. Terada; Y. Kurihara; M. Kuwahara; S. Muramatsu; A. Ueda; S. Notake; K. Nakamura | Suzuki, et. al. May 2022 | Analytical performance of rapid antigen tests for the detection of SARS-CoV-2 during widespread circulation of the Omicron variant | 2022-05-07 | QuickNavi-COVID19 Ag |  |  | CT >= 30 | 9 |  |  | 10 | 47.4 | 24.4,71.1
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-01 | RapidTesta SARS-CoV-2 |  | Without immunoassy device | Overall | 53 | 8 | 1045 | 21 | 71.6 | 59.9,81.5 | 99.2 | 98.5,99.7 | 86.9 | 98 | PPV CI95 [75.8-94.2] NPV CI95 [97.0-98.8]
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-02 | RapidTesta SARS-CoV-2 |  | With immunoassy device | Overall | 58 | 25 | 1028 | 16 | 78.4 | 67.3,87.1 | 97.6 | 96.5,98.5 | 69.9 | 98.5 | PPV CI95 [58.8-79.5] NPV CI95 [97.5-99.1]
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-03 | RapidTesta SARS-CoV-2 |  | Without immunoassy device | Symptomatic | 40 | 5 | 750 | 7 | 85.1 | 71.7,93.8 | 99.3 | 98.5,99.8
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-04 | RapidTesta SARS-CoV-2 |  | With immunoassy device | Symptomatic | 42 | 17 | 738 | 5 | 89.4 | 76.9,96.5 | 97.7 | 96.4,98.7
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-05 | RapidTesta SARS-CoV-2 |  | Without immunoassy device | CT<20 | 27 |  |  | 1 | 96.4 | 81.7,99.99
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-06 | RapidTesta SARS-CoV-2 |  | With immunoassy device | CT<20 | 28 |  |  | 0 | 100 | 87.7,100.0
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-07 | RapidTesta SARS-CoV-2 |  | Without immunoassy device | CT 20-29 | 26 |  |  | 6 | 81.3 | 59.0,91.7
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-08 | RapidTesta SARS-CoV-2 |  | With immunoassy device | CT 20-29 | 27 |  |  | 5 | 84.4 | 63.1,93.9
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-09 | RapidTesta SARS-CoV-2 |  | Without immunoassy device | CT>=30 | 0 |  |  | 14 | 0 | 0,19.5
Amy | Sherli | 271 | H. Suzuki; Y. Akashi; A. Ueda; Y. Kiyasu; Y. Takeuchi; Y. Maehara; Y. Ochiai; S. Okuyama; S. Notake; K. Nakamura; H. Ishikawa | Suzuki, et. al. Oct 2021 | Diagnostic performance of a novel digital immunoassay (RapidTesta SARS-CoV-2): A prospective observational study with nasopharyngeal samples | 2021-10-10 | RapidTesta SARS-CoV-2 |  | With immunoassy device | CT>=30 | 3 |  |  | 11 | 24.1 | 3.8,43.4
Amy | Sherli | 296 | Terpos E. | Terpos, et. al. May 2021 | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 | 2021-05-01 | Custom/Novel/In-house |  | Custom (Zhuhai Lituo Biotechnology) | NP Ct<25 | 70 | 1 | 243 | 0 | 100 | 98.79,100 | 99.6 | 98.07,99.91 | 98.6 | 100 |  | 100 | 99.5902 | 98.5915 | 100
Amy | Sherli | 296 | Terpos E. | Terpos, et. al. May 2021 | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 | 2021-05-01 | Custom/Novel/In-house |  | Custom (Zhuhai Lituo Biotechnology) | NS Ct<25 | 61 | 0 | 250 | 0 | 100 | 94.08,100 | 100 | 98.49,100 | 100 | 100 |  | 100 | 100 | 100 | 100
Amy | Sherli | 296 | Terpos E. | Terpos, et. al. May 2021 | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 | 2021-05-01 | Custom/Novel/In-house |  | Custom (Zhuhai Lituo Biotechnology) | NP Ct<33 | 99 | 1 | 243 | 1 | 99 | 97.28,99.64 | 99.6 | 98.18,99.91 | 99 | 99.6 |  | 99 | 99.5902 | 99 | 99.5902
Amy | Sherli | 296 | Terpos E. | Terpos, et. al. May 2021 | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 | 2021-05-01 | Custom/Novel/In-house |  | Custom (Zhuhai Lituo Biotechnology) | NS Ct<33 | 99 | 0 | 250 | 4 | 96.12 | 90.43,98.49 | 100 | 98.49,100 | 100 | 98.4 |  | 96.1165 | 100 | 100 | 98.4252
Amy | Sherli | 296 | Terpos E. | Terpos, et. al. May 2021 | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 | 2021-05-01 | Custom/Novel/In-house |  | Custom (Zhuhai Lituo Biotechnology) | NS Ct<37 | 100 | 0 | 250 | 9 | 91.74 | 85.05,95.60 | 100 | 98.49,100 | 100 | 96.5 |  | 91.7431 | 100 | 100 | 96.5251
Amy | Sherli | 296 | Terpos E. | Terpos, et. al. May 2021 | Clinical Application of a New SARS-CoV-2 Antigen Detection Kit (Colloidal Gold) in the Detection of COVID-19 | 2021-05-01 | Custom/Novel/In-house |  | Custom (Zhuhai Lituo Biotechnology) | NP Ct<40 | 102 | 1 | 243 | 12 | 89.47 | 85.86,92.24 | 99.6 | 98.22,99.91 | 99.03 | 95.3 |  | 89.4737 | 99.5902 | 99.0291 | 95.2941
Amy | Sherli | 297 | Thakur P. | Thakur, et. al. Nov 2021 | Utility of Antigen-Based Rapid Diagnostic Test for Detection of SARS-CoV-2 Virus in Routine Hospital Settings | 2021-11-01 | PathoCatch (Accucare) |  |  | Overall | 29 | 1 | 592 | 55 | 34.5 | 24.5,45.6 | 99.8 | 99.1,100 | 96.6 | 91.5 |  | 34.5238 | 99.8314 | 96.6667 | 91.4992
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-09 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | CT <25 | 123 | 45 | 362 | 2 | 98.4 | 94.3,99.8 | 88.9 | 85.5,91.8 | 73.2 | 99.5 | PPV CI95 [67.5-78.3] NPV CI95 [97.9-99.9]
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-10 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | CT <20 | 51 | 117 | 364 | 0 | 100 | 93.0,100.0 | 75.7 | 71.6,79.5 | 30.4 | 100 | PPV CI95 [27.1-33.8]
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptomatic | 171 | 3 | 325 | 42 | 80.3 | 74.3,85.4 | 99.1 | 97.4,99.8 | 98.3 | 88.6 | PPVCI95 [94.7-99.4] PPV CI95 [85.5-91.0]
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-02 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall-Emergency Dept | 106 | 3 | 157 | 30 | 77.9 | 70.0,84.6 | 98.1 | 94.6,99.6 | 97.3 | 84 | PPV CI95 [92.0-99.1] NPV CI95 [79.2-87.8]
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-03 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall-Primary Health Care | 65 | 0 | 168 | 12 | 84.4 | 74.4,91.7 | 100 | 97.8.100.0 | 100 | 93.3 | NPV CI95 [89.3-95.9]
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-04 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptoms <3 days | 116 | 1 | 252 | 24 | 82.9 | 75.6,88.7 | 99.6 | 97.8,100.0 | 99.2 | 91.3 | PPV CI95 [94.3-99.9] NPV CI95 [87.9-93.8]
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-05 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptoms 4-7 days | 40 | 2 | 47 | 9 | 81.6 | 68.0,91.2 | 95.9 | 86.0,99.5 | 95.2 | 83.9 | PPV CI95 [83.6-98.7] NPV CI95 [74.3-90.4]
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-06 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptoms >7 days | 15 | 0 | 26 | 9 | 62.5 | 40.6,81.2 | 100 | 86.8,100.0 | 100 | 74.3 | NPV CI95 [63.3-82.9]
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-07 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | CT <40 | 165 | 3 | 327 | 37 | 81.7 | 75.7,86.8 | 99.1 | 97.4,99.8 | 98.2 | 89.8 | PPV CI95 [94.7-99.4] NPV CI95 [86.9-92.2]
Amy | Sherli | 248 | R. Thell; V. Kallab; W. Weinhappel; W. Mueckstein; L. Heschl; M. Heschl; S. Korsatko; F. Toedling; A. Blaschke; T. Herzog; A. Klicpera; C. Koeller; M. Haugk; A. Kreil; A. Spiel; P. Kreuzer; R. Krause; C. Sebesta; S. Winkler; B. Laky; M. Szell | Thell, et. al. Nov 2021 | Evaluation of a novel, rapid antigen detection test for the diagnosis of SARS-CoV-2 | 2021-11-08 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | CT <30 | 160 | 8 | 351 | 13 | 92.5 | 87.5,95.9 | 97.8 | 95.7,99.0 | 95.2 | 96.4 | PPV CI95 [91.0-97.6] NPV CI95 [94.1-97.9]
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-01 | PanBio (Abbott) |  |  | Overall | 256 | 9 | 579 | 216 | 54.2 | 51.2,57.2 | 98.5 | 97.7,99.2 | 96.6 | 72.8 | PPV CI95 [95.5-97.7] NPV CI95 [70.2-75.5]
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-02 | PanBio (Abbott) |  |  | Symptoms <7 days | 190 |  |  | 84 | 69.1 | 65.7,72.5 | 98.4 | 97.5,99.3 | 96.5 | 83.4 | PPV CI95 [95.1-97.8] NPV CI95 [80.7-86.2]
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-03 | PanBio (Abbott) |  |  | Symptoms 7-14 days | 43 |  |  | 56 | 43.6 | 35.3,51.5 | 97.8 | 95.4,100 | 97.8 | 44.1 | PPV CI95 [95.4-100] NPV CI95 [36.1-52.1]
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-04 | PanBio (Abbott) |  |  | Symptoms >7days |  |  |  |  | 36.7 | 30.1,43.3 | 98.4 | 96.7,100 | 98.1 | 41.7 | PPV CI95 [96.2-100] NPV CI95 [34.9-48.5]
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-05 | PanBio (Abbott) |  |  | Asymptomatic |  |  |  |  | 26.3 | 19.1,33.4 | 98.9 | 97.2,100 | 93.8 | 67.7 | PPV CI95 [89.8-97.7] NPV CI95 [60.1-75.3]
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-06 | PanBio (Abbott) |  |  | Symptomatic |  |  |  |  | 58.1 | 54.9,61.3 | 98.4 | 97.6,99.2 | 96.8 | 73.8 | PPV CI95 [95.6-97.9] NPV CI95 [70.9-76.7]
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-07 | PanBio (Abbott) |  |  | CT <15 | 26 |  |  | 2 | 93 | 83.0,100.0
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-08 | PanBio (Abbott) |  |  | CT 15-20 | 66 |  |  | 9 | 88 | 81.0,95.0
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-09 | PanBio (Abbott) |  |  | CT 21-25 | 67 |  |  | 25 | 73 | 64.0,82.0
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-10 | PanBio (Abbott) |  |  | CT 26-30 | 29 |  |  | 58 | 33 | 23.0,43.0
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-11 | PanBio (Abbott) |  |  | CT 31-35 | 14 |  |  | 53 | 21 | 11.0,31.0
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-12 | PanBio (Abbott) |  |  | CT 36-39 | 5 |  |  | 32 | 14 | 2.0,25.0
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-13 | PanBio (Abbott) |  |  | CT =< 25 | 159 |  |  | 35 | 82 | 76.0,87.0
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-14 | PanBio (Abbott) |  |  | CT>25 | 48 |  |  | 144 | 25 | 19.0,31.0
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-15 | PanBio (Abbott) |  |  | All subjects with CT values | 207 |  |  | 179 | 54 | 48.0,58.0
Amy | Sherli | 249 | I. Thirion-Romero; D. S. Guerrero-Zúñiga; D. A. Arias-Mendoza; D. D. P. Cornejo-Juárez; D. P. Meza-Meneses; D. D. S. Torres-Erazo; D. T. Hernández-Gilsoul; D. A. Galindo-Fraga; D. I. Villegas-Mota; D. J. Sepúlveda-Delgado; D. S. Ávila-Ríos; D. E. Becerril-Vargas; R. Fernández-Plata; T. M. Pérez-Kawabe; D. A. Coeto-Cano; D. J. A. Vázquez-Pérez; D. S. Kawa-Karasik; D. G. Reyes-Terán; D. J. R. Pérez-Padilla | Thirion-Romero, et. al. Oct 2021 | Evaluation of Panbio rapid antigen test for SARS-CoV-2 in symptomatic patients and their contacts: a multicenter study | 2021-10-16 | PanBio (Abbott) |  |  | Symptoms > 14 days | 8 |  |  | 32 | 20 | 9.0,36.0
Amy | Sherli | 251 | V. Tonelotto; A. Davini; L. Cardarelli; M. Calderone; P. Marin | Tonelotto, et. al. Jan 2022 | Efficacy of Fluorecare SARS-CoV-2 Spike Protein Test Kit for SARS-CoV-2 detection in nasopharyngeal samples of 121 individuals working in a manufacturing company | 2022-01-01 | Fluorecare (Colloidal Gold/Fluorescent) SARS-CoV-2 Spike Protein Test kit (Shenzen Microprofit) |  |  | Overall | 9 | 0 | 242 | 2 | 84.6 | 54.6,98.1 | 100 | 98.6,100.0
Amy | Sherli | 299 | Toptan T. | Toptan, et. al. Dec 2020 | Evaluation of a SARS-CoV-2 rapid antigen test: Potential to help reduce community spread? | 2020-12-01 | Rida Quick SARS-CoV-2 (R-Biopharm) |  |  | Overall Berlin | 45 | 0 | 9 | 13 | 77.6 |  | 100 |  | 100 | 40.1 |  | 77.5862 | 100 | 100 | 40.9091
Amy | Sherli | 299 | Toptan T. | Toptan, et. al. Dec 2020 | Evaluation of a SARS-CoV-2 rapid antigen test: Potential to help reduce community spread? | 2020-12-01 | Rida Quick SARS-CoV-2 (R-Biopharm) |  |  | Overall Frankfurt | 16 | 0 | 38 | 16 | 50 |  | 100 |  | 100 | 70.3 |  | 50 | 100 | 100 | 70.3704
Amy | Sherli | 305 | Trobajo-Sanmartín | Trobajo-Sanmartín, et. al. Mar 2021 | Evaluation of the rapid antigen test CerTest SARS-CoV-2 as an alternative COVID-19 diagnosis technique | 2021-03-01 | CerTest SARS-CoV-2 (Certest Biotech) |  |  | Overall | 63 | 0 | 160 | 17 | 78.75 | 67.89,86.79 | 100 | 97.08,99.94 | 100 | 90.4 |  | 78.75 | 100 | 100 | 90.3955
Amy | Sherli | 306 | Turcato Gianni | Turcato, et. al. Dec 2020 | Clinical application of a rapid antigen test for the detection of SARS-CoV-2 infection in symptomatic and asymptomatic patients evaluated in the emergency department: A preliminary report | 2020-12-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 179 | 30 | 3157 | 44 | 80.3 | 74.9,85.4 | 99.1 | 98.6,99.3 | 85.64 | 98.6 |  | 80.2691 | 99.0587 | 85.6459 | 98.6254
Amy | Sherli | 306 | Turcato Gianni | Turcato, et. al. Dec 2020 | Clinical application of a rapid antigen test for the detection of SARS-CoV-2 infection in symptomatic and asymptomatic patients evaluated in the emergency department: A preliminary report | 2020-12-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic | 27 | 10 | 2355 | 27 | 50 | 36.0,63.0 | 99.6 | 99.1,99.9 | 72.97 | 98.87 |  | 50 | 99.5772 | 72.973 | 98.8665
Amy | Sherli | 306 | Turcato Gianni | Turcato, et. al. Dec 2020 | Clinical application of a rapid antigen test for the detection of SARS-CoV-2 infection in symptomatic and asymptomatic patients evaluated in the emergency department: A preliminary report | 2020-12-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptomatic | 152 | 20 | 802 | 17 | 89.9 | 85.4,94.4 | 97.6 | 96.5,98.5 | 88.4 | 97.9 |  | 89.9408 | 97.5669 | 88.3721 | 97.9243
Amy | Sherli | 252 | G. Turcato; A. Zaboli; N. Pfeifer; S. Sibilio; G. Tezza; A. Bonora; L. Ciccariello; D. Ausserhofer | Turcato, et. al. Jan 2022 | Rapid antigen test to identify COVID-19 infected patients with and without symptoms admitted to the Emergency Department | 2022-01-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 329 | 32 | 3470 | 68 | 82.9 | 81.0,84.8 | 99.1 | 98.8,99.3 | 91.1 | 98.1 | PPV CI95 [89.7-92.5] NPV CI95 [97.8-98.3]
Amy | Sherli | 252 | G. Turcato; A. Zaboli; N. Pfeifer; S. Sibilio; G. Tezza; A. Bonora; L. Ciccariello; D. Ausserhofer | Turcato, et. al. Jan 2022 | Rapid antigen test to identify COVID-19 infected patients with and without symptoms admitted to the Emergency Department | 2022-01-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Symptomatic | 264 | 21 | 876 | 30 | 89.8 | 88.0,91.5 | 97.6 | 97.1,98.1 | 92.6 | 96.7 | PPV CI95 [91.0-94.1] NPV CI95 [96.0-97.2]
Amy | Sherli | 252 | G. Turcato; A. Zaboli; N. Pfeifer; S. Sibilio; G. Tezza; A. Bonora; L. Ciccariello; D. Ausserhofer | Turcato, et. al. Jan 2022 | Rapid antigen test to identify COVID-19 infected patients with and without symptoms admitted to the Emergency Department | 2022-01-03 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Asymptomatic | 65 | 11 | 2594 | 38 | 63.1 | 58.4,67.8 | 99.6 | 99.5,99.7 | 85.5 | 98.5 | PPV CI95 [81.5-89.5] NPV CI95 [98.3-98.7]
Ellyn | Sherli | 310 | Van der Moeren N. | Van der Moeren, et. al. May 2021 | Evaluation of the test accuracy of a SARS-CoV-2 rapid antigen test in symptomatic community dwelling individuals in the Netherlands | 2021-05-01 | BD Veritor COVID-19 Rapid Antigen Test (Becton-Dickinson) |  |  | Overall | 16 | 0 | 334 | 1 | 94.1 | 71.1,100 | 100 | 98.9, 100 | 100 | 99.7 |  | 94.1176 | 100 | 100 | 99.7015
Ellyn | Sherli | 310 | Van der Moeren N. | Van der Moeren, et. al. May 2021 | Evaluation of the test accuracy of a SARS-CoV-2 rapid antigen test in symptomatic community dwelling individuals in the Netherlands | 2021-05-01 | BD Veritor COVID-19 Rapid Antigen Test (Becton-Dickinson) |  |  | Analyzer | 16 | 2 | 332 | 1 | 94.1 | 71.1,100 | 99.4 | 97.9,100 | 88.9 | 99.7 |  | 94.1176 | 99.4012 | 88.8889 | 99.6997
Ellyn | Sherli | 310 | Van der Moeren N. | Van der Moeren, et. al. May 2021 | Evaluation of the test accuracy of a SARS-CoV-2 rapid antigen test in symptomatic community dwelling individuals in the Netherlands | 2021-05-01 | BD Veritor COVID-19 Rapid Antigen Test (Becton-Dickinson) |  |  | Visual | 16 | 0 | 334 | 1 | 94.1 | 71.1,100 | 100 | 98.9,100 | 100 | 99.7 |  | 94.1176 | 100 | 100 | 99.7015
Amy | Sherli | 253 | E. Van Honacker; K. Van Vaerenbergh; A. Boel; H. De Beenhouwer; I. Leroux-Roels; L. Cattoir | Van Honacker, et. al. Mar 2021 | COMPARISON OF FIVE SARS-COV-2 RAPID ANTIGEN TESTS IN A HOSPITAL SETTING AND PERFORMANCE OF ONE ANTIGEN ASSAY IN ROUTINE PRACTICE. A USEFUL TOOL TO GUIDE ISOLATION PRECAUTIONS? | 2021-03-01 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | Overall | 200 | 12 | 3814 | 169 | 54.2 |  | 99.7
Amy | Sherli | 253 | E. Van Honacker; K. Van Vaerenbergh; A. Boel; H. De Beenhouwer; I. Leroux-Roels; L. Cattoir | Van Honacker, et. al. Mar 2021 | COMPARISON OF FIVE SARS-COV-2 RAPID ANTIGEN TESTS IN A HOSPITAL SETTING AND PERFORMANCE OF ONE ANTIGEN ASSAY IN ROUTINE PRACTICE. A USEFUL TOOL TO GUIDE ISOLATION PRECAUTIONS? | 2021-03-02 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT <20 | 61 |  |  | 0 | 100
Amy | Sherli | 253 | E. Van Honacker; K. Van Vaerenbergh; A. Boel; H. De Beenhouwer; I. Leroux-Roels; L. Cattoir | Van Honacker, et. al. Mar 2021 | COMPARISON OF FIVE SARS-COV-2 RAPID ANTIGEN TESTS IN A HOSPITAL SETTING AND PERFORMANCE OF ONE ANTIGEN ASSAY IN ROUTINE PRACTICE. A USEFUL TOOL TO GUIDE ISOLATION PRECAUTIONS? | 2021-03-03 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT 20- <26 | 100 |  |  | 8 | 92.6
Amy | Sherli | 253 | E. Van Honacker; K. Van Vaerenbergh; A. Boel; H. De Beenhouwer; I. Leroux-Roels; L. Cattoir | Van Honacker, et. al. Mar 2021 | COMPARISON OF FIVE SARS-COV-2 RAPID ANTIGEN TESTS IN A HOSPITAL SETTING AND PERFORMANCE OF ONE ANTIGEN ASSAY IN ROUTINE PRACTICE. A USEFUL TOOL TO GUIDE ISOLATION PRECAUTIONS? | 2021-03-04 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT 26 - <30 | 35 |  |  | 29 | 54.7
Amy | Sherli | 253 | E. Van Honacker; K. Van Vaerenbergh; A. Boel; H. De Beenhouwer; I. Leroux-Roels; L. Cattoir | Van Honacker, et. al. Mar 2021 | COMPARISON OF FIVE SARS-COV-2 RAPID ANTIGEN TESTS IN A HOSPITAL SETTING AND PERFORMANCE OF ONE ANTIGEN ASSAY IN ROUTINE PRACTICE. A USEFUL TOOL TO GUIDE ISOLATION PRECAUTIONS? | 2021-03-05 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT 30 - <36 | 4 |  |  | 80 | 4.8
Amy | Sherli | 253 | E. Van Honacker; K. Van Vaerenbergh; A. Boel; H. De Beenhouwer; I. Leroux-Roels; L. Cattoir | Van Honacker, et. al. Mar 2021 | COMPARISON OF FIVE SARS-COV-2 RAPID ANTIGEN TESTS IN A HOSPITAL SETTING AND PERFORMANCE OF ONE ANTIGEN ASSAY IN ROUTINE PRACTICE. A USEFUL TOOL TO GUIDE ISOLATION PRECAUTIONS? | 2021-03-06 | Standard Q COVID-19 Ag (SD Biosensor) |  |  | CT >36 | 0 |  |  | 52 | 0
Amy | Sherli | 254 | T. von Ahnen; M. von Ahnen; U. Wirth; H. M. Schardey; S. Herdtle | von Ahnen, et. al. Sep 2021 | Evaluation of a rapid-antigen test for COVID-19 in an asymptomatic collective : A prospective study | 2021-09-01 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall | 12 | 0 | 906 | 1 | 92.3 | 78.0,100 | 100 | 100.0,100.0 | 100 | 99.9 | PPV CI95 [100.0-100.0] NPV CI95 [99.7-100.1]
Amy | Sherli | 258 | C. Wertenauer; G. Brenner Michael; A. Dressel; C. Pfeifer; U. Hauser; E. Wieland; C. Mayer; C. Mutschmann; M. Roskos; H. J. Wertenauer; A. P. Moissl; S. Lorkowski; W. März | Wertenauer, et. al. Mar 2022 | Diagnostic Performance of Rapid Antigen Testing for SARS-CoV-2: The COVid-19 AntiGen (COVAG) study | 2022-03-01 | PanBio (Abbott) |  |  | Overall |  |  |  |  | 56.8 |  | 99.9 |  | 99 | 92.8
Amy | Sherli | 258 | C. Wertenauer; G. Brenner Michael; A. Dressel; C. Pfeifer; U. Hauser; E. Wieland; C. Mayer; C. Mutschmann; M. Roskos; H. J. Wertenauer; A. P. Moissl; S. Lorkowski; W. März | Wertenauer, et. al. Mar 2022 | Diagnostic Performance of Rapid Antigen Testing for SARS-CoV-2: The COVid-19 AntiGen (COVAG) study | 2022-03-02 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Overall |  |  |  |  | 60.4 |  | 99.7 |  | 97.6 | 93.3
Amy | Sherli | 258 | C. Wertenauer; G. Brenner Michael; A. Dressel; C. Pfeifer; U. Hauser; E. Wieland; C. Mayer; C. Mutschmann; M. Roskos; H. J. Wertenauer; A. P. Moissl; S. Lorkowski; W. März | Wertenauer, et. al. Mar 2022 | Diagnostic Performance of Rapid Antigen Testing for SARS-CoV-2: The COVid-19 AntiGen (COVAG) study | 2022-03-03 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Symptomatic |  |  |  |  | 75.2 |  | 99.6 |  | 98.2 | 93.2
Amy | Sherli | 258 | C. Wertenauer; G. Brenner Michael; A. Dressel; C. Pfeifer; U. Hauser; E. Wieland; C. Mayer; C. Mutschmann; M. Roskos; H. J. Wertenauer; A. P. Moissl; S. Lorkowski; W. März | Wertenauer, et. al. Mar 2022 | Diagnostic Performance of Rapid Antigen Testing for SARS-CoV-2: The COVid-19 AntiGen (COVAG) study | 2022-03-04 | PanBio (Abbott) |  |  | Symptomatic |  |  |  |  | 74.3 |  | 99.7 |  | 98.8 | 92.9
Amy | Sherli | 258 | C. Wertenauer; G. Brenner Michael; A. Dressel; C. Pfeifer; U. Hauser; E. Wieland; C. Mayer; C. Mutschmann; M. Roskos; H. J. Wertenauer; A. P. Moissl; S. Lorkowski; W. März | Wertenauer, et. al. Mar 2022 | Diagnostic Performance of Rapid Antigen Testing for SARS-CoV-2: The COVid-19 AntiGen (COVAG) study | 2022-03-05 | Roche SARS-CoV-2 Rapid Antigen Test (Roche) |  |  | Asymptomatic |  |  |  |  | 31.9 |  | 99.8 |  | 94.9 | 93.4
Amy | Sherli | 258 | C. Wertenauer; G. Brenner Michael; A. Dressel; C. Pfeifer; U. Hauser; E. Wieland; C. Mayer; C. Mutschmann; M. Roskos; H. J. Wertenauer; A. P. Moissl; S. Lorkowski; W. März | Wertenauer, et. al. Mar 2022 | Diagnostic Performance of Rapid Antigen Testing for SARS-CoV-2: The COVid-19 AntiGen (COVAG) study | 2022-03-06 | PanBio (Abbott) |  |  | Asymptomatic |  |  |  |  | 23.3 |  | 100 |  | 98.2 | 92.7